# Supplementary Material

A new versatile MR-guided High-Intensity Focused Ultrasound (HIFU) device for the treatment of musculoskeletal tumors

Paolo Cabras 1,2, Pierre Auloge3, Fabrice Bing1,4, Pramod Prabakhar Rao1,3, Stéphanie Hoarau2, Erik Dumont2, Alexandre Durand5, Benjamin Maurin5, Benoit Wach1, Loïc Cuvillon1, Elodie Breton1, Afshin Gangi1,3, and Jonathan Vappou1

1ICube, Université de Strasbourg, CNRS, UMR 7357, Strasbourg, France

2Image Guided Therapy, Pessac, France

3Department of Interventional Imaging, Hôpitaux Universitaires de Strasbourg, Strasbourg, France

4Radiology Department, Hôpital d’Annecy, Metz-Tessy, France

5Axilum Robotics, Schiltigheim, France

Corresponding author: jvappou@unistra.fr

- System robustness assessment

Fig S1: Experimental setup for assessment of mechanical robustness: A 2 kg load is attached to the supporting structure. Two configurations are tested, namely, arc-like (a), and strongly inclined (b). The load does not touch the table.

Fig S2: Position (a) and Orientation (b) of the holding system over time (45 mins) for three different configurations: ‘arc configuration’ (cf. Fig S1 (a)) in blue, ‘strongly inclined configuration’ (cf. Fig S1 (b)) in orange. Gray curves correspond to the baseline drift, measured by letting the system ‘at rest’ over the table, without any load.

- Registration process

Fig S3: Schematic view of the different reference frames involved in the registration process.

During the registration process, the aim is to find the transform between the Polaris reference system (with respect to whom the transducer actual pose is measured - “trsd-r”) and the MR reference system (with respect to whom the treatment target and the transducer desired - “trsd-d” -  pose are defined). Colors in figure S3 are used to help identify the different reference frames.

The hybrid tracker is composed of four spheres visible in both modalities. This element provides the missing link between the two measurement systems (MR and Polaris) allowing computing the desired transform.

MRTtrsd-r= PTMR-1PTtrsd-r where PTMR can be computed as

PTMR= PTmarkers-P markers-PTmarkers-MR MRTmarkers-MR-1.

Where:

PTtrsd-r: is known. It is the current pose of the transducer measured and returned by the Polaris system.

PTmarkers-P: is known. It is the pose of the hybrid tracker (IR reflective sphere) given by the Polaris system.

MRTmarkers-MR: is computed (during registration process). The pose of the hybrid tracker with respect to MRI can be computed fitting the tracker model to the MR landmarks.

markers-PTmarkers-MR: is known. Given the particular design of the hybrid tracker, this transform expressing the spatial relationship between the MR visible markers and the Polaris visible marker is the identity.

- Methods for error calculation

The 3d error E for each tested pose is the 2-norm of the difference vector between the desired Xd and the measured Xm position:

E=Xd-Xm2

The angular error is the angle difference between the rotation matrices (desired vs measured) , computed as follows:

- First the difference rotation matrix is computed as R=RdTRm (where d and m are referring to desired and measured respectively).

- The rotation expressed by R is then expressed as axis-angle rotation where the angle θ can be retrieved according to the formula: Trace(R) = 1 + 2cos(θ)

This gives the absolute angle of a rotation matrix. It is the minimum angle needed to rotate from the identity matrix to R, and hence it is a useful pseudo-norm for 3D rotations.