Supplementary Table S2. Current dialyzer classification based on urea and β2-microglobulin clearance and albumin sieving coefficient in Japan since 2013

| Type of dialyzers | I |  | II |  | S |
| --- | --- | --- | --- | --- | --- |
|  | Ia | Ib | IIa | IIb |  |
| Measurement condition* |  |  |  |  |  |
| Surface area, m2 | 1.5 |  |  |  |  |
| Blood flow rate, mL/min | 200 ± 4 |  |  |  |  |
| Dialysate flow rate, mL/min | 500 ± 15 |  |  |  |  |
| Substitution flow rate, mL/min | 15 ± 1 |  |  |  |  |
| Performance standards |  |  |  |  |  |
| Urea clearance (mL/min) | ≥150 |  | ≥185 |  | 150 |
| β2MG clearance (mL/min) | <70 |  | ≥70 |  | 0 |
| Albumin sieving coefficient | <0.03 | ≥0.03 | <0.03 | ≥0.03 |  |

*The average temperature of the bovine blood and dialysate solution in the storage tank should be 37°C ± 1°C. Measurements were taken 60 and 240 min after the start of the experiment. β2MG, β2-microglobulin.