Supplementary Table S2. Current dialyzer classification based on urea and β2-microglobulin clearance and albumin sieving coefficient in Japan since 2013

	Type of dialyzers
	I
	II
	S

	
	Ia
	Ib
	IIa
	IIb
	

	Measurement condition*
	　

	    Surface area, m2
	1.5

	    Blood flow rate, mL/min
	200 ± 4

	    Dialysate flow rate, mL/min
	500 ± 15

	    Substitution flow rate, mL/min
	15 ± 1

	Performance standards
	　
	　
	　

	    Urea clearance (mL/min)
	≥150
	≥185
	150

	    β2MG clearance (mL/min)
	<70
	≥70
	0

	    Albumin sieving coefficient
	<0.03
	≥0.03
	<0.03
	≥0.03
	


*The average temperature of the bovine blood and dialysate solution in the storage tank should be 37°C ± 1°C. Measurements were taken 60 and 240 min after the start of the experiment. β2MG, β2-microglobulin.


