Scan analysis conditions 
[bookmark: OLE_LINK3][bookmark: OLE_LINK4]LC Conditions:
· Mobile phase: 
	Solvent C
	ACN + 0.1 % Formic Acid
	

	Solvent B
	 + 0.1 % Formic Acid
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· Column: C18 3µ, 4.6×150 mm
· Column temperature: 35 ̊C
· Flow rate: 0.3 ml/min
· Injection volume:  5  µL

MASS conditions: 
	Mode: ESI -/+

	Cone Volt: 25 V

	Capillary Volt: 4 kV

	Extractor: 3 V

	RF Lens: 0.1 V

	Collision Energy: 30 eV

	Gas nebulizer:  (grade 5)

	Flow gas: 300 L/h

	[bookmark: OLE_LINK5][bookmark: OLE_LINK6]Source temperature: 120   ̊C

	Desolvation temperature: 300   ̊C



Device model:
HPLC: Waters Alliance 2695 
Mass Spectrometer: Micromass Quattro micro API 
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