| Total samples tested (n=28) | Exon 5 (HEX) Ct value |  |  |  |  | Exon 7 (ROX) Ct value |  |  |  |  |  | Exon 10 (FAM) Ct value |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates |  | Avg | Stdev | Ct Min | Ct Max |  |  |
| Mother RhD- Fetus RhD+ (n=18) | 36/36 | 33.47 | 1.42 | 30.07 | 36.31 | 36/36 | 31.96 | 1.35 | 28.6 | 34.18 | 36/36 |  | 32.56 | 1.39 | 29.06 | 35.05 |  |  |
|  | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates |  | Avg | Stdev | Ct Min | Ct Max |  |  |
| Mother RhD- Fetus RhD- (n=10) | 3/20 | 36.30 | 1.00 | 35.59 | 37 | 4/20 | 36.02 | 0.97 | 34.98 | 37.11 | 2/20 |  | 35.13 | 1.63 | 33.28 | 36.37 |  |  |

RhD-  = RhD-negative; RhD+ = RhD-positive; Avg = Average; Stdev = Standard Deviation; Ct = Cycle Threshold; Min = Minimum; Max = Maximum

Ct values for samples when both maternal and fetal samples are negative (bottom row) are only for the positive replicates

Table S1: Ct values of Eonis RhD qPCR kit using 20µL cfDNA sample volume

| Total samples tested (n=28) | Exon 5 (HEX) Ct value |  |  |  |  | Exon 7 (ROX) Ct value |  |  |  |  |  | Exon 10 (FAM) Ct value |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates |  | Avg | Stdev | Ct Min | Ct Max |  |  |
| Mother RhD- Fetus RhD+ (n=18) | 36/36 | 34.71 | 1.35 | 31.12 | 36.46 | 36/36 | 33.25 | 1.28 | 30.01 | 35.28 | 36/36 |  | 33.41 | 1.21 | 30.14 | 34.83 |  |  |
|  | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates | Avg | Stdev | Ct Min | Ct Max | Positive replicates |  | Avg | Stdev | Ct Min | Ct Max |  |  |
| Mother RhD- Fetus RhD- (n=10) | 6/20 | 37.48 | 1.28 | 36.22 | 38.78 | 3/20 | 35.73 | 0.63 | 35.28 | 36.17 | 2/20 |  | 35.50 | 1.29 | 34.2 | 37.68 |  |  |

RhD-  = RhD-negative; RhD+ = RhD-positive; Avg = Average; Stdev = Standard Deviation; Ct = Cycle Threshold; Min = Minimum; Max = Maximum

Ct values for samples when both maternal and fetal samples are negative (bottom row) are only for the positive replicates

Table S2: Ct values of Eonis RhD qPCR kit using 10µL cfDNA sample volume