Supplementary Information

For

Design and development of a portable multiwavelength LED-based diffuse reflectance spectroscopy tool for rapid breast cancer identification

Arif Mohd. Kamala, Kirteyman Singh Rajputa,ǂ, Apurva Dahakea,ǂ, Gayatri Gogoib, Ajay Krishnan Ajitha, VSN Sitaramgupta Vangetia, Dilip Kilingc, Jayant S. Vaidyad, and Hardik J. Pandyaa,e,*

aDepartment of Electronic Systems Engineering, Indian Institute of Science, Bangalore, India; bMedical Research Unit Laboratory and Department of Pathology, Assam Medical College, Dibrugarh, India;  cDepartment of Surgery, Assam Medical College, Dibrugarh, India; dDivision of Surgery and Interventional Science, University College London, London, UK; eDepartment of Design and Manufacturing, Indian Institute of Science, Bangalore, India.

*Correspondence: Hardik J. Pandya (hjpandya@iisc.ac.in)

ǂEqual Contribution

Tables:

Table S1: Clinical diagnosis and Histopathological impressions of the human breast tissue samples

| Sample No. | Clinical Diagnosis |  |  |  | Histopathology Impression |
| --- | --- | --- | --- | --- | --- |
|  | Menopause | Type of Breast Cancer | Age (Yrs.) | BMI (Kg/m2) |  |
| 1 | Post | Invasive lobular Carcinoma | 56 | 23.4 | Malignant |
|  |  |  |  |  | Adjacent Normal |
| 2 | Post | Invasive Ductal Carcinoma | 50 | 19.9 | Malignant |
|  |  |  |  |  | Adjacent Normal |

Table S2: Hyperparameters used in grid search during model tuning

| Algorithm | Hyperparameter | Range |
| --- | --- | --- |
| Logistic Regression | classifier__C | 0.001, 0.01, 0.1, 1, 10, 100 |
|  | classifier__penalty | 'l1', 'l2', 'elasticnet', None |
|  | classifier__solver | 'newton-cg', 'lbfgs', 'liblinear', 'sag', 'saga' |
|  | classifier__max_iter | 50, 100, 200, 500 |
| Random Forest | classifier__n_estimators | 100, 200, 300, 400, 500 |
|  | classifier__max_depth | 3, 4, 5, 6, None |
|  | classifier__min_samples_split | 2, 5, 10 |
|  | classifier__min_samples_leaf | 1, 2, 4 |
|  | classifier__max_features | 'sqrt', 'log2' |
|  | classifier__bootstrap | True, False |
| XGBoost | classifier__n_estimators | 50, 100, 200 |
|  | classifier__max_depth | 3, 5, 7 |
|  | classifier__learning_rate | 0.01, 0.1, 0.2 |
|  | classifier__subsample | 0.6, 0.8, 1.0 |
|  | classifier__colsample_bytree | 0.6, 0.8, 1.0 |
|  | classifier__gamma | 0, 0.1, 0.2 |
|  | classifier__reg_alpha | 0, 0.1, 1 |
|  | classifier__reg_lambda | 0, 0.1, 1 |
| CatBoost | classifier_iterations | 100, 200, 300 |
|  | classifier_depth | 4, 6, 8 |
|  | classifier__learning_rate | 0.01, 0.05, 0.1 |
|  | classifier__l2_leaf_reg | 1, 3, 5 |
|  | classifier__border_count | 32, 64, 128 |
|  | classifier__random_strength | 0.5, 1, 2 |
|  | classifier__bagging_temperature | 0, 0.5, 1 |
|  | classifier__verbose | False |
| SVM | classifier__C | 0.1, 1, 10, 100 |
|  | classifier__gamma | 'scale', 'auto', 0.01, 0.1, 1 |
|  | classifier__kernel | 'linear', 'rbf', 'poly' |
|  | classifier__degree | 2, 3, 4 |
|  | classifier__probability | True |
| KNN | classifier__n_neighbors | 3, 5, 7, 9, 11, 13 |
|  | classifier__weights | 'uniform', 'distance' |
|  | classifier__p | 1, 2 |
|  | classifier__metric | 'minkowski', 'cosine' |
| Gaussian Naive Bayes | classifier__var_smoothing | 1e-9, 1e-8, 1e-7, 1e-6, 1e-5 |

Table S3: Detected voltage values for formalin-fixed tissue samples (N=25 patients) at 850 nm, 940 nm, and 1050 nm.

|  | Detected Voltage (Vd) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Patient No. | 850 nm |  | 940 nm |  | 1050 nm |  |
|  | N | C | N | C | N | C |
| 1 | 1.62 | 2.81 | 0.86 | 1.74 | 0.69 | 1.12 |
| 2 | 1.45 | 2.49 | 0.89 | 1.90 | 0.69 | 0.94 |
| 3 | 1.87 | 2.36 | 1.19 | 1.34 | 0.88 | 0.89 |
| 4 | 2.15 | 2.43 | 1.24 | 1.43 | 0.80 | 0.94 |
| 5 | 1.34 | 2.34 | 0.74 | 1.81 | 0.59 | 1.23 |
| 6 | 2.14 | 2.82 | 1.16 | 1.86 | 0.69 | 1.12 |
| 7 | 1.83 | 2.97 | 0.90 | 1.89 | 0.57 | 1.24 |
| 8 | 2.04 | 2.73 | 0.99 | 1.29 | 0.62 | 0.99 |
| 9 | 2.12 | 2.62 | 1.17 | 1.65 | 0.81 | 0.91 |
| 10 | 2.08 | 3.14 | 1.21 | 1.89 | 0.74 | 1.22 |
| 11 | 2.10 | 2.72 | 1.15 | 1.84 | 0.69 | 1.25 |
| 12 | 2.11 | 2.79 | 1.24 | 1.53 | 0.72 | 0.87 |
| 13 | 2.21 | 2.82 | 1.18 | 1.72 | 0.63 | 0.78 |
| 14 | 2.27 | 2.78 | 1.32 | 1.52 | 0.71 | 0.91 |
| 15 | 2.14 | 2.70 | 1.22 | 1.67 | 0.83 | 0.86 |
| 16 | 2.23 | 2.61 | 1.16 | 1.48 | 0.74 | 0.83 |
| 17 | 2.11 | 2.74 | 1.24 | 1.54 | 0.85 | 0.69 |
| 18 | 2.21 | 2.81 | 1.41 | 1.49 | 0.62 | 0.84 |
| 19 | 2.32 | 2.77 | 1.39 | 1.48 | 0.84 | 0.86 |
| 20 | 0.98 | 1.58 | 0.38 | 0.91 | 0.18 | 0.19 |
| 21 | 2.6 | 3.62 | 1.55 | 1.89 | 0.82 | 1.25 |
| 22 | 2.14 | 3.71 | 0.89 | 1.96 | 0.29 | 1.54 |
| 23 | 2.28 | 3.65 | 0.87 | 1.84 | 0.44 | 0.95 |
| 24 | 2.47 | 2.96 | 1.17 | 1.49 | 0.49 | 0.79 |
| 25 | 1.97 | 3.63 | 1.24 | 1.89 | 0.40 | 1.53 |
| Mean | 2.03 | 2.82 | 1.11 | 1.64 | 0.65 | 0.99 |
| SEM | 0.07 | 0.10 | 0.05 | 0.05 | 0.04 | 0.06 |

Table S4: Detected voltage values for fresh tissue samples (N=6 patients) at 850 nm, 940 nm, and 1050 nm.

|  | Detected Voltage (VD) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Patient No. | 850 nm |  | 940 nm |  | 1050 nm |  |
|  | N | C | N | C | N | C |
| 1 | 2.37 | 2.28 | 1.16 | 1.35 | 0.36 | 0.91 |
| 2 | 1.54 | 2.79 | 0.93 | 1.61 | 0.96 | 1.36 |
| 3 | 2.28 | 2.62 | 1.77 | 1.86 | 0.73 | 0.91 |
| 4 | 1.19 | 2.92 | 0.62 | 1.63 | 0.92 | 1.41 |
| 5 | 0.98 | 2.71 | 0.46 | 1.83 | 0.86 | 1.19 |
| 6 | 0.84 | 2.85 | 0.48 | 1.70 | 0.34 | 1.46 |
| Mean | 1.53 | 2.7 | 0.9 | 1.66 | 0.7 | 1.21 |
| SEM | 0.27 | 0.09 | 0.26 | 0.08 | 0.11 | 0.10 |

Figures:

Fig. S1: Engineering design of Multispectral Pen and GUI. (a) Exterior view of the Pen, (b) Interior view of the Pen, (c) Different views of the Pen’s tip, (d)Associated parts of the Pen, and (e) Snapshot of the developed LabVIEW-based GUI.

Fig. S2: Detected voltage (V) of (A) Formalin-fixed tissue samples and (B) Fresh tissue samples obtained at 850 nm, 940 nm, and 1050 nm, respectively.

Fig. S3: Process flow chart of the algorithm and model selection.

(A)

(B)

(C)

(D)

(E)

Fig. S4: Adjacent normal tissue measurements with: (A) Black backing, (B) White backing. Malignant tissue measurements with: (C) Black backing, (D) White backing, and (E) Bar plot of comparison of mean reflectance values with black and white backing for adjacent normal and malignant tissue samples obtained at 850 nm, 940 nm, and 1050 nm, respectively.