Supplementary material

AI-assisted versus manual sustainability assessment of a high-throughput LC–MS/MS method for psychotropic and OTC drugs of abuse in human plasma

Hend Z. Yamani1*, Khaled Hesham2,Shereen M. Tawakkol2,, Lobna A. Hussein1, Nesma M. Fahmy2

1Pharmaceutical Analytical Chemistry Department, Faculty of Pharmacy, Ain Shams University, Cairo 11566, Egypt

2Pharmaceutical Chemistry Department, Faculty of Pharmacy, Ahram Canadian University, Giza, Egypt

*Corresponding author e-mail: hend.z.yamani@pharma.asu.edu.eg

Figure S1: The chemical structure of  DXM, PSE, OLA, and FLU

Table S1: Optimized LC-MS/MS Parameters for DXM, PSE, FLU and OLA.

| Analyte | Q1 (m/ z) | Q3 (m/ z) | Collision Energy (V) | Dwell Time (ms) |
| --- | --- | --- | --- | --- |
| DXM | 271.90 | 215.10 | -12 | 100 |
| PSE | 166.25 | 148.10 | -13 | 100 |
| OLA | 312.95 | 256.05 | -10 | 100 |
| FLU | 309.85 | 148.10 | -9 | 100 |
| Nebulizing gas flow | 3 L/min |  |  |  |
| Drying gas flow | 15 L/min |  |  |  |
| DL temperature | 250°C |  |  |  |
| Heat block temperature | 400°C |  |  |  |
| Interface voltage | 4.5 kV |  |  |  |

Table S2: Comparison of the proposed LC–MS/MS method with previously reported LC–MS/MS methods for simultaneous determination of olanzapine and fluoxetine in human plasma.

| Reference | Plasma Volume | Extraction Strategy | Solvents in Sample Preparation | IS | Column Specifications | Mobile Phase | Flow Rate (mL/min) | Run Time (min) | Linear Range (ng/mL) | Remarks |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ni et al. [28] | 0.2 | LLE | Ethyl acetate, methanol, and NaOH | Stable isotope labeled IS: Olanzapine‑d3 Fluoxetine‑d5 | Agilent Eclipse plus C18 column (4.6 × 100 mm, 5 μm) | Methanol-water containing 20mM ammonium formate (82.5:17.5, v/v) | 1.0 | 3.3 | OLA: 0.2 -25 ng/mL FLU: 0.2 -25 ng/mL | Relatively short run time (3.3 min); however, LLE-based extraction involves more complex sample preparation steps compared with simple protein precipitation. Relatively low fluoxetine recovery (59.96%) reported. Reliance on isotope-labeled IS increases analytical cost. Conventional 4.6 mm i.d., 5 µm column operated at 1.0 mL min⁻¹, resulting in higher mobile phase consumption per run (3.3 mL). |
| Fan et al. [29] | 0.2 | SPE | 1% Acetonitrile in water containing 10mM ammonium formate (v/v) | Diphenhydramine | Hypersil Gold C18 column (4.6×150 mm, 5 μm) | Gradient elution. Solvents A and B: acetonitrile and water containing10 mM ammonium formate and 0.01% formic acid, respectively. | 1.0 | 11.0 | OlA: 0.25–50 ng/mL FLU: 0.5–100 ng/mL | Automated online SPE improves cleanup efficiency; however, it requires specialized and costly valve-switching instrumentation. Longer analysis time (11.0 min). Conventional 4.6 mm i.d., 5 µm column operated at 1.0 mL min⁻¹, resulting in higher mobile phase consumption per run (11.0 mL). |
| Bonde et al. [30] | 0.2 | SPE | 5% ortho phosphoric acid in water | Stable isotope labeled IS: Olanzapine‑d3 Fluoxetine‑d5 | Gold C18 column (4.6×50 mm, 5 μm) | Methanol-water containing 2mM ammonium acetate (90:10, v/v) | 1.0 | 2.0 | OlA: 0.1–20 ng/mL FLU: 0.50–50 ng/mL | Short run time (2.0 min); however, SPE requires multi-step processing and costly cartridges. Reliance on isotope-labeled IS increases analytical cost. Conventional 4.6 mm i.d., 5 µm column operated at 0.5 mL min⁻¹, resulting in higher mobile phase consumption per run (2.0 mL). |
| Gopinath et al. [31] | 0.5 | SPE | 2mM ammonium acetate buffer (pH 8.5, adjusted with ammonia, methanol, solution), 5% methanol in water, 5% formic acid in acetonitrile | Duloxetine | Hypersil Gold C18 column (4.6×50 mm, 5 μm) | Acetonitrile–water containing 2% formic acid (70:30, v/v) | 0.5 | 4.0 | OLA: 0.12–25.03 ng/mL FLU: 1.00–150.20 ng/mL | Multi-step extraction with costly SPE cartridges. Longer analysis time. Conventional 4.6 mm i.d., 5 µm column operated at 0.5 mL min⁻¹, resulting in higher mobile phase consumption per run (2.0 mL). |
| Ravinder et al. [32] | 0.25 | LLE | Methyl tertiary butyl ether and n-hexane (80:20) | Duloxetine | X-terra RP8 column (4.6×50 mm, 5 μm) | Acetonitrile–water containing 30mM ammonium formate (90:10, v/v, pH 5.0) | 0.4 | 3.2 | OLA: 0.1–50 ng/mL FLU: 0.05–25.044 ng/mL | Relatively short run time (3.2 min); however, LLE-based extraction involves multi-step handling. Conventional 4.6 mm i.d., 5 µm column operated at 0.4 mL min⁻¹, resulting in higher mobile phase consumption per run (1.28 mL). |
| This work | 0.5 | Simple protein precipitation | Acetonitrile | Dextromethorphan | Hypersil Gold column (100 mm × 3 mm, 1.9 μm) column (3 × 100 mm, 1.9 μm) | 0.1% aqueous formic acid and acetonitrile (30:70, v/v) | 0.3 | 2.5 | OLA: 0.2–20.0 ng/mL FLU: 0.5–50.0 ng/mL for FLU | Simple and rapid protein precipitation sample preparation. Short run time (2.5 min). UHPLC 3.0 mm i.d., 1.9 µm column operated at 0.3 mL min⁻¹, resulting in reduced mobile phase consumption per run (0.75 mL), corresponding to a 41.4–93.2% reduction relative to the compared methods. Unified UHPLC–MS/MS platform for the simultaneous determination of two pharmacologically distinct binary mixtures under identical chromatographic conditions. |

Table S3: Extraction recovery results (n = 6)

| Analyte | QC Level | Recovery (%) | % CV* |
| --- | --- | --- | --- |
| Mixture 1 |  |  |  |
| DXM | LQC | 92.64 | 1.832 |
|  | MQC | 93.15 | 1.491 |
|  | HQC | 96.73 | 2.576 |
| PSE | LQC | 92.83 | 1.571 |
|  | MQC | 96.85 | 1.912 |
|  | HQC | 92.75 | 1.886 |
| IS | -- | 95.64 | 3.157 |
| Mixture 2 |  |  |  |
| OLA | LQC | 94.56 | 3.526 |
|  | MQC | 95.26 | 2.824 |
|  | HQC | 97.92 | 2.102 |
| FLU | LQC | 92.91 | 1.894 |
|  | MQC | 93.06 | 1.579 |
|  | HQC | 105.70 | 0.892 |
| IS | -- | 102.36 | 2.458 |

Table S4. System suitability parameters evaluated using six consecutive injections (n = 6) of MQC-level standard solution for all analytes.

| Peak Area Ratio %CV | Retention Time %CV | Analyte/ IS |
| --- | --- | --- |
| Mixture 1 |  |  |
| 0.629 | 0.104 | DXM |
| 1.236 | 0.354 | PSE |
| Mixture 2 |  |  |
| 1.861 | 0.004 | OLA |
| 0.382 | 0.335 | FLU |

Table S5. AI-assisted results generated by Gemini Pro for the greenness assessment of the proposed LC-MS/MS method using the Analytical GREEnness (AGREE) metric.

| AGREE Principle | GAC Principle Evaluated | Proposed Method Characteristic | Calculated Score | Color Scale |
| --- | --- | --- | --- | --- |
| 1 | Sample Treatment | External sample pre-treatment and batch analysis (reduced number of steps). | 0.30 | Orange |
| 2 | Sample Volume | 0.5 mL of human plasma per analysis. | 0.75 | Light Green |
| 3 | Device Positioning | Off-line analysis. | 0.00 | Deep Red |
| 4 | Sample Preparation Steps | ≤ 3 macroscopic steps (solvent addition, vortexing, and centrifugation). | 1.00 | Dark Green |
| 5 | Automation & Miniaturization | Semi-automated autosampler non-miniaturized standard LC-MS/MS system. | 0.25 | Orange |
| 6 | Derivatization | Direct MS detection; no derivatizing agents required. | 1.00 | Dark Green |
| 7 | Waste Generation | 2.25 mL of total analytical waste per sample. | 0.59 | Light Green |
| 8 | Analysis Throughput | 48 analytes determined per hour (24 runs/hour × 2 target analytes/run). | 0.89 | Dark Green |
| 9 | Energy Consumption | High energy consumption (LC-MS/MS system). | 0.00 | Deep Red |
| 10 | Source of Reagents | Aqueous mobile phase component utilized (some reagents from renewable sources). | 0.50 | Yellow |
| 11 | Toxicity | 1.53 mL of toxic/hazardous reagents per sample. | 0.45 | Yellow |
| 12 | Operator Safety | Operator exposed to 1 GHS hazard classes during prep (ACN: flammable) | 0.80 | Green |
|  |  | Overall Composite Score | 0.54 | Yellow |