| Supplemental Table 2. Analyte performance in serum, EDTA plasma (EDTA), and 100 plasma (P100) for 24 subjects. |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Re-liability | CV (%) | Comments |
| Simplex | ALB | Microalbumin | ug/mL | 2140 | Serum | 0 | 21 | 59554 | 5753 | 0.30 | 9.7 | Low reliability (<0.60); three pairs of samples not reported due to assay failure. |
|  |  |  |  |  | EDTA | 0 | 21 | 57990 | 4058 | 0.54 | 7.0 |  |
|  |  |  |  |  | P100 | 0 | 21 | 54238 | 4433 | 0.29 | 8.2 |  |
| HCVD4 | FABP3 | Fatty acid binding protein 3, muscle and heart (FABP, heart) | ng/mL | 15 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | MDA-LDL | Malondialdehyde-modified low-density lipoprotein | ng/mL | 245 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected consistently in any sample. |
|  |  |  |  |  | EDTA | 92 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 92 | nd | nd | nd | nd | nd |  |
|  | NPPB_PH | Natriuretic peptide (N-terminal prohormone) (NT proBNP) | pg/mL | 76 | Serum | 4 | 23 | 792 | 65.2 | 0.99 | 8.2 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 988 | 57.9 | >0.99 | 5.9 |  |
|  |  |  |  |  | P100 | 2 | 23 | 1031 | 56.8 | >0.99 | 5.5 |  |
|  | OLR1 | Oxidized low density lipoprotein (lectin-like) (LOX-1) | ng/mL | 0.9 | Serum | 42 | 13 | 1.64 | 0.24 | 0.91 | 14.4 | Only detected in serum, but not in all samples. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | THBD | Thrombomodulin (TM) | ng/mL | 0.24 | Serum | 0 | 24 | 4.28 | 0.14 | 0.98 | 3.4 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 4.57 | 0.28 | 0.90 | 6.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 4.34 | 0.23 | 0.91 | 5.4 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMP8 | A2M | Alpha-2-macroglobulin (A2Macro) | mg/mL | 0.017 | Serum | 0 | 24 | 1.76 | 0.11 | 0.93 | 6.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.55 | 0.09 | 0.83 | 6.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.10 | 0.06 | 0.86 | 5.6 |  |
|  | ADIPOQ | Adiponectin, C1Q and collagen domain containing | ug/mL | 0.06 | Serum | 0 | 24 | 5.26 | 1.08 | 0.92 | 20.5 | Performs better in EDTA plasma compared to serum. CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 4.93 | 0.41 | 0.99 | 8.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 4.63 | 0.67 | 0.96 | 14.5 |  |
|  | B2M | Beta-2-microglobulin | ug/mL | 0.01 | Serum | 0 | 24 | 1.55 | 0.21 | 0.81 | 13.4 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.37 | 0.13 | 0.86 | 9.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.26 | 0.17 | 0.74 | 13.3 |  |
|  | CCL5 | Chemokin (C-C motif) ligand 5(RANTES) | ng/mL | 0.21 | Serum | 0 | 24 | 20.34 | 2.71 | 0.91 | 13.3 | Performs better in EDTA plasma compared to serum. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 9.38 | 0.71 | 0.98 | 7.6 |  |
|  |  |  |  |  | P100 | 0 | 24 | 6.70 | 0.70 | 0.97 | 10.4 |  |
|  | FTL_FTH1 | Ferritin (dimer, light and heavy chain) (FT(L/H1)(FRTN) | ng/mL | 4.3 | Serum | 4 | 23 | 99.91 | 6.97 | 0.99 | 7.0 |  |
|  |  |  |  |  | EDTA | 4 | 23 | 95.11 | 8.60 | 0.99 | 9.0 |  |
|  |  |  |  |  | P100 | 4 | 23 | 90.17 | 5.50 | >0.99 | 6.1 |  |
|  | MB | Myoglobin | ng/mL | 8.1 | Serum | 0 | 24 | 51.65 | 5.99 | 0.95 | 11.6 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 48.15 | 4.66 | 0.97 | 9.7 |  |
|  |  |  |  |  | P100 | 0 | 24 | 45.13 | 3.30 | 0.98 | 7.3 |  |
|  | SERPINE1 | Plasminogen activator inhibitor 1(PAI-1) | ng/mL | 3.1 | Serum | 0 | 24 | 186.4 | 29.2 | 0.67 | 15.7 | Performs better in EDTA plasma compared to serum. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 46.94 | 7.04 | 0.90 | 15.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 39.71 | 6.97 | 0.88 | 17.6 |  |
|  | TIMP1 | Tissue inhibitor of metalloproteinase 1 | ng/mL | 3.2 | Serum | 0 | 24 | 136.9 | 17.84 | 0.76 | 13.0 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 71.71 | 7.96 | 0.85 | 11.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 68.24 | 8.28 | 0.84 | 12.1 |  |
|  | TNFRSF1B | Tumor necrosis factor receptor superfamily, member 1B (TNFR2) | ng/mL | 1.1 | Serum | 0 | 24 | 6.23 | 1.12 | 0.63 | 18.0 | Performs better in EDTA plasma compared to serum. Mean values higher in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 5.52 | 0.89 | 0.76 | 16.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 4.94 | 0.83 | 0.76 | 16.8 |  |
|  | VCAM1 | Vascular cell adhesion molecule-1 | ng/mL | 11 | Serum | 0 | 24 | 580.8 | 63.9 | 0.70 | 11.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 551.5 | 49.41 | 0.82 | 9.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 520.7 | 51.34 | 0.77 | 9.9 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC19 | AXL | AXL receptor tyrosine kinase | ng/mL | 0.05 | Serum | 0 | 24 | 9.20 | 0.74 | 0.96 | 8.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 10.16 | 0.68 | 0.97 | 6.7 |  |
|  |  |  |  |  | P100 | 0 | 24 | 10.00 | 0.75 | 0.96 | 7.5 |  |
|  | CCL16 | Chemokine (C-C motif) ligand 16 (hCC-4) | ng/mL | 0.047 | Serum | 0 | 24 | 4.22 | 0.42 | 0.92 | 10.0 | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 4.48 | 0.27 | 0.97 | 6.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 4.48 | 0.28 | 0.97 | 6.3 |  |
|  | CCL18 | Chemokine (C-C motif) ligand 18 (PARC) | ng/mL | 6.2 | Serum | 0 | 24 | 89.44 | 9.71 | 0.96 | 10.9 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 93.90 | 9.41 | 0.97 | 10.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 96.15 | 8.02 | 0.98 | 8.3 |  |
|  | FAS | Fas cell surface death receptor | ng/mL | 5.8 | Serum | 0 | 24 | 19.71 | 1.68 | 0.94 | 8.5 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 4 | 22 | 15.00 | 2.73 | 0.79 | 18.2 |  |
|  |  |  |  |  | P100 | 2 | 23 | 15.14 | 2.48 | 0.80 | 16.4 |  |
|  | HGF | Hepatocyte growth factor (hepapoietin A; scatter factor) | ng/mL | 1 | Serum | 0 | 24 | 13.01 | 1.04 | 0.97 | 8.0 | Performs better in serum compared to EDTA plasma. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 3.58 | 0.52 | 0.86 | 14.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 3.50 | 0.58 | 0.81 | 16.4 |  |
|  | TNFRSF10C | Tumor necrosis factor receptor superfamily, member 10c, decoy without an intracellular domain (TRAIL-R3) | ng/mL | 0.96 | Serum | 0 | 24 | 16.82 | 0.98 | 0.97 | 5.8 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 15.09 | 0.98 | 0.96 | 6.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 13.13 | 0.79 | 0.97 | 6.0 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC35 | AGER | Advanced glycosylation end product-specific receptor (RAGE) | ng/mL | 0.35 | Serum | 0 | 24 | 3.32 | 0.30 | 0.97 | 9.0 | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 3.03 | 0.16 | 0.99 | 5.2 |  |
|  |  |  |  |  | P100 | 0 | 24 | 2.99 | 0.22 | 0.98 | 7.4 |  |
|  | CHGA | Chromogranin-A (CgA) | ng/mL | 13 | Serum | 0 | 24 | 648.27 | 50.14 | 0.99 | 7.7 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 554.06 | 36.28 | >0.99 | 6.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 607.94 | 43.79 | 0.99 | 7.2 |  |
|  | NGF | Nerve growth factor (beta polypeptide)(NGFB, NGF-beta) | ng/mL | 0.078 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | NRCAM | Neuronal cell adhesion molecule | ng/mL | 0.2 | Serum | 6 | 22 | 0.83 | 0.10 | 0.94 | 11.8 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.21 | 0.11 | 0.98 | 8.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.09 | 0.11 | 0.97 | 9.9 |  |
|  | S100B | S100 calcium-binding protein B | ng/mL | 0.5 | Serum | 92 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | SOD1 | Superoxide dismutase 1, soluble | ng/mL | 0.12 | Serum | 0 | 24 | 15.21 | 1.48 | 0.96 | 9.7 | Mean values higher in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 34.90 | 3.11 | 0.96 | 8.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 31.33 | 2.31 | 0.97 | 7.4 |  |
|  | SORT1 | Sortilin 1 | ng/mL | 0.22 | Serum | 0 | 24 | 6.99 | 0.43 | 0.90 | 6.2 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 4.43 | 0.27 | 0.92 | 6.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 5.54 | 0.39 | 0.91 | 7.0 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC42 | CXCL10 | Chemokine (C-X-C motif) ligand 10 (IP-10) | pg/mL | 127 | Serum | 0 | 24 | 359.40 | 15.73 | 0.99 | 4.4 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 360.17 | 21.34 | 0.97 | 5.9 |  |
|  |  |  |  |  | P100 | 4 | 23 | 313.61 | 15.08 | 0.98 | 4.8 |  |
|  | CCL8 | Chemokine (C-C motif) ligand 8 (MCP-2) | pg/mL | 6.7 | Serum | 0 | 24 | 36.31 | 3.00 | 0.95 | 8.3 |  |
|  |  |  |  |  | EDTA | 2 | 23 | 29.72 | 2.10 | 0.97 | 7.1 |  |
|  |  |  |  |  | P100 | 2 | 23 | 27.55 | 1.90 | 0.97 | 6.9 |  |
|  | CCL13 | Chemokine (C-C motif) ligand 13 (MCP-4) | pg/mL | 445 | Serum | 0 | 24 | 3460.8 | 331.3 | 0.82 | 9.6 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 1829.4 | 171.4 | 0.81 | 9.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1665.3 | 148.4 | 0.89 | 8.9 |  |
|  | CXCL9 | Chemokine (C-X-C motif) ligand 9 (MIG) | pg/mL | 110 | Serum | 0 | 24 | 1548.3 | 154.1 | 0.98 | 10.0 | CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 1463.9 | 110.7 | 0.99 | 7.6 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1376.0 | 161.3 | 0.97 | 11.7 |  |
|  | CCL20 | Chemokine (C-C motif) ligand 20 (MIP-3a) | pg/mL | 30 | Serum | 15 | 18 | 49.53 | 4.96 | 0.96 | 10.0 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 8 | 20 | 52.33 | 7.93 | 0.96 | 15.2 |  |
|  |  |  |  |  | P100 | 19 | 18 | 45.69 | 5.39 | 0.98 | 11.8 |  |
|  | CCL23 | Chemokine (C-C motif) ligand 23 (MPIF-1) | ng/mL | 0.14 | Serum | 0 | 24 | 1.57 | 0.06 | 0.99 | 4.1 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.37 | 0.06 | 0.98 | 4.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.23 | 0.08 | 0.97 | 6.6 |  |
|  | IL6R | Interleukin 6 receptor | ng/mL | 0.027 | Serum | 0 | 24 | 24.83 | 2.25 | 0.91 | 9.1 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 25.24 | 2.01 | 0.91 | 8.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 24.36 | 1.61 | 0.94 | 6.6 |  |
|  | SELE | Selectin E | ng/mL | 0.21 | Serum | 0 | 24 | 12.13 | 0.70 | 0.98 | 5.7 | Does not meet Myriad-RBM performance specification for EDTA plasma |
|  |  |  |  |  | EDTA | 0 | 24 | 8.44 | 0.51 | 0.97 | 6.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 8.31 | 0.45 | 0.98 | 5.4 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC49 | APOA4 | Apolipoprotein A-IV (Apo A-IV) | ug/mL | 0.8 | Serum | 0 | 24 | 13.10 | 1.47 | 0.94 | 11.2 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 13.08 | 1.25 | 0.92 | 9.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 8.93 | 1.12 | 0.78 | 12.6 |  |
|  | HSPD1 | Heat shock 60kDa protein 1 (chaperonin)(HSP-60) | ng/mL | 37 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | INS-intact | Proinsulin, intact | pM | 3.6 | Serum | 94 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 79 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 88 | nd | nd | nd | nd | nd |  |
|  | INS-total | Proinsulin, total | pM | 34 | Serum | 100 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 92 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 94 | nd | nd | nd | nd | nd |  |
|  | SERPINA3 | Serpin peptidase inhibitor, clade A (alpha-1-antiproteinase, antitrypsin), member 3)(AACT) | ug/mL | 6.9 | Serum | 0 | 24 | 892.73 | 176.8 | 0.79 | 19.8 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 711.46 | 120.1 | 0.61 | 16.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 765.23 | 116.1 | 0.66 | 15.2 |  |
| HMPC62 | CCL24 | Chemokine (C-C motif) ligand 24 | pg/mL | 38 | Serum | 0 | 24 | 1165.7 | 61.79 | 0.99 | 5.3 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 623.27 | 47.13 | 0.99 | 7.6 |  |
|  |  |  |  |  | P100 | 0 | 24 | 477.85 | 31.54 | 0.99 | 6.6 |  |
|  | IL2RA | Interleukin 2 receptor, alpha (IL-2 receptor alpha) | pg/mL | 92 | Serum | 0 | 24 | 2437.5 | 256.7 | 0.83 | 10.5 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 2121.8 | 186.4 | 0.93 | 8.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 2043.3 | 127.4 | 0.96 | 6.2 |  |
|  | MICA | MHC Class I polypeptide-related sequence A | pg/mL | 73 | Serum | 77 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 63 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 56 | nd | nd | nd | nd | nd |  |
|  | TGFB1_LAP | Transforming growth factor, beta 1 (TGFB1; latency associated peptide)(TGFB1) | ng/mL | 0.077 | Serum | 0 | 24 | 10.09 | 0.85 | 0.88 | 8.4 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 4.75 | 0.52 | 0.96 | 10.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 3.65 | 0.38 | 0.95 | 10.4 |  |
|  | TNFRSF11B | Tumor necrosis factor receptor superfamily, member 11b (OPG) | pM | 0.85 | Serum | 0 | 24 | 6.43 | 0.60 | 0.90 | 9.3 | Performs better in serum compared to EDTA plasma. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 3.72 | 0.51 | 0.80 | 13.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 3.36 | 0.36 | 0.86 | 10.7 |  |
|  | TNFRSF1A | Tumor necrosis factor receptor superfamily, member 1A (TNFR1) | pg/mL | 38 | Serum | 0 | 24 | 1682.5 | 83.44 | 0.98 | 5.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1662.3 | 71.8 | 0.98 | 4.3 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1635.8 | 78.84 | 0.98 | 4.8 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC83 | ANGPT1 | Angiopoietin 1 (ANG1) | ng/mL | 2.1 | Serum | 0 | 24 | 36.00 | 2.97 | 0.92 | 8.3 | Performs better in serum compared to EDTA plasma. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 5.79 | 0.84 | 0.88 | 14.5 |  |
|  |  |  |  |  | P100 | 2 | 23 | 5.79 | 0.93 | 0.77 | 16.1 |  |
|  | CA9 | Carbonic anhydrase IX | ng/mL | 0.22 | Serum | 8 | 20 | 0.91 | 0.23 | 0.57 | 25.6 | Performs better in serum compared to EDTA plasma; however, reliability <60 and CV ≥ 20%. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 42 | nd | 0.52 | 0.22 | nd | nd |  |
|  |  |  |  |  | P100 | 52 | nd | 0.43 | 0.13 | nd | nd |  |
|  | CDH13 | Cadherin 13, H-cadherin (heart) (T-cad) | ng/mL | 2.2 | Serum | 0 | 24 | 4.26 | 0.32 | 0.86 | 7.5 | Mean values higher in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 17.83 | 1.42 | 0.91 | 8.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 16.94 | 1.82 | 0.87 | 10.7 |  |
|  | CEACAM1 | Carcinoembryonic antigen-related cell adhesion molecule 1 | ng/mL | 3.1 | Serum | 4 | 23 | 10.81 | 1.71 | 0.72 | 15.8 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 9.73 | 1.87 | 0.49 | 19.2 |  |
|  |  |  |  |  | P100 | 0 | 24 | 9.00 | 1.93 | 0.45 | 21.5 |  |
|  | DCN | Decorin | ng/mL | 0.13 | Serum | 0 | 24 | 1.38 | 0.14 | 0.66 | 10.4 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.86 | 0.15 | 0.76 | 8.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.72 | 0.11 | 0.84 | 6.1 |  |
|  | IL18BP | Interleukin 18 binding protein | ng/mL | 0.096 | Serum | 0 | 24 | 11.73 | 0.39 | 0.99 | 3.3 | CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 11.42 | 0.45 | 0.98 | 4.0 |  |
|  |  |  |  |  | P100 | 0 | 24 | 9.75 | 0.71 | 0.95 | 7.3 |  |
|  | MDK | Midkine (neurite growth-promoting factor 2) | ng/mL | 0.46 | Serum | 0 | 24 | 1.89 | 0.14 | 0.97 | 7.4 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 2.00 | 0.15 | 0.96 | 7.3 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.67 | 0.17 | 0.95 | 10.1 |  |
|  | PECAM1 | Platelet/endothelial cell adhesion molecule 1 | ng/mL | 11 | Serum | 0 | 24 | 46.54 | 3.23 | 0.89 | 6.9 | CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 45.96 | 2.54 | 0.93 | 5.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 39.85 | 3.53 | 0.85 | 8.8 |  |
|  | SFTPD | Surfactant protein D (SP-D) | ng/mL | 0.19 | Serum | 0 | 24 | 9.78 | 0.63 | >0.99 | 6.5 | CV lower in EDTA plasma compared to P100 plasma. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 5.98 | 0.28 | >0.99 | 4.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 5.81 | 0.91 | 0.96 | 15.7 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMPC84 | CSTB | Cystatin B | ng/mL | 0.4 | Serum | 0 | 24 | 8.33 | 1.10 | 0.90 | 13.2 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 9.75 | 0.89 | 0.95 | 9.2 |  |
|  |  |  |  |  | P100 | 0 | 24 | 8.94 | 0.93 | 0.94 | 10.4 |  |
|  | CDH1 | Cadherin, type 1, E-cadherin (epithelial)(ECAD) | ng/mL | 6 | Serum | 0 | 24 | 3428.9 | 511.5 | 0.81 | 14.9 | Does not meet Myriad-RBM performance specification for serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 2700.4 | 273.6 | 0.91 | 10.1 |  |
|  |  |  |  |  | P100 | 0 | 24 | 2672.9 | 238.1 | 0.92 | 8.9 |  |
|  | LTF | Lactotransferrin | ng/mL | 1.8 | Serum | 0 | 24 | 29.12 | 2.73 | 0.98 | 9.4 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 11.75 | 1.39 | 0.86 | 11.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 12.26 | 1.67 | 0.86 | 13.6 |  |
|  | KIT | v-kit Hardy-Zuckerman 4 feline sarcoma viral (SCFR) | ng/mL | 0.45 | Serum | 0 | 24 | 8.36 | 1.18 | 0.75 | 14.1 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 8.11 | 1.05 | 0.78 | 12.9 |  |
|  |  |  |  |  | P100 | 0 | 24 | 7.99 | 1.14 | 0.79 | 14.2 |  |
|  | SLPI | Secretory leukocyte peptidase inhibitor (ALP) | ng/mL | 0.63 | Serum | 0 | 24 | 33.00 | 1.86 | 0.90 | 5.6 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 35.25 | 2.04 | 0.93 | 5.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 32.60 | 1.73 | 0.97 | 5.3 |  |
|  | SPINK1 | Serine peptidase inhibitor, Kazal type 1 (TATI) | ng/mL | 0.27 | Serum | 0 | 24 | 15.62 | 1.87 | 0.99 | 12.0 | CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 15.32 | 1.72 | 0.99 | 11.2 |  |
|  |  |  |  |  | P100 | 0 | 24 | 15.73 | 2.76 | 0.98 | 17.5 |  |
|  | TIMP2 | Tissue inhibitor of metalloproteinase 2 | ng/mL | 2.7 | Serum | 0 | 24 | 69.4 | 3.15 | 0.92 | 4.5 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 61.21 | 3.38 | 0.85 | 5.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 60.48 | 3.14 | 0.86 | 5.2 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMP-CORE1 | CCL2 | Chemokine (C-C motif) ligand 2 (MCP-1) | pg/mL | 23 | Serum | 0 | 24 | 387.94 | 25.85 | 0.95 | 6.7 | Performs better in serum compared to EDTA plasma. Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 166.58 | 18.90 | 0.81 | 11.3 |  |
|  |  |  |  |  | P100 | 0 | 24 | 164.02 | 13.46 | 0.91 | 8.2 |  |
|  | CCL3 | Chemokine (C-C motif) ligand 3 (MIP1A) | pg/mL | 35 | Serum | 73 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 60 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 81 | nd | nd | nd | nd | nd |  |
|  | CCL4 | Chemokine (C-C motif) ligand 4 (MIP1B) | pg/mL | 29 | Serum | 0 | 24 | 306.6 | 22.09 | 0.98 | 7.2 | CV lower in P100 plasma compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 247.48 | 22.77 | 0.97 | 9.2 |  |
|  |  |  |  |  | P100 | 0 | 24 | 237.31 | 13.15 | 0.99 | 5.5 |  |
|  | CSF2 | Granulocyte-macrophage colony-stimulating factor (GMCSF) | pg/mL | 49 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IFNG | Interferon, gamma (IFN-gamma) | pg/mL | 1.7 | Serum | 98 | 0 | nd | nd | nd | nd | Only detected in plasma; however, reliability <60 and CV ≥ 20%. |
|  |  |  |  |  | EDTA | 0 | 24 | 8.26 | 2.30 | 0.32 | 27.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 7.76 | 2.53 | 0.28 | 32.6 |  |
|  | IL10 | Interleukin 10 | pg/mL | 4.7 | Serum | 75 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 69 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 85 | nd | nd | nd | nd | nd |  |
|  | 1L18 | Interleukin 18 (interferon-gamma-inducing-factor) | pg/mL | 15 | Serum | 0 | 24 | 261.38 | 17.90 | 0.99 | 6.8 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 269.10 | 31.14 | 0.98 | 11.6 |  |
|  |  |  |  |  | P100 | 0 | 24 | 248.71 | 21.78 | 0.99 | 8.8 |  |
|  | IL2 | Interleukin 2 | pg/mL | 4.5 | Serum | 100 | 0 | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 67 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 81 | nd | nd | nd | nd | nd |  |
|  | IL3 | Interleukin 3 | ng/mL | 0.0069 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IL4 | Interleukin 4 | pg/mL | 18 | Serum | 100 | 0 | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 81 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 88 | nd | nd | nd | nd | nd |  |
|  | IL5 | Interleukin 5 (colony stimulating factor, eosinophil) | pg/mL | 16 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |

| HMP-CORE1 (cont.) | IL6 | Interleukin 6 (interferon, beta 2) | pg/mL | 6.3 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IL7 | Interleukin 7 | pg/mL | 12 | Serum | 100 | 0 | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 98 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IL8 | Interleukin 8 | pg/mL | 3.1 | Serum | 0 | 24 | 11.8 | 1.48 | 0.81 | 12.5 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 8.70 | 1.14 | 0.85 | 13.1 |  |
|  |  |  |  |  | P100 | 2 | 23 | 8.41 | 1.16 | 0.73 | 13.7 |  |
|  | LTA | Lymphotoxin alpha (tumor necrosis factor beta) (TNFB, TNF-beta) | pg/mL | 9.1 | Serum | 100 | 0 | nd | nd | nd | nd | Not detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | MMP2 | Matrix metalloproteinase 2 (gelatinase A, 72kDa gelatinase, 72kDa type IV collagenase) | ng/mL | 12 | Serum | 100 | 0 | nd | nd | nd | nd | Only detected in plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 1743.8 | 118.9 | 0.94 | 6.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1676.9 | 112.8 | 0.95 | 6.7 |  |
|  | TNF | Tumor necrosis factor (TNF-alpha) | pg/mL | 12 | Serum | 88 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 96 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 92 | nd | nd | nd | nd | nd |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMP-CORE2 | BDNF | Brain-derived neurotrophic factor | ng/mL | 0.045 | Serum | 0 | 24 | 23.83 | 2.48 | 0.90 | 10.4 | Mean values higher in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 1.89 | 0.22 | 0.97 | 11.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 2.29 | 0.18 | 0.99 | 7.7 |  |
|  | CCL11 | Chemokine (C-C motif) ligand 11 (eotaxin-1) | pg/mL | 51 | Serum | 4 | 23 | 201.17 | 17.79 | 0.97 | 8.8 | Mean values higher in serum compared to EDTA plasma |
|  |  |  |  |  | EDTA | 31 | 14 | 107.39 | 13.66 | 0.88 | 12.7 |  |
|  |  |  |  |  | P100 | 31 | 15 | 98.70 | 13.03 | 0.87 | 13.2 |  |
|  | F7 | Coagulation factor VII | ng/mL | 2.8 | Serum | 0 | 24 | 552.85 | 38.88 | 0.84 | 7.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 602.44 | 38.92 | 0.85 | 6.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 570.08 | 34.30 | 0.86 | 6.0 |  |
|  | ICAM1 | Intercellular adhesion molecule 1 | ng/mL | 1.4 | Serum | 0 | 24 | 161.69 | 11.44 | 0.95 | 7.1 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 162.38 | 12.70 | 0.93 | 7.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 152.85 | 11.66 | 0.94 | 7.6 |  |
|  | IL12B | Interleukin 12B (natural killer cell stimulatory factor 2, cytotoxic lymphocyte maturation factor 2, p40)(IL-12p40) | ng/mL | 0.34 | Serum | 42 | 12 | 0.42 | 0.04 | 0.00 | 9.7 | Only detected in serum, but not consistently and reliability <0.60. |
|  |  |  |  |  | EDTA | 94 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IL12A/IL12B | Interleukin12 subunit p70 (IL12A/IL12B heterodimer)(IL-12p70) | pg/mL | 43 | Serum | 100 | 0 | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 94 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 94 | nd | nd | nd | nd | nd |  |
|  | IL15 | Interleukin 15 | ng/mL | 0.63 | Serum | 54 | nd | nd | nd | nd | nd | Best detected in P100, but not consistently. |
|  |  |  |  |  | EDTA | 75 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 44 | 12 | 1.54 | 0.24 | 0.99 | 15.3 |  |
|  | IL17A | Interleukin 17 | pg/mL | 4.8 | Serum | 96 | nd | nd | nd | nd | nd | Not detected consistently in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 98 | 0 | nd | nd | nd | nd |  |
|  | IL1A | Interleukin 1, alpha (IL-1 alpha) | ng/mL | 0.0053 | Serum | 33 | 16 | 0.01 | 0.00 | 0.73 | 13.5 | Best detected in serum, but not consistently. |
|  |  |  |  |  | EDTA | 69 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 71 | nd | nd | nd | nd | nd |  |
|  | IL1B | Interleukin 1, beta (IL-1 beta) | pg/mL | 2.7 | Serum | 96 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
|  |  |  |  |  | EDTA | 100 | 0 | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 100 | 0 | nd | nd | nd | nd |  |
|  | IL1RN | Interleukin1 receptor antagonist (IL1RA) | pg/mL | 127 | Serum | 33 | 13 | nd | nd | 0.33 | 20.7 | Only detected in serum; however, reliability <0.6 and CV >20%. |
|  |  |  |  |  | EDTA | 83 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 77 | nd | nd | nd | nd | nd |  |

| HMP-CORE2 (cont.) | IL23B | Interleukin 23, alpha subunit p19 (IL-23) | ng/mL | 0.8 | Serum | 71 | nd | nd | nd | nd | nd | Not consistently detected in any sample. |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | EDTA | 79 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | P100 | 92 | nd | nd | nd | nd | nd |  |
|  | KITLG | KIT ligand (stem cell factor)(SCF) | pg/mL | 105 | Serum | 2 | 23 | 276.22 | 38.66 | 0.74 | 14.0 | Performs better in serum compared to EDTA plasma. |
|  |  |  |  |  | EDTA | 6 | 22 | 227.18 | 35.86 | 0.56 | 15.8 |  |
|  |  |  |  |  | P100 | 2 | 23 | 215.46 | 29.47 | 0.65 | 13.7 |  |
|  | MMP3 | Matrix metalloproteinase 3 (stromelysin 1, progelatinase) | ng/mL | 0.046 | Serum | 0 | 24 | 16.53 | 0.91 | 0.99 | 5.5 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 11.21 | 0.70 | 0.99 | 6.3 |  |
|  |  |  |  |  | P100 | 0 | 24 | 10.83 | 0.77 | 0.98 | 7.1 |  |
|  | MMP9 | Matrix metalloproteinase 9 (gelatinase B, 92kDA type IV collagenase) | ng/mL | 16 | Serum | 48 | nd | nd | nd | nd | nd | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 130.5 | 7.60 | 0.98 | 5.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 179.38 | 14.95 | 0.94 | 8.3 |  |
|  | VEGFA | Vascular endothelial growth factor | pg/mL | 47 | Serum | 0 | 24 | 1303.8 | 174.2 | 0.86 | 13.4 | Performs better in EDTA plasma compared to serum. Mean values higher in serum compared to EDTA plasma. CV lower in EDTA plasma compared to P100 plasma. |
|  |  |  |  |  | EDTA | 0 | 24 | 717.33 | 54.06 | 0.76 | 7.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 695.38 | 78.62 | 0.55 | 11.3 |  |

| Plex | Analyte Abbreviation* | Analyte Description (alternative commonly-used abbreviations) | Units | LLOQ | Sample | % <LLOQ | N of Pairs | Mean | SD | Reliability | CV (%) | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| HMP-CORE4 | C3 | Complement component 3 | mg/mL | 0.023 | Serum | 0 | 24 | 0.91 | 0.15 | 0.62 | 16.0 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 0.91 | 0.10 | 0.77 | 11.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 0.83 | 0.09 | 0.62 | 10.6 |  |
|  | CRP | C-reactive protein | ug/mL | 1.4 | Serum | 25 | 18 | 5.27 | 1.54 | 0.91 | 29.1 | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 23 | 18 | 5.71 | 1.00 | 0.97 | 17.5 |  |
|  |  |  |  |  | P100 | 29 | 16 | 5.95 | 1.00 | 0.97 | 16.9 |  |
|  | FGA_FGB_FGG | Fibrinogen (trimer; alpha chain, beta chain, gamma chain)(FG(A/ B/ G)) | mg/mL | 0.06 | Serum | 92 | nd | nd | nd | nd | nd |  |
|  |  |  |  |  | EDTA | 0 | 24 | 3.41 | 0.39 | 0.77 | 11.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 3.02 | 0.30 | 0.75 | 10.0 |  |
|  | HP | Haptoglobin | mg/mL | 0.035 | Serum | 0 | 24 | 1.64 | 0.33 | 0.88 | 20.1 | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 2 | 23 | 1.48 | 0.17 | 0.95 | 11.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.25 | 0.11 | 0.98 | 8.9 |  |
|  | IgA | Immunoglobulin A | mg/mL | 0.035 | Serum | 0 | 24 | 1.80 | 0.41 | 0.88 | 22.5 | . |
|  |  |  |  |  | EDTA | 0 | 24 | 1.84 | 0.32 | 0.93 | 17.5 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.68 | 0.21 | 0.96 | 12.5 |  |
|  | IgM | Immunoglobulin M | mg/mL | 0.19 | Serum | 6 | 22 | 1.89 | 0.34 | 0.94 | 17.8 | Performs better in EDTA plasma compared to serum. CV lower in EDTA plasma compared to P100 plasma |
|  |  |  |  |  | EDTA | 6 | 22 | 1.93 | 0.21 | 0.98 | 10.7 |  |
|  |  |  |  |  | P100 | 8 | 22 | 1.71 | 0.28 | 0.95 | 16.4 |  |
|  | GC | Group-specific component (vitamin D-binding protein)(VDBP) | ug/mL | 13 | Serum | 0 | 24 | 196.06 | 27.10 | 0.88 | 13.8 | Performs better in EDTA plasma compared to serum. |
|  |  |  |  |  | EDTA | 0 | 24 | 188.02 | 10.95 | 0.97 | 5.8 |  |
|  |  |  |  |  | P100 | 0 | 24 | 174.90 | 11.04 | 0.97 | 6.3 |  |
|  | SERPINA1 | Serpin peptidase inhibitor, clade A (alpha-1-antiproteinase, antitrypsin)(AAT) | mg/mL | 0.014 | Serum | 0 | 24 | 1.31 | 0.19 | 0.68 | 14.4 |  |
|  |  |  |  |  | EDTA | 0 | 24 | 1.31 | 0.18 | 0.67 | 13.4 |  |
|  |  |  |  |  | P100 | 0 | 24 | 1.19 | 0.14 | 0.67 | 11.6 |  |

*The Analyte Abbreviation corresponds to the terminology of NCBI when the analyte is a protein product of a specific gene (ncbi.nlm.hig.gov/gene).  In some instances, alternative, commonly used abbreviations are provided in non-bolded font.  If the analyte is not directly a gene product, then the designation for the analyte as used by Myriad-RBM is provided.  In some instances, especially with the chemokines/cytokines, the official name is presented along with one other commonly used name (found in parenthesis).

Fibrinogen is a three subunit complex consisting of three protein FGA, FGB, FGG;    Interleukin12 subunit p70 is a dimer consisting of IL12B/IL12A;  Ferritin is a dimer consisting of FTL and FTH1.