Supplementary Table 1. Chromatographic data of the measured TRP metabolites. Operating software was MassLynx V4.2 SCN977. 11 calibrators were prepared to this study. Narrower linear range was obtained in the case of 5-HIAA (first 6 solution).

| Metabolite | Retention time (min) | Linear range (ng/mL) | Calibration curve | R2 |
| --- | --- | --- | --- | --- |
| 3-HK | 1.40 | 0.625 – 20.0 | 2.53847 x - 0.867225 | 0.9981 |
| KYN | 2.23 | 46.875 – 1 500.0 | 1.81225 x + 4.48078 | 0.9999 |
| 5-HT | 2.28 | 9.375 – 300.0 | 2.43416 x + 5.72209 | 0.9998 |
| 3-HANA | 3.01 | 0.625 – 20.0 | 0.581564 x + 0.00773113 | 0.9990 |
| TRP | 3.39 | 937.5 – 30 000.0 | 17.0744 x + 2285.12 | 0.9998 |
| 5-HIAA | 3.75 | 5.625 – 30.0 | 6.39825 x - 25.2444 | 0.9991 |
| ANA | 3.94 | 0.625 – 20.0 | 1.04401 x + 1.547 | 0.9994 |
| KYNA | 3.97 | 0.938 – 30.0 | 2.33303 x - 0.778719 | 0.9996 |
| QUIN | 4.51 | 4.688 – 180.0 | 0.940585 x + 21.4087 | 0.9994 |
| PICA | 4.69 | 0.625 –20.0 | 1.37843 x + 1.14019 | 0.9977 |
| XA | 4.87 | 0.625 – 20.0 | 2.64667 x + 0.0711093 | 0.9995 |

Chromatographic data of the deuterated internal standards.

| IS | Retention time (min) | Conc. (ng/mL) | SD | IS | Retention time (min) | Conc. (ng/mL) | SD (%) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| d3-3-HK | 1.39 | 20.0 | 4.142 | d5-KYNA | 3.96 | 20.0 | 2.685 |
| d4-KYN | 2.17 | 800.0 | 1.500 | d3-QUIN | 4.51 | 150.0 | 3.129 |
| d4-5-HT | 2.26 | 300.0 | 1.662 | d4-PICA | 4.68 | 20.0 | 14.101 |
| d3-3-HANA | 2.98 | 20.0 | 4.990 | d4-XA | 4.87 | 20.0 | 2.817 |
| d5-TRP | 3.36 | 4 000.0 | 6.055 |  |  |  |  |
| d5-5-HIAA | 3.73 | 40.0 | 5.107 |  |  |  |  |

QC samples – runtime 30.5 hrs (mean concentrations, 15-15 replicates of each, the given n is the samples size of the pooled individual samples).

| Metabolite | Control group (n = 9) |  | Cluster headache group (n = 10) |  |
| --- | --- | --- | --- | --- |
|  | Concentration (ng/mL) | SD value | Concentration (ng/mL) | SD value |
| 3-HK | 8.254 | 0.487 | 9.960 | 0,526 |
| KYN | 353.420 | 2.826 | 387.036 | 3.555 |
| 5-HT | 4.087 | 0.908 | 36.497 | 0.815 |
| 3-HANA | 7.241 | 0.400 | 9.504 | 0.544 |
| TRP | 12 793.123 | 527.794 | 12 558.435 | 544.574 |
| 5-HIAA | 8.376 | 0.917 | 8.991 | 0.755 |
| ANA | 5.627 | 0.470 | 4.209 | 0.277 |
| KYNA | 6.615 | 0.170 | 6.709 | 0.160 |
| QUIN | 52.144 | 2.239 | 54.888 | 2.640 |
| PICA | 4.319 | 0.736 | 3.371 | 0.857 |
| XA | 2.865 | 0.158 | 2.750 | 0.144 |

PICA, QUIN, and XA were analysed by their mono butylated forms.

According to the absence of relative internal standard for ANA, d5-KYA was applied to the quantification process in each sample, as their retention times are comparable.

Timing of the QC samples

5 replicates of control QC  5 replicates of cluster QC  duplicates of the standards  1-1 QC of both groups  (duplicates of 7 or 8 randomized samples  1-1 QC) x 4  duplicates of vortexed, same standards  1-1 QC of both groups  (duplicates of 7 or 8 randomized samples  1-1 vortexed, same QC) x 4

QC samples – range of measured concentration levels.

| Compound | Control group (n = 9) |  | Cluster headache group (n = 10) |  |
| --- | --- | --- | --- | --- |
|  | Min. value (ng/mL) | Max. value (ng/mL) | Min. value (ng/mL) | Max. value (ng/mL) |
| 3-HK | 7.554 | 9.078 | 9.044 | 10.912 |
| KYN | 350.225 | 359.489 | 382.282 | 396.047 |
| 5-HT | 2.306 | 5.766 | 34.890 | 37.935 |
| 3-HANA | 6.599 | 7.858 | 8.055 | 10.238 |
| TRP | 11 849.924 | 13 663.727 | 11 734.006 | 13 261.491 |
| 5-HIAA | 6.942 | 9.755 | 7.778 | 10.342 |
| ANA | 4.809 | 6.622 | 3.698 | 4.624 |
| KYNA | 6.336 | 6.879 | 6.450 | 6.951 |
| QUIN | 48.922 | 56.465 | 50.904 | 60.022 |
| PICA | 3.048 | 5.687 | 2.146 | 4.910 |
| XA | 2.465 | 3.061 | 2.564 | 2.988 |