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Organism Type
	
N
	Autof MS1000
	Bruker Biotyper
	Statistical results

	
	
	Agreement %
	No identification %
	Discrepancy %
	Agreement %
	No identification %
	Discrepancy %
	X2
	P

	Gram-negative Enterobacteriaceae and Non-Enterobacteriaceae
	69
	94.2
	0
	2.9
	94.2
	0
	2.9
	0
	1

	Gram-negative fastidious
	30
	100
	0
	0
	100
	0
	0
	-
	1

	Gram-positive aerobic
	66
	97.0
	0
	3.0
	98.5
	0
	1.5
	0
	1

	Anaerobic
	30
	100
	0
	0
	100
	0
	0
	0
	1

	Yeast and yeast-like
	34
	100
	0
	0
	97.1
	0
	2.9
	-
	1

	Total
	229
	97.4
	0
	2.6
	97.4
	0
	2.6
	0
	1


-：Fisher test without X2 value.
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