Additional file 3. Verification of microbial identification on the Autof MS1000 and Bruker Biotyper

| Organism Type | N | Autof MS1000 |  |  | Bruker Biotyper |  |  | Statistical results |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Agreement % | No identification % | Discrepancy % | Agreement % | No identification % | Discrepancy % | X2 | P |
| Gram-negative Enterobacteriaceae and Non-Enterobacteriaceae | 69 | 94.2 | 0 | 2.9 | 94.2 | 0 | 2.9 | 0 | 1 |
| Gram-negative fastidious | 30 | 100 | 0 | 0 | 100 | 0 | 0 | - | 1 |
| Gram-positive aerobic | 66 | 97.0 | 0 | 3.0 | 98.5 | 0 | 1.5 | 0 | 1 |
| Anaerobic | 30 | 100 | 0 | 0 | 100 | 0 | 0 | 0 | 1 |
| Yeast and yeast-like | 34 | 100 | 0 | 0 | 97.1 | 0 | 2.9 | - | 1 |
| Total | 229 | 97.4 | 0 | 2.6 | 97.4 | 0 | 2.6 | 0 | 1 |

-：Fisher test without X2 value.