Table S1. Comparison of identification accuracy for Aeromonas species by different methods
	Species (WGS)
	No. of isolates
	MALDI‑TOF MS correct identification (%)
	Major misidentification by MALDI‑TOF MS
	VITEK 2 COMPACT system correct identification (%)
	Major misidentification by VITEK 2 COMPACT system

	A. caviae
	187
	180(96.26)
	A. hydrophila (7)
	0 (0)
	A. hydrophila/caviae (186), A. veronii (1)

	A. veronii
	11
	10(90.91)
	A. hydrophila (1)
	8 (72.73)
	A. hydrophila/caviae (3)

	A. trota
	7
	0 (0)
	A. hydrophila (6), A. veronii (1)
	0 (0)
	A. veronii (5), A. hydrophila/caviae (2)

	A. taiwanensis
	4
	0 (0)
	A. caviae (4)
	0 (0)
	A. hydrophila/caviae (4)

	A. sanarellii
	3
	0 (0)
	A. caviae (3)
	0 (0)
	A. hydrophila/caviae (3)

	A. dhakensis
	3
	0 (0)
	A. hydrophila (3)
	0 (0)
	A. hydrophila/caviae (3)

	A. hydrophila
	2
	2 (100)
	0 (0)
	0 (0)
	A. hydrophila/caviae (2)

	Total
	217
	192 (88.48)
	
	8 (3.69)
	


Note: Numbers in parentheses indicate the number of isolates misidentified as the listed species. MALDI TOF MS, matrix assisted laser desorption ionization–time of flight mass spectrometry. 
