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	Sample condition
	

	Sample type
	Assay
	
	ZIKV-
spiked
	Mock-
infected
	Total number 
of samples

	Whole blood
	RT-iiPCR
	Positive
	74
	0
	74

	
	
	Negative
	26
	20
	46

	
	
	Total
	100
	20
	120

	
	CDC-PAHO 
RT-qPCR
	Positive
	60
	0
	60

	
	
	Negative
	40
	20
	60

	
	
	Total
	100
	20
	120

	Plasma
	RT-iiPCR
	Positive
	86
	0
	86

	
	
	Negative
	14
	20
	34

	
	
	Total
	100
	20
	120

	
	CDC-PAHO 
RT-qPCR
	Positive
	74
	0
	74

	
	
	Negative
	26
	20
	46

	
	
	Total
	100
	20
	120

	Serum
	RT-iiPCR
	Positive
	90
	0
	90

	
	
	Negative
	10
	20
	30

	
	
	Total
	100
	20
	120

	
	CDC-PAHO 
RT-qPCR
	Positive
	78
	0
	78

	
	
	Negative
	22
	20
	42

	
	
	Total
	100
	20
	120

	Semen
	RT-iiPCR
	Positive
	17
	0
	17

	
	
	Negative
	3
	4
	7

	
	
	Total
	20
	4
	24

	
	CDC-PAHO 
RT-qPCR
	Positive
	15
	0
	15

	
	
	Negative
	5
	4
	9

	
	
	Total
	20
	4
	24

	Urine
	RT-iiPCR
	Positive
	73
	0
	73

	
	
	Negative
	27
	20
	47

	
	
	Total
	100
	20
	120

	
	CDC-PAHO 
RT-qPCR
	Positive
	57
	0
	57

	
	
	Negative
	43
	20
	63

	
	
	Total
	100
	20
	120

	Mosquito pools
	RT-iiPCR
	Positive
	5
	0
	5

	
	
	Negative
	1
	1
	2

	
	
	Total
	6
	1
	7

	
	CDC-PAHO 
RT-qPCR
	Positive
	5
	0
	5

	
	
	Negative
	1
	1
	2

	
	
	Total
	6
	1
	7



