
Supplemental Table 1 MR parameters
	1.5T GE  Signa HDxt 

	Parameter
	FIESTA
	T2WI
	DWI
b=800
	TIWI
FS
	TIWI
DUAL
	T1WI+C
Mph
	T1WI+C
	T1WI+C
Delay

	Image mode
	2D
	2D
	2D
	2D
	2D
	3D
	3D
	3D

	Special parameters
	GRE
	FRFSE
	EPI
	FSPGR
	FSPGR
	LAVA
	LAVA
	LAVA

	Breath
	Hold breath
	Respiratory triggering
	Hold breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath

	Scan plane
	Cor
	Ax
	Ax
	Ax
	Ax
	Ax
	Cor
	Ax

	Fat suppression
	
No
	
Yes
	
No
	
Yes
	In-Phase
Out-Phase
	
Yes
	
Yes
	
Yes

	TR (ms)
	3.4
	8000
	6667
	220
	220
	3.3
	3.5
	3.3

	TE (ms)
	1.5
	86
	69
	2.2
	2.2/4.8
	1.6
	1.7
	1.6

	FOV (cm)
	38×38
	38×29
	38×29
	38×34
	38×34
	38×34
	40×40
	38×34

	Slice thickness         (mm)
	5
	6
	6
	6
	6
	4
	3
	4

	Spacing
	0.5
	1.2
	1.2
	1.2
	1.2
	-2
	-1.5
	-2

	Slices
	20
	24
	24
	24
	24
	92
	88
	92

	Matrix
	224×
224
	288×
224
	128×
128
	288×
170
	288×
170
	288×
160
	300×
224
	288×
160

	NEX
	1
	2
	4
	1
	1
	0.72
	0.71
	0.72






Supplemental Table 1-continued MR parameters
	3.0T GE  Discovery MR750w

	Parameters
	T2WI
	T2WI
	DWI
b=800
	LAVA
FLEX
	IDEAL
-IQ
	T1W+C
Mph
	T1W+C
	T1W+C
delay

	Image mode
	2D
	2D
	2D
	3D
	3D
	3D
	3D
	3D

	Special parameters
	SSFSE
	FSE-XL
	EPI
	LAVA
	IDEAL
	LAVA
	LAVA
	LAVA

	Breath
	Hold breath
	Respiratory triggering
	Respiratory triggering
	Hold breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath

	Scan plane
	Cor
	Ax
	Ax
	Ax
	Ax
	Ax
	Cor
	Ax

	Fat suppression
	No
	Yes
	Yes
	No/Yes
	-
	Special
	Special
	Special

	TR (ms)
	Min
	8000
	8000
	6.6
	7.3
	4.1
	4.2
	4.1

	TE (ms)
	Min
	85
	Min
	Min
Full
	Min
	1.9
	2
	1.9

	FOV (cm)
	42×42
	38×38
	38×38
	38×34
	40×40
	38×34
	40×40
	38×34

	Slice thickness         (mm)
	5
	6
	6
	4
	8
	4
	3
	4

	Spacing (mm)
	0.5
	1.2
	1.2
	-2
	0
	-2
	-1.5
	-2

	Slices
	24
	24
	24
	96
	32
	96
	136
	96

	
Matrix
	288×
256
	288×
224
	128×
128
	288×
180
	160×
160
	288×
170
	300×
224
	288×
170

	NEX
	0.53
	2
	2
	1
	0.5
	0.72
	0.71
	0.72









Supplemental Table 1-continued MR parameters
	3.0T Siemens MAGNETOM Prisma  

	Parameters
	T2WI
	T2WI
	Diff
	q-dixon
	Vibe
	T1WI+C
Vibe
	T1WI+C
Vibe
	T1WI+C
Delay

	Image mode
	2D
	2D
	2D
	3D
	3D
	3D
	3D
	2D

	Special parameters
	Haste
	TSE
	EPI
	-
	Dixon
	Dixon
	Dixon
	Dixon

	Breath mode
	Hold breath
	Respiratory triggering
	Free breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath

	Scan plane
	Cor
	Tra
	Tra
	Tra
	Tra
	Tra
	Cor
	Tra

	Fat suppression
	No
	SPAIR
	SPAIR
	-
	-
	Yes
	Yes
	Yes

	
TR (ms)
	800
	2000-
5000
	2400-
3000
	9
	3.89
	3.89
	4.21
	3.89

	TE (ms)
	91
	83
	58
	1.05
	1.23
	1.23
	1.34
	1.23

	FOV (mm) 
× %
	360×
100%
	400×
87.5%
	420×
65.7﹪
	450×
87.5%
	400×
81.3%
	400×
81.3%
	400×
81.3%
	400×
81.3%

	Slice thickness            (mm)
	
5
	
6
	
6
	
3.5
	
3
	
3
	
1.5
	
3

	Spacing
	20%
	30%
	30%
	20%
	20%
	20%
	20%
	20%

	Slices
	20
	30
	30
	64
	72
	72×5 phases
	144
	72

	Matrix size
	320×
100%
	384×
80﹪
	134×
100%
	160×
79%
	288×
75%
	288×
75%
	320×
90%
	288×
75%

	NEX
	1
	3
	1
	1
	1
	1
	1
	1





Supplemental Table 1-continued MR parameters
	3.0T Philips Ingenia

	Parameters
	T2WI
	T2WI
	Diff
	Mdxion Quant
	mDIXON
	mDIXON
	mDIXON cor
	mDIXON

	Image mode
	2D
	2D
	2D
	3D
	3D
	3D
	3D
	2D

	Special parameters
	SSTSE
	TSE
	EPI
	-
	Dixon
	Dixon
	Dixon
	Dixon

	Breath mode
	Hold breath
	Respiratory triggering
	Free breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath
	Hold breath

	Scan plane
	Cor
	Tra
	Tra
	Tra
	Tra
	Tra
	Cor
	Tra

	Fat suppression
	No
	SPIR
	SPAIR
	-
	Yes
	Yes
	Yes
	Yes

	
TR (ms)
	924
	2000-
5000
	2000
	5.7
	3.7
	3.6
	3.5
	3.6

	TE (ms)
	80
	78
	69
	0.97
	1.32
	1.31
	1.14
	1.31

	FOV (mm)
	450×
398
	380×
380
	400×
272
	400×
300
	400×
362
	400×
352
	380x
418
	400×
352

	Slice thickness            (mm)
	
5
	
6
	
6
	
6
	
4
	
4
	
3
	
4

	Spacing (mm)
	0.6
	1.2
	1.2
	-3
	-2
	-2
	-1.5
	-2

	Slices
	24
	24
	24
	80
	100
	100×2 phases
	80
	100

	Matrix
	376×
290
	280×
280
	124×
75
	168×
119
	252×
228
	252×
222
	200×
220
	252×
222

	
NEX
	
1
	
1
	
2
	
1
	
1
	
1
	
1
	
1




Note: FIESTA, fast imaging employing steatly-state acquisition; T2WI, T2 weighted imaging; DWI, diffusion weighted imaging; T1WI, T1 weighted imaging; LAVA, Liver Acquisition with Volume Acceleration; GRE, gradient echo; FRFSE, fast relaxation fast spin echo; EPI, echo planar imaging; FSPGR, fast spoiled gradient recalled echo; TR, repetition time; TE, echo time; FOV, field of view; NEX, number of average; IDEAL-IQ, Iterative Decomposition of water and fat with Echo Asymmetry and Least squares estimation; SSFSE, single-shot fast spin echo; TSE, turbo spin echo; HASTE, half-Fourier acquisition single shot turbo spin echo; SPIR, spectral presaturation inversion recovery; SPAIR, spectrally selective attenuated inversion recovery; VIBE, volume interpolated body examination. 

