## Sheet1
S. No | Device name | K-number | Product code | Year of review | Applicant | Indication for use | AI/ML algorithm used | Clinical application/utility | Predicate device | Geographical distribution | Real-World Deployment Status | Specialty
1 | Clevaligner Software (V1.0.0) | K233616 | PNN,LLZ | November 13, 2023 → January 11, 2024 | Clevaligner Ltd | The Clevaligner Software is intended for use as a medical device standalone software, providing tools for management of orthodontic models, systematic inspection, detailed analysis, treatment simulation and virtual design of a series of dental casts, which may be used for sequential aligner trays or retainers, based on 3D models of the patients dentition before the start of an orthodontic treatment. It can also be applied during the treatment to inspect and analyze the progress of the treatment. It can be used at the end of the treatment to evaluate if the outcome is consistent with the planned/desired treatment objectives. The use of Clevaligner Software requires the user to have the necessary training and domain knowledge in the practice of orthodontics, as well to have received a dedicated training in the use of the software. | Did not specify | A performance study was conducted to validate the functionality, safety, and effectiveness of the Clevaligner Software. The software testing followed the software life cycle processes defined in IEC 62304 and the FDA Guidance Document on General Principles of Software Validation. The study aimed to ensure that all user needs and performance requirements were met according to the design inputs. Verification and validation testing were performed prior to the release of the Clevaligner Software to validate its functionality and support its safety and effectiveness. The testing included the validation of implemented mitigations. | SoftSmile, Inc. Vision (K212770). | Jerusalem, Israel | Deployed | Orthodontics
2 | Overjet Charting Assist | K233590 | QIH,LLZ | November 8, 2023 → February 23, 2024 | Overjet, Inc | Overjet Charting Assist is a Medical Image Management and Processing System (MIMPS) intended to detect natural dental structures including detection of tooth anatomy (enamel, pulp), and tooth numbering, as well as dental structures added through past restorative treatments: implants, crowns, endodontic treatment (previous root canal treatment), fillings. The device is intended to assist dental professionals in producing dental charts based on image analysis. The Overjet Charting Assist detects these findings on bitewing (BW) and periapical (PA) images for patients with primary and/or permanent teeth (Ages 5 and above), and panoramic (Pano) radiographs for patients with only permanent teeth. The device is not intended as a replacement for a complete clinician's review or clinical judgment that considers other relevant information from the image or patient history. | Did not specify | A standalone performance study was conducted to evaluate the performance of the Overjet Charting Assist device. The study included 634 images obtained from male and female patients of various age groups and from different clinics across the United States. The primary endpoints of the study included tooth level sensitivity of past restorative treatment, dental tooth anatomy, and tooth level sensitivity of tooth numbering. The device was compared to a consensus reference standard established by trained dentists using majority pixel voting. The results of the study demonstrated that the device performed well when compared to the reference standard. The tooth level sensitivity for past restorative structures ranged from 83.6% to 90.5% across different age groups, with a specificity ranging from 95.1% to 98.8%. The tooth level sensitivity for dental tooth anatomy ranged from 89.2% to 98.4%. The performance of the device varied across different clinical sites and regions in the United States. Overall, the performance study showed that the Overjet Charting Assist device performed effectively in detecting dental structures and producing dental charting data based on image analysis. | Denti.AI Auto-Chart | Boston, MA, US | Deployed | Oral radiology
3 | NemoScan | K232698 | QIH,LLZ | September 5, 2023 → January 18, 2024 | Software Nemotec S.L. | NemoScan is an implant planning and surgery planning software tool intended for use by dental professionals who have appropriate knowledge in the field of application. The software reads imaging information output from medical scanners such as CBCT or CT scanners. It is indicated for pre-operative simulation and evaluation of patient anatomy, dental implant placement, surgical instrument positioning, and surgical treatment options, in edentulous, partial edentulous, or dentition situations, which may require a surgical guide. It is further indicated for the user to design such guides for, alone or in combination, the guiding of a surgical path along a trajectory or a profile, or to help evaluate a surgical preparation or step. NemoScan software allows for surgical guide export to a validated manufacturing center or to the point of care. Manufacturing at the point of care requires a validated process using CAM equipment (additive manufacturing system, including software and associated tooling) and compatible material (biocompatible and sterilizable). A surgical guide may require to be used with accessories. NemoScan is intended to be used only for Adolescents (12-21 years of age) to adults (>21 years of age) patients. | Did not specify | The document does not provide specific details about the design of any performance study that was conducted for the device NemoScan (K232698). | coDiagnostiX Implant Planning Software (K130724). | Madrid, Spain | Deployed | Implant and oral tissue regeneration
4 | Adravision Perio | K232440 | QIH,CLA,DEV,LLZ | August 14, 2023 → December 5, 2023 | Adra Corporation | Adravision Perio is a radiological semi-automated image processing software device intended to aid dental professionals in the measurements of mesial and distal bone levels associated with each tooth from bitewing and periapical radiographs. It should not be used in-lieu of full patient evaluation or solely relied upon to make or confirm a diagnosis. The intended users of Adravision Perio are trained professionals including, but not limited to, dentists and dental hygienists. The intended patient population of the device is patients living in the United States, who are 22 years old or older, and that do not have any remaining primary teeth. Adra Corporation has not evaluated the performance of the device on primary dentition. | Yes | A performance study was conducted to evaluate the design of the Adravision Perio device. The study included bench testing, which consisted of evaluating the device's performance based on a ground truth data set. The evaluation involved assessing precision, recall, sensitivity, specificity, and mean absolute error for measurements obtained from dental radiographs. The study analyzed various factors such as patient age, gender, clinic site, region of the mouth, and sensor type. The results showed that Adravision Perio met the prespecified performance requirements, demonstrating robustness and reliability in drawing line segments and accurately measuring bone levels on relevant teeth. No adverse effects or complications were recorded during the study. Based on the study findings, Adravision Perio was determined to be substantially equivalent to the predicate device, Overjet Dental Assist. | Overjet Dental Assist | San Jose, CA, US | Deployed | Periodontology
5 |  |  |  |  | CureMetrix, Inc.
6 | Videa Dental Assist | K232384 | MYN | August 8, 2023 → December 15, 2023 | VideaHealth, Inc. | Videa Dental Assist is a computer-assisted detection (CADe) device that analyzes intraoral radiographs to identify and localize the following features. Videa Dental Assist is indicated for the review of bitewing, periapical, and panoramic radiographs acquired from patients aged 3 years or older. Suspected Dental Findings: - Caries - Attrition - Broken/Chipped Tooth - Restorative Imperfections - Pulp Stones - Dens Invaginatus - Periapical Radiolucency - Widened Periodontal Ligament - Furcation - Calculus Historical Treatments: - Crown - Filling - Bridge - Post and Core - Root Canal - Endosteal Implant - Implant Abutment - Bonded Orthodontic Retainer - Braces Normal Anatomy: - Maxillary Sinus - Maxillary Tuberosity - Mental Foramen - Mandibular Canal - Inferior Border of the Mandible - Mandibular Tori - Mandibular Condyle - Developing Teeth - Erupting Teeth - Non-matured Erupted Teeth - Exfoliating Teeth - Impacted Teeth - Crowding Teeth | Yes | A performance study was conducted to evaluate the design of the Videa Dental Assist device. The study involved analyzing intraoral radiographs to identify and localize various dental findings, historical treatments, and normal anatomy. The device was tested with bitewing, periapical, and panoramic radiographs obtained from patients aged 3 years and older. The study aimed to assess the device's ability to accurately detect and highlight regions of interest in the radiographic images. The results of the performance study demonstrated that the Videa Dental Assist device effectively identified and localized the intended features in the radiographs, providing valuable assistance to dental professionals in their diagnostic activities. | Videa Caries Assist (K213795). | Boston, MA, US | Deployed | Oral radiology
7 | DTX Studio Clinic (4.0) | K231898 | LLZ,MYN | 2023 | Nobel Biocare AB | DTX Studio Clinic is a software program for the acquisition, management, transfer, and analysis of dental and craniomaxillofacial image information, and can be used to provide design input for dental restorative solutions. It displays and enhances digital images from various sources to support the diagnostic process and treatment planning. It stores and provides these images within the system or across computer systems at different locations. It can be used to support guided implant surgery whereby the results can be exported. DTX Studio Clinic is a computer-assisted detection (CADe) device that analyzes intraoral radiographs to identify and localize dental findings, which include caries, calculus, periapical radiolucency, root canal filling deficiency, discrepancy at the margin of an existing restoration, and bone loss. The DTX Studio Clinic CADe functionality is indicated for use by dentists for the concurrent review of bitewing and periapical radiographs of permanent teeth in patients 15 years of age or older. | Yes | A summary of the design of any performance study that was conducted is not provided in the document. | DTX Studio Clinic 3.0 (Focus Area Detection) with product code MYN (K221921). | Sweden | Deployed | Oral radiology
8 |  |  |  |  | Qure.ai Technologies
9 | CEPHX- Cephalometric Analysis Software | K231396 | LLZ,QIH | May 15, 2023 → January 31, 2024 | Orca Dental AI LTD | CEPHX - Cephalometric Analysis Software is indicated for use by dentists who provide orthodontic treatment for image analysis, simulation, profilogram, VTO (Visual Treatment Objective), and patient consultation. Results produced by the software's diagnostic, treatment planning, and simulation tools are dependent on the interpretation of trained and licensed practitioners or dentists. The device is only for use on patients 14 years old and above. | Yes | A performance study was conducted to verify the accuracy of the Al-generated landmarks in cephalometric tracing compared to manually generated landmarks by orthodontic specialists. The study aimed to evaluate the reliability and precision of the Cephx software in detecting critical cephalometric landmarks. The study involved comparing 21 clinically significant landmarks detected automatically by the Al algorithm to manually detected landmarks by three orthodontic specialists. The pass criteria were set as a margin of up to 2.0mm, and any margin above this range was considered a fail. The results showed that 99% of landmarks were identified, meeting the study endpoint and predefined hypothesis. The secondary endpoints also demonstrated very good accuracy without any outliers. The device's stand-alone performance was established against the ground truth, with at least 85% of cases meeting the pass criteria. | WebCeph | Israel | Deployed | Orthodontics
10 |  |  |  |  | Neusoft Medical Systems Co., Ltd.
11 | Denti.AI Detect | K230144 | MYN | January 18, 2023 → October 6, 2023 | Denti.AI Technology, Inc. | Denti.AI Detect 1s a Computer-Assisted Detection (CADe) software device intended to be used by dental professionals, comprising dentists and dental specialists, while reading extraoral and intraoral 2D dental radiographs. The device aims to assist in detecting and highlighting uncategorized regions of interest (ROIs) within the teeth area, which include caries and periapical radiolucency, as a second reader. The device is also intended to aid in the measurements of mesial and distal bone levels associated with each tooth. The device is aimed to be used with images from the patients of 22 years age and older without remaining primary dentition. The device is not intended to replace a complete clinician's review or clinical judgment that considers other relevant information from the image or patient history. | Yes | A performance study was conducted for the Denti.AI Detect system. The study included two main components: detecting suspicious regions of interest (ROIs) such as caries and periapical radiolucencies, and making measurements of distal and mesial bone levels associated with each tooth. For the detection of suspicious ROIs, 709 images were obtained from six sites in the US. Ground truthing was performed by three independent dentists, and the primary endpoint was to show that the lower bound of the 95% confidence interval (CI) of the across-category weighted Alternative Free-response Receiver Operating Characteristic (wAFROC) area under the curve (AUC) exceeds 0.6. The study met its primary endpoint with an across-category wAFROC AUC of 0.737 and a 95% CI of (0.713, 0.761). For the measurement of bone levels, 193 images were obtained from nine sites in the US. Ground truthing was performed by three independent dentists, and the primary study endpoints included sensitivity, specificity, mean absolute error (MAE), and ratios. The results showed high sensitivity and specificity for various measurements, with MAE values ranging from 0.513 mm to 0.735 mm and ratio MAE values ranging from 3.8% to 4.3%. These performance studies demonstrate the effectiveness of the Denti.AI Detect system in detecting suspicious ROIs and making accurate measurements of bone levels associated with each tooth. | The primary predicate device for the Denti.AI Detect is the Pearl Second Opinion (K210365). The secondary predicate device is the Overjet Dental Assist (K210187). | Toronto, Ontario, Canada | Deployed | Oral radiology
12 | Atomica Planner | K230012 | LLZ | January 3, 2023 → February 2, 2023 | Atomica Technology, Inc. | "Atomica Planner" is stand-alone software indicated for use by highly qualified dental professionals in the preoperative simulation-based planning of implant placements, and designing patients' 3D model surgical guides to support dental treatment surgeries (such as teeth replacement for Adult patients with missing/damaged teeth; taking into account dental functional & esthetic aspects). | Yes | A performance study was conducted to evaluate the design of the Atomica Planner software. The study focused on the software's image manipulation and processing methods, including 2D/3D visualization and segmentation tools. Dental professionals used the software to simulate pre-operative implant placements and design surgical guides. The study aimed to assess the software's ability to accurately detect suitable implant placements based on anatomical measurements and evaluations of surrounding tissues. Additionally, the study examined the software's functionality in exporting plan datasets and generating final surgical instructions and 3D models compatible with manufacturing systems. The results of the performance study indicated that the design of the Atomica Planner software met the intended objectives and demonstrated its effectiveness in supporting dental treatment surgeries. | 3Shape Implant Studio | Atlanta, GA, US | Deployed | Implant and oral tissue regeneration
13 | Videa Perio Assist | K223296 | QIH | October 26, 2022 → February 6, 2023 | VideaHealth, Inc. | Videa Perio Assist is a radiological semi-automated image processing software device intended to aid dental professionals in the measurements and visualization of mesial and distal bone levels associated with each tooth from bitewing and periapical radiographs. Measurements are made available as linear distances or relative percentages. It should not be used in-lieu of full patient evaluation or solely relied upon to make or confirm a diagnosis. The system is to be used by trained professionals including, but not limited to, dentists and dental hygienists. | Yes | A performance study was conducted for the Videa Perio Assist device. The study included both bench testing and clinical testing. For the bench testing, 996 radiographs and 16,131 landmarks were used. Bitewing and periapical radiographs were labeled as ground truth, and the Videa Perio Assist measurements were compared to the ground truth. The acceptance criteria for recall and precision were set at greater than 82%. The Videa Perio Assist met the acceptance criteria with a recall of 94.4% and precision of 84.3% for bitewing radiographs. For periapical radiographs, the recall was 91.9%. In a subgroup study for CEJ-ABL measurements, the precision was 79.1%. Clinical testing involved 189 radiographs and over 2,350 lines. US licensed dentists labeled the data, and two US licensed periodontists adjudicated the labels to establish a reference standard. The Videa Perio Assist predictions were compared to the reference standard. The acceptance criteria for sensitivity, specificity, and mean absolute error were set at greater than 82%, greater than 81%, and less than 1.5mm, respectively. The Videa Perio Assist met the acceptance criteria with a sensitivity of 92.8%, specificity of 89.4%, and mean absolute error below the thresholds for bitewing radiographs. For periapical radiographs, the sensitivity was 88.3%, specificity was 87.0%, and mean absolute error was below the thresholds. The subgroups analysis showed that all subgroups, except for CEJ-ABL in periapical radiographs, met the pre-specified acceptance criteria. No adverse events were observed during the clinical study. Overall, the performance study demonstrated that the Videa Perio Assist device meets the performance requirements. | Overjet Dental Assist (Overjet, Inc.). | Boston, MA, US | Deployed | Periodontology
14 | Ez3D-i/E3 | K222069 | LLZ | July 14, 2022 → September 6, 2022 | Ewoosoft Co., Ltd | Ez3D-i is dental imaging software that is intended to provide diagnostic tools for maxillofacial radiographic imaging. These tools are available to view and interpret a series of DICOM compliant dental radiology images and are meant to be used by trained medical professionals such as radiologists and dentists. Ez3D-1 is intended for use as software to load, view and save DICOM images from CT, panorama, cephalometric and intraoral imaging equipment and to provide 3D visualization, 2D analysis, in various MPR (Multi-Planar Reconstruction) functions. | Did not specify | A performance study was conducted to establish the performance, functionality, and reliability characteristics of the modified devices. The study included software verification/validation and measurement accuracy tests. The device passed all of the tests based on pre-determined Pass/Fail criteria. The study followed the guidelines outlined in IEC 62304 and ISO 14971. The purpose of the study was to validate the performance of the device and ensure its effectiveness for its intended use. | Ez3D-i v5.3(E3) | Hwaseong-si, Gyeonggido , Republic of Korea | Deployed | Oral radiology
15 | Denti.AI Auto-Chart | K222054 | LLZ | July 12, 2022 → November 22, 2022 | Denti.AI Technology Inc. | Denti.AI Auto-Chart is a Medical Image Management and Processing System (MIMPS) device aimed to assist dental professionals (Users), comprising general dentists, dental specialists, and dental hygienists, in detecting dental structures and producing dental charting data based on the interpretation of intraoral and extraoral 2D X-Ray images. | Yes | A performance study was conducted for the Denti.AI Auto-Chart device. The study included the following design: - Dataset: The testing dataset consisted of 336 images taken from multiple dental clinics across the US. The patient population was roughly uniformly distributed by age and gender. The dataset included various imaging modalities and sensor manufacturers. - Reference Standard: Ground truth annotations (GT) were used as the reference standard. Two experienced dental hygienists established the GT with the help of an experienced dentist in cases of disagreement. - Study Results: The study evaluated various metrics, including sensitivity, positive predictive value (PPV), overall classification accuracy, specificity, classification accuracy for restorative findings, binding dental findings to teeth, classifying fillings by type, classifying fillings by surface, and classifying crowns by type. The results showed high accuracy and confidence intervals for the evaluated metrics. Overall, the performance study demonstrated the effectiveness and accuracy of the Denti.AI Auto-Chart device in detecting dental structures and producing dental charting data based on 2D X-Ray images. | Ewoosoft EzOrtho V1.3 | Toronto, Ontario, Canada | Deployed | Oral radiology
16 |  |  |  |  | Mediso Medical Imaging Systems, LTtd.
17 | DTX Studio Clinic 3.0 | K221921 | MYN | July 1, 2022 → March 28, 2023 | Nobel Biocare AB | DTX Studio Clinic is a computer-assisted detection (CADe) device that analyses intraoral radiographs to identify and localize dental findings, which include caries, calculus, periapical radiolucency, root canal filling deficiency, discrepancy at margin of an existing restoration, and bone loss. The DTX Studio Clinic CADe functionality is indicated for use by dentists for the concurrent review of bitewing and periapical radiographs of permanent teeth in patients 15 years of age or older. | Yes | A performance study was conducted to evaluate the design of the DTX Studio Clinic 3.0. The device is a computer-assisted detection (CADe) system that analyzes intraoral radiographs to identify and localize dental findings. The study aimed to assess the accuracy and effectiveness of the device in detecting various dental conditions, including caries, calculus, periapical radiolucency, root canal filling deficiency, discrepancy at the margin of an existing restoration, and bone loss. The study design involved collecting a dataset of intraoral radiographs from adult patients aged 15 years and older. The radiographs were independently reviewed by dental professionals, who marked the presence or absence of each dental finding. The same set of radiographs was then processed by the DTX Studio Clinic 3.0 CADe system, which automatically detected and annotated suspected areas of interest. The performance of the device was evaluated by comparing its detection results with the independent reviews of the dental professionals. Measures such as sensitivity, specificity, positive predictive value, and negative predictive value were calculated to assess the accuracy and reliability of the device in detecting the different dental conditions. Based on the results of the performance study, the DTX Studio Clinic 3.0 demonstrated high sensitivity and specificity in detecting caries, calculus, periapical radiolucency, root canal filling deficiency, discrepancy at the margin of an existing restoration, and bone loss. The device showed promising performance and can be considered as a reliable tool for assisting dentists in the detection and localization of dental findings in intraoral radiographs. | The predicate device for the DTX Studio Clinic 3.0 is the Videa Caries Assist (K213795) and the Second Opinion (K210365). | Belgium | Deployed | Oral radiology
18 | Overjet Calculus Assist | K220928 | MYN | March 31, 2022 → December 16, 2022 | Overjet Inc. | Did not specify | Yes | A performance study was conducted to evaluate the design of the Overjet Calculus Assist software. The study included standalone testing and a clinical reader evaluation. In the standalone testing, the Overjet AI algorithm was tested on bitewing and periapical radiographs. The sensitivity and specificity of the algorithm were summarized based on surfaces, and confidence intervals were provided. The results showed an overall sensitivity of 74.1% for bitewing radiographs and 72.9% for periapical radiographs, with high specificity values. In the clinical reader evaluation, 14 US licensed dentists evaluated a set of radiographs with and without calculus. The dentists were asked to draw a box around suspected calculus and review the predictions from the Overjet Calculus Assist model. The readers provided a confidence rating for their labels. The evaluation showed improved sensitivity for the readers when assisted by the software, compared to unassisted evaluation. Overall, the performance study demonstrated that the Overjet Calculus Assist software met the prespecified requirements and improved the sensitivity of dentists in detecting calculus on radiographs. | Logicon Caries Detector, manufactured by Carestream Dental. | Boston, MA, US | Deployed | Periodontology
19 |  |  |  |  | Siemens Medical Solutions USA Inc.
20 | Dentis s-Clean Abutment Mini | K220440 | NHA | February 16, 2022 → June 16, 2022 | Dentis Co., Ltd. | The Dentis s-Clean Abutment Mini is an endosseous dental implant that is indicated for surgical placement in the upper and lower jaw arches, to provide a root form means for single or multiple-units prosthetic appliance attachment to restore a patients chewing function. Implants can be placed with a conventional two-stage surgical process with an option for transmucosal healing or they can be placed in a single-stage surgical process for immediate loading when good primary stability has been achieved and with appropriate occlusal loading. | Did not specify | A performance study was not mentioned in the document, so there is no information available to summarize about the design of any performance study conducted. | Dentis Dental Implant System (K171027). | California, USA | Deployed | Implant and oral tissue regeneration
21 |  |  |  |  | Aidoc Medical, Ltd.
22 | Videa Caries Assist | K213795 | MYN | December 6, 2021 → April 21, 2022 | VideaHealth, Inc | Did not specify | Yes | A standalone performance assessment study was conducted to evaluate the performance of the Videa Caries Assist device. The dataset used in the study consisted of 1034 adult radiographs collected from 10 US sites. The radiographs were ground-truthed by three US board-certified dentists. The study measured the performance of Videa Caries Assist in detecting carious lesions in dental radiographs. The device returned a set of bounding boxes representing the carious lesions detected. The study reported an overall average Alternative Free-response Receiver Operating Characteristic Figure of Merit (AFROC FOM) of 0.740 with a corresponding average image-based sensitivity of 70.8% and positive predictive value of 59.5%. The false positive fraction (FPF) was observed to be 0.335, and the non-lesion fraction was 0.599. The performance of Videa Caries Assist was assessed against potential subject and image confounders such as age, sex, number of lesions per image, image quality, imaging sensor model, effective resolution, bit-depth, and image size. The study observed very good generalizability for all confounders, except for the image sensor model. However, further analysis showed that the aided performance for the Schick 33 sensor improved compared to unaided reads. No animal testing was conducted for the device, and clinical data from a reader study was also included in the submission. | Logicon Caries Detection Software (Carestream Dental LLC). | Boston, MA, US | Deployed | Oral radiology
23 | ERI Handheld Dental X-ray System (Model AG100) | K213282 | EHD,MUH | October 1, 2021 → April 4, 2022 | Energy Resources International Co., Ltd. Hsinchu Branch | ERI Handheld Dental X-Ray System (Model AG100) is intended for the purpose of using it for both adult and pediatric patients by qualified operators (e.g. dentist and/or radiographer, etc.) regulated by competent authority of each country for producing diagnostic dental X-ray images. | Did not specify | A non-clinical performance study was conducted for the AG100 handheld dental X-ray system. The study included non-clinical design verification and validation tests, such as electrical performance evaluations and usability tests. The results of the performance evaluations demonstrated that the device design is well-suited for its intended use. No clinical testing has been conducted on this product, and successful bench testing results were deemed sufficient to demonstrate substantial equivalence for the AG100 handheld system. | KaVo NOMAD Pro 2 Handheld X-Ray System, cleared by the FDA under the 510(k) number K173319. | Taiwan | Unknown | Oral radiology
24 | Overjet Caries Assist | K212519 | MYN | August 10, 2021 → May 10, 2022 | Overjet, Inc. | Did not specify | Yes | For the Overjet Caries Assist software, performance testing was conducted according to FDA's guidance. The software is a radiological, automated, concurrent read, computer-assisted detection (CAD) system intended to aid in the detection and segmentation of caries on bitewing radiographs. The performance study aimed to evaluate the system's ability to detect carious lesions accurately. The study utilized a dataset of bitewing radiographs and compared the software's annotations of suspected carious lesions to human expert annotations. The performance metrics assessed included sensitivity, specificity, and positive predictive value. The results showed that the software achieved an increase in dentists' sensitivity of approximately 20% compared to expert annotations. The level of concern for the performance study was determined to be moderate. The study demonstrated that the Overjet Caries Assist software performed effectively in identifying carious lesions on bitewing radiographs and provided additional information for dentists to use in their diagnosis. Overall, the performance study confirmed the software's substantial equivalence to the predicate device, the Carestream Logicon Caries Detector, which was down-classified to Class II. Both systems aim to support dental professionals in their diagnosis and treatment planning for caries, utilizing automated annotation and visualization of suspected carious lesions on dental radiographs. | Logicon Caries Detector, manufactured by Carestream Dental. The Logicon Caries Detector was approved under PMA P980025 and later down-classified to Class II. | Boston, MA, US | Deployed | Oral radiology
25 | Will3D | K211114 | LLZ | April 14, 2021 → December 15, 2021 | HDX WILL CORP. | Will3D is a software application used for the display and 3D visualization of medical image files from scanning devices, such as Dental CT scanner. It is intended for use by radiologists, clinicians, referring physicians, and other qualified individuals to retrieve, process, render, review, store, print, assist in diagnosis, and distribute images utilizing standard PC hardware. Additionally, Will3D is a preoperative software application used for the simulation and evaluation of dental implants, orthodontic planning, and surgical treatments. | Did not specify | A performance study was conducted to evaluate the design of the Will3D software application. The study aimed to assess the software's safety and effectiveness. The specific details and methodology of the performance study are not provided in the document. | The predicate device for the Will3D is the device with the 510(k) number K170180, also manufactured by HDX WILL Corp. | Republic of Korea | Unknown | Oral radiology
26 | Second Opinion | K210365 | MYN | February 8, 2021 → March 4, 2022 | Pearl Inc. | Second Opinion is a computer-aided detection ("CADe") software to identify and mark regions in relation to suspected dental findings which include Caries, Discrepancy at the margin of an existing restoration, Calculus, Periapical radiolucency, Crown (metal, including zirconia & non-metal), Filling (metal & non-metal), Root canal, Bridge, and Implants. It is designed to aid dental health professionals to review bitewing and periapical radiographs of permanent teeth in patients 12 years of age or older as a second reader. | Yes | A performance study was conducted for Second Opinion, which is a computer-aided detection (CADe) software device. The study aimed to assess the effectiveness of the device in detecting pathologic dental features in radiographs of permanent teeth. It included multiple-reader, multiple-case (MRMC) studies to evaluate the performance of the device. The analysis involved calculating the changes in sensitivity and the number of false-positive detections for different types of dental caries and other pathologic features. The study demonstrated statistically significant improvement in aided reader performance across all pathologic dental features. Overall, the performance study focused on evaluating the accuracy and effectiveness of Second Opinion in aiding dental health professionals in the detection of abnormal dental features in radiographs. | Logicon Caries Detection by Carestream Dental LLC. | Beverly Hills, CA, US | Unknown | Oral radiology
27 | Overjet Dental Assist | K210187 | LLZ | January 25, 2021 → May 19, 2021 | Overjet, Inc. | Overjet Dental Assist is a radiological semi-automated image processing software device intended to aid dental professionals in the measurements of mesial and distal bone levels associated with each tooth from bitewing and periapical radiographs. | Did not specify | A performance study was conducted to evaluate the design of the Overjet Dental Assist software. The study included bench testing and clinical testing. Bench testing involved evaluating the software's performance using a ground truth dataset and Object Keypoint Similarity assessment. The software was tested with two datasets of bitewing and periapical radiographs, evaluating precision and recall against labeled keypoints. The software met acceptable performance criteria with average precision of 82.3% for bitewing and 83.0% for periapical, and average recall of 89.5% for bitewing and 90.0% for periapical. Clinical testing utilized retrospective clinical subject data from bitewing and periapical radiographs. Three US licensed dentists labeled interproximal bone levels using the measurement tool in Dexis Webview software, which were then adjudicated by two US Dental Radiologists. The software met pre-specified acceptance criteria of >85% sensitivity and specificity compared to the ground truth, with an accuracy of <1.5mm mean absolute difference versus ground truth. Overall, the performance study demonstrated that the Overjet Dental Assist software performed well in measuring interproximal bone levels in radiographs, meeting prespecified performance requirements and showing high sensitivity, specificity, and accuracy. | EZOrtho, manufactured by Ewoosoft Co., Ltd. The 510(k) number for the predicate device is K192888. | Boston, MA, US | Deployed | Oral radiology
28 |  |  |  |  | Aidoc Medical, Ltd.
29 |  |  |  |  | Optellum Ltd
30 |  |  |  |  | BriefCase
31 | VARO Plan | K200078 | LLZ,PNN | January 14, 2020 → April 10, 2020 | 3D Industrial Imaging Co., Ltd. | VARO Plan Software is a stand-alone Windows-based software application to support the treatment planning for dental implantation. It is designed for qualified dental practitioners, including dentists and lab technicians. The software imports the medical image dataset of the patient in DICOM format from medical CT or dental CBCT scanners for pre-operative planning and simulation of dental placement. It is intended for use as pre-operative planning software for the placement of dental implant(s) based on imported CT image data, optionally aligned to an optical 3D surface scan. Virtual Crowns can be used for optimized implant positioning under the prosthetic aspect. The digital three-dimensional model of a surgical guide for a guided surgery can be designed based on the approved implant position. This 3D data can be exported to manufacture a separate physical product. Indications of the dental implants do not change with guided surgery compared to conventional surgery. Use of the software requires that the user has the necessary medical training in implantology and surgical dentistry. | Did not specify | The document does not provide specific information about the design of any performance study conducted for the device. | The predicate devices mentioned in the document are: 1. Device 1: - 510(k) Number: K180629 - Device Name: Dentig Guide - Manufacturer: 3D Industrial Imaging Co. Ltd 2. Device 2: - 510(k) Number: K152078 - Device Name: Implant StudioTM - Manufacturer: 3Shape Medical A/S | Republic of Korea | Unknown | Implant and oral tissue regeneration
32 | NemoScan | K192571 | LLZ | September 18, 2019 → December 11, 2019 | Software Nemotec S.L. | NemoScan is software indicated for supporting the diagnostic and treatment planning process for Dental Implantology. NemoScan is software that is also used as an image segmentation system and for the transfer of imaging information from a scanner such as a CT scanner. | Did not specify | According to the provided document, there is no specific information regarding the design of any performance study that was conducted for the device NemoScan (K192571). | The predicate device for K192571, NemoScan, is a legally marketed device in interstate commerce prior to May 28, 1976, or a device that has been reclassified in accordance with the provisions of the Federal Food, Drug, and Cosmetic Act (Act) that do not require approval of a premarket approval application (PMA). | Colorado, USA | Available; used for diagnosis and planning of implantology treatments. | Implant and oral tissue regeneration
33 | Ez3D-i/E3 | K190791 | LLZ | March 27, 2019 → April 19, 2019 | EwooSoft Co., Ltd. | The indications for use are as follows: " Ez3D-i is dental imaging software that is intended to provide diagnostic tools for maxillofacial radiographic imaging. These tools are available to view and interpret a series of DICOM compliant dental radiology images and are meant to be used by trained medical professionals such as radiologist and dentist. Ez3D-i is intended for use as software to load, view and save DICOM images from CT, panorama, cephalometric and intraoral imaging equipment and to provide 3D visualization, 2D analysis, in various MPR (Multi-Planar Reconstruction) functions." | Did not specify | The document does not provide specific information about the design of any performance study that was conducted for the device mentioned (Ez3D-i/E3). Therefore, no summary of the design of a performance study can be provided based on the given document. | The predicate device for the Ez3D-i/E3 (K190791) is the Ez3D-i/E3 (K173863) by Ewoosoft Co., Ltd. | Republic of Korea | Available; provides 3D imaging solutions for dental diagnostics. | Oral radiology
34 | AnatomicAligner | K190767 | LLZ | March 25, 2019 → March 16, 2020 | The Methodist Hospital Research Institute d/b/a Houston | AnatomicAligner is software for pre-operative simulation of orthognathic surgical treatment options, based on imaging information from a medical scanner such as CT or MRI, in patients who have reached skeletal and dental maturity. | Did not specify | The 510(k) submission for the AnatomicAligner software included a summary of the design and development process. Extensive verification and validation testing was conducted to ensure that all requirements for the software were met. The testing results provided objective evidence that the outputs of the software design activity met all specified requirements. The AnatomicAligner software version 1.0 was verified and validated as meeting all requirements. No clinical testing of the AnatomicAligner software was required or performed. | The predicate device for AnatomicAligner is ProPlan, which is a software for pre-operative simulation of orthognathic surgical treatment options. | Houston, Texas, USA | Unknown | Oral and maxillofacial surgery
35 | EzDent-i / E2 / Prora View | K190087 | LLZ | January 17, 2019 → February 27, 2019 | EWOO SOFT Co., Ltd. | EzDent-i is dental imaging software that is intended to provide diagnostic tools for maxillofacial radiographic imaging. These tools are available to view and interpret a series of DICOM compliant dental radiology images and are meant to be used by trained medical professionals such as radiologists and dentists. EzDent-i is intended for use as software to acquire, view, save 2D image files, and load DICOM project files from panorama, cephalometric, and intra-oral imaging equipment. | Did not specify | A performance study was conducted to establish the performance, functionality, and reliability characteristics of the modified device. Verification, validation, and testing activities were carried out to ensure that the device met predetermined pass/fail criteria. The device passed all of the tests based on these criteria. The study aimed to demonstrate that the modified device performs effectively and reliably in accordance with its intended use. | The predicate device for the EzDent-i / E2 / Prora View is the EzDent-i v3.0, which is also manufactured by Ewoosoft Co., Ltd. The predicate device has the same intended use and technical characteristics as the subject device. | Irvine, CA, USA | Available; dental imaging software solutions. | Oral radiology
36 | DentiqAir | K183676 | LLZ | December 28, 2018 → December 19, 2019 | 3D Industrial Imaging Co., Ltd. | DentiqAir is a software application for the visualization and segmentation of imaging information of the oral-maxillofacial region. The imaging data originates from medical scanners such as CT or CBCT scanners. The dental professionals' planning data may be exported from DentiqAir and used as input data for CAD or Rapid Prototyping Systems. | Yes | A summary of the design of the performance study conducted for the DentiqAir device (K183676) is not provided in the document. | The predicate device for the product mentioned in the document is "Dolphin Imaging" by PATTERSON DENTAL SUPPLY, INC. | Republic of Korea | Unknown | Oral radiology
37 | uWS-CT | K183170 | LLZ | November 16, 2018 → June 24, 2019 | Shanghai United Imaging Healthcare Co., Ltd. | uWS-CT is a software solution intended to be used for viewing, manipulation, communication, and storage of medical images. It supports interpretation and evaluation of examinations within healthcare institutions. It has the following additional indications: The CT Oncology application is intended to support fast-tracking routine diagnostic oncology, staging, and follow-up, by providing a tool for the user to perform the segmentation and volumetric evaluation of suspicious lesions in lung or liver. The CT Colon Analysis application is intended to provide the user a tool to enable easy visualization and efficient evaluation of CT volume data sets of the colon. The CT Dental application is intended to provide the user a tool to reconstruct panoramic and paraxial views of jaw. The CT Lung Density Analysis application is intended to segment pulmonary, lobes, and airway, providing the user quantitative parameters, structure information to evaluate the lung and airway. The CT Lung Nodule application is intended to provide the user a tool for the review and analysis of thoracic CT images, providing quantitative and characterizing information about nodules in the lung in a single study, or over the time course of several thoracic studies. The CT Vessel Analysis application is intended to provide a tool for viewing, manipulating, and evaluating CT vascular images. The Inner view application is intended to perform a virtual camera view through hollow structures (cavities), such as vessels. The CT Brain Perfusion application is intended to calculate the parameters such as: CBV, CBF, etc. in order to analyze functional blood flow information about a region of interest (ROT) in the brain. The CT Heart application is intended to segment heart and extract coronary artery. It also provides analysis of vascular stenosis, plaque and heart function. The CT Calcium Scoring application is intended to identify calcifications and calculate the calcium score. The CT Dynamic Analysis application is intended to support visualization of the CT datasets over time with the 3D/4D display modes. The CT Bone Structure Analysis application is intended to provide visualization and labels for the ribs and spine, and support batch function for intervertebral disk. The CT Liver Evaluation application is intended to provide processing and visualization for liver segmentation and vessel extraction. It also provides a tool for the user to perform liver separation and residual liver segments evaluation. | Did not specify | The document does not provide specific information about the design of any performance study that was conducted for the uWS-CT device. | The predicate device for K183170 is the uWS-CT, which has a 510(k) number of K173001. | Shanghai, China | Unknown | Oral diagnostic imaging
38 | Portable X-ray System (Model: MiniX-V, Mini X-S) | K181891 | EHD | July 13, 2018 → August 10, 2018 | Digimed Co., LTD. | The indications for use are as follows: "The device is a diagnostic x-ray system, which is intended to be used by trained dentists and dental technicians as an extra-oral x-ray source for producing diagnostic x-ray images using intra-oral image receptors. Its use is intended for both adults and pediatric subjects." | No | A performance study was conducted for the Portable X-ray System (Models: MiniX-V, MiniX-S) as part of the submission. Details about the design of the performance study are not provided in the document. | The predicate device for the Portable X-ray System (Models: MiniX-V, MiniX-S) is the Portable X-Ray System (Model: MiniX-V, MiniX-S), which is classified as an Extra-oral Source X-Ray System under regulation 21 CFR 872.1800. | Washington, USA | Unknown | Oral radiology
39 | Reflect Implant System | K180924 | DZE,NHA | April 9, 2018 → March 29, 2019 | ids - integrated dental systems | ReflectTM Dental Implants are indicated for use in partially or fully edentulous patients to support maxillary and mandibular single-unit, multiple-unit, and overdenture dental restorations. ReflectTM Dental Implants are indicated for immediate loading when good primary stability is achieved and with appropriate occlusal loading. | No | A performance study was conducted as part of the submission. The non-clinical data included sterilization validation, biocompatibility testing, dynamic compression-bending testing, and engineering and dimensional analysis of the implant bodies, abutments, and fixation screws. However, no clinical data were included in this submission. | The primary predicate device for the Reflect Implant System (K180924) is K120414, the OsseoSpeed Plus by Astra Tech AB. | San Diego, CA, USA | Unknown | Implant and oral tissue regeneration
40 | DentiqGuide | K180629 | LLZ | March 9, 2018 → July 19, 2018 | 3D Industrial Imaging Co., Ltd | NA | Did not specify | A performance study was conducted to verify the design of the DentiqGuide software. The study included various tests to evaluate the accuracy and reliability of the software. These tests included: - Accuracy test for measurements made in the software using a phantom, comparing the true values of the phantom to the measured values in the software. - Accuracy test to ensure that the generated output surgical guide matched the user input requirements, such as guide thickness and offset from teeth to guide. - Accuracy test to verify the size of the implants provided by the software's implant library, comparing them to the real size values of the implants. The results of these tests demonstrated that the DentiqGuide software performed accurately and met the specified criteria for performance. This study confirmed that the software is designed to effectively support treatment planning for dental implantation. | The predicate device for the DentiqGuide (K180629) is the Implant StudioTM by 3Shape Medical A/S (K152078). | Irvine, CA, USA | Unknown | Implant and oral tissue regeneration
41 | Ez3D-i /E3 | K173863 | LLZ | December 20, 2017 → February 23, 2018 | Ewoo Soft Co., Ltd. | Ez3D-i is dental imaging software that is intended to provide diagnostic tools for maxillofacial radiographic imaging. These tools are available to view and interpret a series of DICOM compliant dental radiology images and are meant to be used by trained medical professionals such as radiologists and dentists. | Did not specify | A performance study was conducted to verify, validate, and test the performance, functionality, and reliability characteristics of the modified device. The study aimed to assess the device's performance in terms of its technical characteristics, general function, application, and intended use. The device underwent testing based on predetermined pass/fail criteria, and it successfully passed all tests. The results of the study demonstrated that the device is safe and effective in accordance with its intended use. | The predicate device for K173863 is the Ez3D-i /E3 (K163539) by EWOO SOFT Co., Ltd. | Irvine, CA, USA | Available; provides 3D imaging solutions for dental diagnostics. | Oral radiology
42 | CS Imaging | K173622 | LLZ | November 24, 2017 → January 16, 2018 | Trophy | NA | Did not specify | A performance study was conducted for the CS Imaging device. The device is digital imaging software intended for use with Carestream Dental's digital imaging devices. Healthcare professionals use the software to display, adjust, make measurements, print, export, and store digital or digitized images to support image diagnosis in medical care, predominantly in dentistry. The study evaluated the performance of the software in terms of its ability to display, adjust, measure, print, export, and store images. The specific details and results of the performance study are not provided in the given document. | The predicate device for K173622 is legally marketed predicate devices marketed in interstate commerce prior to May 28, 1976, the enactment date of the Medical Device Amendments, or devices that have been reclassified in accordance with the provisions of the Federal Food, Drug, and Cosmetic Act (Act) that do not require approval of a premarket approval application (PMA). | France | Unknown | Oral radiology
43 | Curve Image 2.0 PACS | K173101 | LLZ | September 29, 2017 → October 18, 2017 | Curve Dental LTD | Indications for Use: Curve Image is an internet-based, image management software (PACS) that enables dental offices to keep records of hard and soft tissue charts in a form of digital images. The system uses a Web-based interface and includes acquisition, editing, and storage of digital images. Images and data can be stored, communicated, processed, and displayed within the system or across computer networks at distributed locations. Images can be acquired from standard dental imaging devices or can be uploaded directly from the user's computer. Images can be edited (e.g., zoomed, contrast, rotated, etc.), as well as exported. The system is designed to provide images for diagnostic use. | Did not specify | Based on the provided document, it does not explicitly mention a performance study or provide specific details about its design. Therefore, there is no information available to summarize regarding the design of any performance study conducted for the device mentioned. | The predicate device for the Curve Image 2.0 PACS is a legally marketed device in interstate commerce prior to May 28, 1976, or a device that has been reclassified in accordance with the provisions of the Federal Food, Drug, and Cosmetic Act (Act) that do not require approval of a premarket approval application (PMA). Unfortunately, the specific details of the predicate device are not mentioned in the provided document. | Calgary, Canada | Unknown | Oral diagnostic imaging
44 | 3DIEMME RealGUIDE | K173041 | LLZ | September 28, 2017 → December 20, 2018 | 3DIEMME Ltd. | PC/MAC version: The RealGUIDE software is intended for the following uses: 1. Support to the diagnosis for trained professionals. The input DICOM files acquired by a CT/CBCT scanner are not modified in any way but they are showed to the doctor through the classical imaging and volume rendering techniques. It is a stand-alone product. No information of the patient is modified, all the parameters used for the image processing are read from the DICOM file itself. Neither automatic diagnosis is made, nor automatic disease detection is performed. This software is not connected to any medical instrumentation and it doesn't control any medical or energy supplying device. The user imports DICOM data coming from any CT/CBCT imaging device and the software enables him to view the Patient exam in different multi-planar 2D images and easily reconstruct the 3D volume for an immediate visualization of bone structures and surrounding tissues. 2. Virtual oral and maxillofacial surgery planning. Doctors can plan virtual implants and surgeries on 2D/3D reconstructions and export the projects in open or proprietary format for further processing. The user can choose different implant models (for example dental implants models) from a library provided by the Manufacturers and simulate the positioning in the Patient reconstructed volume (this operation is called virtual plan) 3. Dental/maxillofacial surgical guides and prosthetic modelling. The virtual plan is used to design a surgical guide that is used by the doctor to drive the surgery drills according to the planned implants direction and depth. This surgical guide can be manufactured by any 3D printer working from STL files. The user can also design the patient prosthesis (typically a denture) with the surface and volume free-form tools implemented in the software. The result is exported in STL format for 3D printing or CAD/CAM technologies. Mobile version: The RealGUIDE software APP is intended for the following uses: 1. Projects visualization and editing. The input PROJECT files, pre-processed with the RealGUIDE desktop version, are used by trained professionals to evaluate the implants projects, edit them and share them with other colleagues through the cloud, as well as for a more effective Patient treatment communication. The RealGUIDE APP version is NOT INTENDED for managing a 3D diagnosis starting from DICOM images, due to the mobile devices screen resolution limitations. For this reason, the APP is not reading directly the DICOM files but only pre-processed project files, exported through the cloud by the RealGUIDE desktop version. | Did not specify | A performance study was conducted to verify the implementation and functionality of the RealGUIDE system. The study followed the applicable procedures and documents provided by 3DIEMME. The verification, validation, and testing procedures were conducted according to the RealGUIDE Software System Validation Procedure (SVP) documents. The study included non-clinical tests and a comparison between the RealGUIDE system and the predicate device's performance. The hardware requirements for the software use were also evaluated and found to be equivalent to the predicate and reference devices. The testing confirmed that the implementation functions as intended and that there are no additional safety or effectiveness concerns compared to the predicate device. | The predicate device for the 3DIEMME RealGUIDE is the Implant StudioTM 2015-1, manufactured by 3Shape Medical A/S. | Virginia, USA | Unknown | Oral and maxillofacial surgery/oral radiology/implant and tissue regeneration
45 | uWS-CT | K173001 | LLZ | September 27, 2017 → November 7, 2018 | Shanghai United Imaging Healthcare Co.,Ltd. | Indications for Use: uWS-CT is a software solution intended to be used for viewing, manipulation, communication, and storage of medical images. It supports interpretation and evaluation of examinations within healthcare institutions. It has the following additional indications: - The CT Oncology application is intended to support fast-tracking routine diagnostic oncology, staging, and follow-up, by providing a tool for the user to perform the segmentation and volumetric evaluation of suspicious lesions in lung or liver. - The CT Colon Analysis application is intended to provide the user a tool to enable easy visualization and efficient evaluation of CT volume data sets of the colon. - The CT Dental application is intended to provide the user a tool to reconstruct panoramic and paraxial views of the jaw. - The CT Lung Density application is intended to provide the user a number of density parameters and structure information for evaluating tomogram scans of the lung. - The CT Lung Nodule application is intended to provide the user a tool for the review and analysis of thoracic CT images, providing quantitative and characterizing information about nodules in the lung in a single study, or over the time course of several thoracic studies. - The CT Vessel Analysis application is intended to provide a tool for viewing, manipulating, and evaluating CT vascular images. - The Inner view application is intended to perform a virtual camera view through hollow structures (cavities), such as vessels. | Did not specify | A performance study was conducted for the uWS-CT software solution. The study aimed to evaluate the functionality and effectiveness of the software in processing, reviewing, and analyzing CT studies. The software allows for image browsing, processing, and advanced information retrieval for diagnosis. The study assessed the software's ability to display, review, and process images in both 2D and 3D formats, including functionalities such as volume rendering, multi-planar reconstruction, maximum intensity projection, and statistical analysis. Additionally, specific applications within the software, such as CT oncology, CT colon analysis, CT dental, CT lung density, CT lung nodule, CT vessel analysis, and inner view, were evaluated for their intended purposes, including segmentation, visualization, and evaluation of specific anatomical structures. The performance study aimed to validate the software's capabilities and ensure its compliance with regulatory requirements for medical image processing and analysis. | The predicate device for the uWS-CT is the syngo.via (K123920) and Philips Plus CT Scanner. | Shanghai, China | Unknown | Oral diagnostic imaging
46 | Straumann Dental Implant System | K171784 | DZE | June 15, 2017 → January 19, 2018 | Straumann USA, LLC | NA | NA | NA | NA | Massachusetts, USA | Available; widely used implant system in dental practices. | Implant and oral tissue regeneration
47 | Straumann PURE Ceramic Implants | K171769 | DZE,NHA | June 14, 2017 → November 14, 2017 | Straumann USA, LLC (on behalf of Institut Straumann AG) | The Straumann PURE Ceramic Implant (Monotype) is indicated for restoration in single tooth gaps and in an edentulous or partially edentulous jaw. The prosthetic restorations used are single crowns, fixed partial or full dentures, which are connected to the implants through the corresponding components. The 3.3 mm reduced diameter implants are recommended for central and lateral incisors only. | Did not specify | A performance study was conducted to evaluate the design of the Straumann PURE Ceramic Implants. The study included the following aspects: 1. Sterilization and Shelf-Life: A sterilization validation was performed using the Half Cycle Overkill Approach, demonstrating the capability of reliably sterilizing the implants to a minimum Sterility Assurance Level (SAL) of 10. Transportation and packaging stability studies were also conducted to ensure the integrity of the sterile barrier system throughout handling, distribution, and storage. 2. Biocompatibility Testing: The biological safety of the implants was re-evaluated, and further biocompatibility testing was conducted according to ISO and FDA standards. Chemical analysis confirmed the absence of harmful residues, and cytotoxicity testing showed no adverse effects. 3. Bench Testing: Mechanical testing, including static, dynamic fatigue, and torque testing, was performed to assess the implants' performance. These tests were conducted following ISO and FDA guidance for root-form endosseous dental implants. The implants demonstrated satisfactory mechanical performance and surface chemistry, even after hydrothermal aging. Overall, the performance study confirmed that the design of the Straumann PURE Ceramic Implants met the necessary safety and effectiveness requirements. | The predicate device for the Straumann PURE Ceramic Implants is the K151328 Straumann PURE Ceramic Implants. | Massachusetts, USA | Available; ceramic implants used in dental practices. | Implant and oral tissue regeneration
48 | Planmeca Romexis | K171385 | LLZ | May 11, 2017 → November 14, 2017 | Planmeca Oy | Indications for Use (Describe) Planmeca Romexis is a medical imaging software intended for use in dental and medical care as a tool for displaying and visualizing dental and medical 2D and 3D image files from imaging devices, such as projection radiography and CBCT. It is intended for use by radiologists, clinicians, referring physicians, and other qualified individuals to retrieve, process, render, diagnose, review, store, print, and distribute images of both adult and pediatric patients. Planmeca Romexis is also a preoperative software application used for dental implant planning. Based on the planned implant position, a model of a surgical guide for a guided implant surgery can be designed. The designed data can be exported to manufacture a separate physical product. Additionally, Planmeca Romexis includes monitoring features for Planmeca devices for maintenance purposes. The software is designed to work as a stand-alone or as an accessory to Planmeca imaging and Planmeca dental unit products in a standard PC. The software is for use by authorized healthcare professionals. Use of the software for implant planning requires that the user has the necessary medical training in implantology and surgical dentistry. Indications of the dental implants do not change with guided surgery compared to conventional surgery. Planmeca Romexis is also intended to be used for monitoring, recording, storing, and displaying mandibular jaw positions and movements relative to the maxilla. This device is not indicated for mammography use. | Did not specify | A performance study was conducted to evaluate the design of Planmeca Romexis. The study involved non-clinical bench testing to compare the images rendered by Planmeca Romexis with a predicate software version. The results confirmed that the software applications are equally effective in performing the essential functions and provide substantially equivalent clinical data. | The predicate device for K171385 is the Anatomage InVivoDental (K123519). | Helsinki, Finland | Available; dental imaging software platform. | Implant and oral tissue regeneration/oral radiology
49 | EzRay Air W | K163705 | EHD | December 29, 2016 → January 24, 2017 | Vatech Co., Ltd | The indications for use of the device are as follows: "The EzRay Air W (Model: VEX-S300W) is an intra-oral dental X-ray system (extra-oral X-ray source system) intended for use by a trained and qualified dentist or dental technician for both adult and pediatric subjects for producing diagnostic dental radiographs for treatment of diseases of the teeth, jaw, and other oral structures using intra-oral image receptors." | Did not specify | A performance study was conducted to evaluate the design and performance of the EzRay Air W (Model: VEX-S300W) dental X-ray system. The study involved comparing the subject device to the predicate device (VEX-S100W) and assessing various parameters. The performance testing confirmed that the focal spot to skin distance for both devices met the minimum length requirement of 20 cm. The accuracy of loading factors and reproducibility of Air KERMA (kinetic energy released per unit mass) also met the essential performance requirements. Both devices demonstrated similar performance outcomes in terms of half-value layer (HVL), limitation of the X-ray field, and leakage radiation test, meeting the X-ray performance standards specified in Federal Standard 21CFR 1020.30 and 31. The safety, electromagnetic compatibility (EMC), and performance data of the subject device were provided and demonstrated compliance with the relevant standards. The study confirmed the safety and effectiveness of the Auto Mode function, which was previously demonstrated and confirmed in the reference device (EzRay Air, Model: VEX-P300). Overall, the performance study concluded that the subject device, EzRay Air W (Model: VEX-S300W), is substantially equivalent to the predicate device in terms of design, technology, functions, and principle of operation. | The predicate device for the EzRay Air W (Model: VEX-S300W) is the VEX-S100W. | Houston, Texas, USA | Unknown | Oral radiology
50 | Ez3D-i / E3 | K163539 | LLZ | December 16, 2016 → December 23, 2016 | EWOO SOFT Co., Ltd. | Indications for Use: Ez3D-i is dental imaging software that is intended to provide diagnostic tools for maxillofacial radiographic imaging. These tools are available to view and interpret a series of DICOM compliant dental radiology images and are meant to be used by trained medical professionals such as radiologists and dentists. Ez3D-1 is intended for use as software to load, view and save DICOM images from CT, panorama, cephalometric and intraoral imaging equipment and to provide 3D visualization, 2D analysis, in various MPR (Multi-Planar Reconstruction) functions. | Did not specify | A performance study was conducted to verify, validate, and test the performance, functionality, and reliability characteristics of the modified device. The study aimed to assess the device's performance in relation to its technical characteristics, general function, application, and intended use. The study utilized pre-determined pass/fail criteria to evaluate the device's performance. Based on the results of the study, it was concluded that the modified device is substantially equivalent to the predicate device in terms of its design and performance. | The predicate device for the K163539 application is the Ez3D-i / E3 (K161246) by EWOO SOFT Co., Ltd. | Irvine, CA, USA | Available; provides 3D imaging solutions for dental diagnostics. | Oral radiology
51 | Neo GBR System | K160991 | DZL | April 8, 2016 → January 10, 2017 | NEOBIOTECH CO., LTD. | Indications for Use: Neo GBR System is intended to fixate and stabilize non-resorbable barrier membranes used for regeneration of tissue in the oral cavity or in dental situations that require membrane use or fixation. | Did not specify | A performance study was conducted to evaluate the Neo GBR System. The study involved bench testing to confirm the safe and effective physical properties of the device. The testing included visual tests, dimension tests, compressive loads, pull out, driving torque, and insertion-elimination tests in accordance with ASTM F543-13. The results of the testing met the pre-defined acceptance criteria and intended uses of the device. Additionally, biocompatibility tests such as cytotoxicity, irritation, and sensitization tests were performed in accordance with ISO 10993-1. The stability and compatibility of the Tent Screw and CTi-Cover Screw were also evaluated to assess adaptation accuracy. | The predicate device for the Neo GBR System is the GBR system manufactured by Jeil Medical Corporation, with the 510(k) number K143730. | CA, USA | Unknown | Implant and oral tissue regeneration
52 | DENTAL IMAGING SOFTWARE | K141607 | LLZ | June 16, 2014 → September 18, 2014 | TROPHY | The indications for use are as follows: "Carestream's Dental Imaging Software is intended to be used with Carestream's dental digital imaging devices by healthcare professionals to display, adjust, measure, print and store 2D digital images of the dento-maxillofacial region of the human anatomy to support image diagnosis." | Did not specify | According to the provided document, there is no specific information available about the design of any performance study that was conducted for the Dental Imaging Software (product code LLZ). | The predicate device for the Dental Imaging Software is a legally marketed device that was marketed in interstate commerce prior to May 28, 1976, or a device that has been reclassified in accordance with the provisions of the Federal Food, Drug, and Cosmetic Act (Act) that do not require approval of a premarket approval application (PMA). The specific details of the predicate device are not mentioned in the provided document. | France | Unknown | Oral radiology
53 | ORTHOVISION | K131570 | LLZ | May 30, 2013 → August 15, 2013 | EWOO SOFT CO., LTD | Indications for Use: OtrhoVision software is indicated for use by orthodontists for image analysis, simulation, profilogram, growth forecast, VTO/STO and patient consultation. Results produced by the softwares diagnostic, treatment planning and simulation tools are dependent on the interpretation of trained and licensed practitioners or dentists. | Did not specify | A performance study was conducted to evaluate the design of the OrthoVision software. The study aimed to assess the safety and effectiveness of the software in performing image analysis, simulation, profilogram, growth forecast, VTO/STO, and patient consultation. The software was tested for its ability to accurately interpret and analyze images, provide diagnostic information, and assist in treatment planning. The study evaluated the software's performance in comparison to a predicate device, CS Orthodontic and OMS of Carestream Dental LLC, to demonstrate substantial equivalence in terms of technical characteristics and intended use. The study results showed that OrthoVision meets the requirements for safety and effectiveness, making it suitable for use by orthodontists and licensed practitioners in orthodontic analysis and treatment planning. | The predicate device for the Ortho Vision software is the CS Orthodontic and OMS of Carestream Dental LLC. | Republic of Korea | Unknown | Orthodontics
54 | ORTOMED | K131348 | LLZ | May 10, 2013 → August 5, 2014 | INFOMED SERVICIOS INFORMATICOS S.L. | ORTOMED is intended for use by specialized dental practices for capturing, storing and presenting patient images and radiographs and to aid in cephalometric analysis, orthodontic and orthognathic surgery treatment planning and communication as well as case follow-up. Results produced by the software tools are to be interpreted by trained and licensed dental practitioners. | Did not specify | A performance study was conducted to evaluate the design of the Ortomed software. The study involved the identification, evaluation, and control of potential hazards according to ISO 14971 standards. The design and development of the software followed documented procedures and specific protocols in line with FDA guidance documents. Integration, verification, and validation testing were successfully completed following ISO 62304 standards. The study aimed to establish the safety and performance of the Ortomed software and concluded that it is substantially equivalent to the predicate devices and does not pose any new or increased risks compared to them. | The predicate device for the ORTOMED software is Dolphin Imaging 11.5 (K110430) and CS Orthodontic Imaging and CS OMS Imaging Software (K122427). | Barcelona, Spain | Unknown | Orthodontics
55 | MAVEN PRO | K130242 | LLZ | January 31, 2013 → August 29, 2013 | NSEQUENCE | Maven Pro is intended for use as a front-end software interface for the transfer of imaging information from a medical scanner such as a Dental CT scanner. It is also intended for use as a planning and simulation software in the placement of dental implants and surgical treatment. | Did not specify | The document does not provide specific information about the design of any performance study that was conducted. | The predicate device is not explicitly mentioned in the document provided. | Nevada, USA | Unknown | Implant and oral tissue regeneration
56 | NOBELCLINICIAN | K123976 | LLZ | December 26, 2012 → March 26, 2013 | NOBEL BIOCARE AB | NA | Did not specify | A performance study was conducted to verify and validate the performance of the NobelClinician software. The study followed the guidance provided in the "Guidance for the Content of Premarket Submissions for Software Contained in Medical Devices." The software's performance was assessed based on its ability to transfer and visualize imaging information from equipment such as CT scanners or magnetic resonance scanners. The study aimed to demonstrate the safety and effectiveness of the software in supporting diagnosis and treatment planning in dental and cranio-maxillofacial regions. The results of the study showed that the NobelClinician software is substantially equivalent to the predicate device. | The predicate device for the NobelClinician software is the Swissmeda Dental Planning System (K112251), The Guided Surgery Concept (K050393), and Maxilim (K052424). | Yorba Linda, CA, USA | Unknown | Implant and oral tissue regeneration
57 | SIMPLANT GO | K120733 | LLZ | March 9, 2012 → July 5, 2012 | MATERIALISE DENTAL NV | Materialise Dentals SimPlant Go software is indicated for use as a medical front-end software that can be used by medically trained people for the purpose of visualizing gray value images. It is intended for use as a pre-operative software program for simulating /evaluating dental implant placement and surgical treatment options. | Yes | A design validation was conducted to ensure the compatibility of all system components and to support the safety and effectiveness of the device. The validation activities included design validation, beta validation using mockups and prototypes, internal validation, and clinical beta validation. The design validation involved interviews and usability tests performed by an external usability company. The beta validation included additional usability tests with both external and internal users. The clinical case planning in SimPlant GO was also validated, and all validation criteria were met. The performance testing of the software consisted of unit testing, integration testing, IR testing, smoke testing, formal testing, acceptance testing, alpha testing, and beta testing. The results of the testing, including bugs and suggestions for improvements, were documented in the Design History File. The software validation activities were performed in addition to benchtop performance testing. | The predicate device for the SimPlant Go software is the SimPlant 2011 (K110300). | Leuven, Belgium | Unknown | Implant and oral tissue regeneration
58 | SIMPLANT IMMEDIATE SMILE DENTAL CARE SYSTEM | K113739 | LLZ,DZE,EBG | December 20, 2011 → May 30, 2012 | MATERIALISE DENTAL NV | NA | Yes | A performance study was conducted for the SimPlant Immediate Smile System. The study included software validation and benchtop performance testing to ensure the compatibility of all system components. Clinical data was also gathered, where the system was evaluated by 11 doctors with 20 patients. The evaluation included feedback on usability, aesthetic result, and material properties. Predefined acceptance thresholds were met for all safety and effectiveness criteria. The clinical validation confirmed the safety and effectiveness of the system. Non-clinical testing, including engineering tests, were undertaken to demonstrate the conformity of the system. The SimPlant Immediate Smile System was shown to be safe and effective, meeting all pre-set criteria on component and system levels. | The predicate device for the SimPlant Immediate Smile System is the SimPlant 2011; (K110300), the Nobel Biocare Guided Surgery Concept; (K050393), and Ivoclar Telio CAD (K093708). | Leuven, Belgium | Unknown | Implant and oral tissue regeneration
59 | DENTAL 3D PLANNING SOFTWARE | K113404 | LLZ | November 18, 2011 → December 6, 2011 | SYMBYO DENTAL, LLC | IPS is prescription use software used by dentists and dental lab technicians for the visualization and image segmentation of DICOM data from medical scanners such as CT. The software aids the user in the creation of 3D models of oral maxillofacial region and in planning dental surgical treatments and placement of dental implants. | Yes | A performance study was not mentioned in the provided document, so there is no information available to summarize regarding the design of such a study. | The predicate devices for K113404 are the following: - MATERIALISE N.V. Simplant system (K033849) - siCAT GmbH & Co. KG Sicat implant (K090119) | Florida, USA | Unknown | Implant and oral tissue regeneration
60 | SIMPLANT NAVIGATOR PERSONALIZED DENTAL CARE SYSTEM | K112679 | LLZ,DZA,NDP,NOF | September 14, 2011 → February 22, 2012 | MATERIALISE DENTAL NV | NA | Yes | A performance study was conducted to validate the SimPlant Navigator Personalized Dental Care System. The study included software validation and benchtop performance testing to ensure the compatibility of all system components. However, no specific details or results of the performance study are provided in the document. | The predicate device for the SimPlant Navigator Personalized Dental Care System is the SimPlant 2011 (K110300) and the Nobel Biocare Guided Surgery Concept (K050393). | Leuven, Belgium | Unknown | Implant and oral tissue regeneration
61 | RADIOVISION | K111102 | LLZ | April 20, 2011 → May 5, 2011 | HAAS SOFTWARE | Indications for Use: "RadioVision is a Macintosh-based image management database, or system used primarily by dentists to acquire, archive, display, enhance, print, email, and import/export digital images." | Did not specify | A performance study was conducted to evaluate the design of RadioVision, the picture archiving and communications system. The study involved the development of Software Requirement Specifications (SRS) and the execution of a Verification & Validation (V&V) plan through a series of Test Cases. The results of the V&V testing demonstrated that RadioVision performs as indicated and is substantially equivalent to the VixWin Pro software manufactured by Gendex Dental Systems. | The predicate device for K111102 is the VixWin Pro software manufactured by Gendex Dental Systems. | San Jose, CA, US | Unknown | Oral radiology
The devices highlighted in yellow were excluded from the analysis because they were not related to dentistry.
Table 1: Characteristics of included FDA-approved dental devices