AUDIT TOOL FOR MEDICAL DEVICE REPROCESSING AND STEAM STERILIZATION PRACTICES

Hospital No: _____    			Date: _ _/_ _/_ _ _ _    		Observation No: ___/___

AA. GENERAL

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AA1 | Decontamination activities take place in a dirty to clean workflow |  |  |  |  |
| AA2 | Single-use items are reprocessed |  |  |  |  |

AA3. Design of the reprocessed medical devices

☐Solid, hollow		☐Pin and box joints		☐Lumen, tubing, tortuous paths

☐Porous			☐Other (specify_____________________)

AA4. Material of the reprocessed medical devices

☐Metal 			☐Non-metal

AB. TRANSPORT

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AB1 | Medical devices are transported to the decontamination area using a rigid, durable, leak-proof container that has a tight-fitting lid |  |  |  |  |
| AB2 | Container used for transporting medical devices is easy to clean and disinfect |  |  |  |  |

AC. CLEANING & DISINFECTION

AC1. Medical devices are cleaned before sterilization.	☐Yes 			☐No

AC2. Time period between use and cleaning of medical devices	 _______ minutes

AC3. Used medical devices are soaked in or sprayed with water before cleaning to prevent drying.

☐Yes 		☐No

AC3. Personnel involved in cleaning of medical devices

☐Doctors       	☐Nurses 	      ☐HA/AHW/ANM      ☐Support staff

☐Other (Specify) ___________

AC4. Cleaning methods used

☐Manual 		☐Automated 	      ☐Both		☐None

AC5. Specific procedures and solutions used for cleaning and disinfection of medical devices before sterilization

☐Water			☐Water and detergent/soap 		☐Ultrasonic washers

☐Enzymatic cleaner	☐Disinfectant solution		☐Other (Specify)_______________

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AC6 | Cleaning is done in a separate area from where the instrument will be used (i.e., designated dirty area) |  |  |  |  |
| AC7 | Medical devices are pre-disinfected before cleaning (e.g. with hypochlorite solution) |  |  |  |  |
| AC8 | Following personal protective equipment are used during cleaning of used instruments |  |  |  |  |
|  | - Eye protection |  |  |  |  |
|  | - Gloves |  |  |  |  |
|  | - Protective clothing |  |  |  |  |
|  | - Facemask |  |  |  |  |
| AC9 | Medical devices are opened/dismantled for cleaning purpose |  |  |  |  |
| AC10 | Medical devices are submerged in water while washing them manually using a brush |  |  |  |  |
| AC11 | For instruments with lumens, all channels are cleaned using cleaning brushes of appropriate size |  |  |  |  |
| AC12 | Cleaning brushes are single use, disposable items |  |  |  |  |
| AC13 | After completion of cleaning process, reusable brushes are cleaned and either high level disinfected or sterilized |  |  |  |  |
| AC14 | Instruments are rinsed thoroughly with water after cleaning |  |  |  |  |
| AC15 | Medical devices are dried with low-linting (disposable or reusable) towels immediately after rinsing |  |  |  |  |
| AC16 | Enzymatic cleaner, detergent, and/or disinfectant are used according to manufacturer’s instructions |  |  |  |  |
| AC17 | Enzymatic cleaner, detergent, and/or disinfectant are discarded according to manufacturer’s instructions |  |  |  |  |

AD. INSPECTION

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AD1 | All instruments are inspected every time after cleaning |  |  |  |  |
| AD2 | An illuminated magnifier is used to inspect instruments |  |  |  |  |

AE. PACKAGING

AE1. Sterile barrier system used

☐Single wrapped/pouch

☐Double wrapped in wrapping material or pouches, double wrapped container or tray, reusable sterilization container according to manufacturer’s instructions

☐Combination of two or more systems, for example, a reusable sterilization container with an inner sterile barrier system

☐None

AE2. Wrapping material used

☐Paper

☐Cellulose/non-cellulose based non-woven wrapping materials

☐Cellulose/non-cellulose based woven wrapping materials

☐Linen

☐Other (specify___________)

AE3. Wrapping technique used

☐Envelope-fold wrapping technique

☐Square-fold wrapping technique

☐Other (specify____________)

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AE4 | Hinged devices are open and devices are disassembled (if indicated by the manufacturer) while packaging them |  |  |  |  |
| AE5 | Packages to be sterilized are labelled with |  |  |  |  |
|  | - The sterilizer used |  |  |  |  |
|  | - The cycle or load number |  |  |  |  |
|  | - The date of sterilization |  |  |  |  |
|  | - The expiration date |  |  |  |  |

AF. STERILIZATION (AUTOCLAVING)

AF1. Personnel involved in sterilization of medical devices (autoclaving)

☐Doctors       	☐Nurses 	      ☐HA/AHW/ANM      ☐Support staff

☐Other (Specify) ___________

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AF2 | Timer is used to monitor holding period of the autoclave cycle |  |  |  |  |
| AF3 | Holding period of the autoclave cycle starts when the pressure gauze shows the reading of required pressure (e.g.15 lbs) |  |  |  |  |
| AF4 | The following parameters are recorded for each sterilization cycle: |  |  |  |  |
|  | - Cycle/load number |  |  |  |  |
|  | - Operator |  |  |  |  |
|  | - Date and Time |  |  |  |  |
|  | - Pressure |  |  |  |  |
|  | - Temperature and exposure time |  |  |  |  |
|  | - Holding period |  |  |  |  |
| AF5 | Indicators used for monitoring sterilization process |  |  |  |  |
|  | - Autoclave tape |  |  |  |  |
|  | - Chemical Indicator |  |  |  |  |
|  | - Biological Indicator |  |  |  |  |
| AF6 | Results for indicator recorded |  |  |  |  |
|  | - Autoclave tape |  |  |  |  |
|  | - Chemical Indicator |  |  |  |  |
|  | - Biological Indicator |  |  |  |  |
| AF7 | Sterilizer physical parameters are reviewed after each run |  |  |  |  |
| AF8 | Indicator tape is used on the outside of each wrapped package |  |  |  |  |
| AF9 | Sterilized packs are intact and dry |  |  |  |  |

AG. TRANSPORT AND STORAGE

| S.No. | Check Points | Yes | No | NA | Comments |
| --- | --- | --- | --- | --- | --- |
| AG1 | Sterilized packages are checked for integrity and compromised packages are repackaged and re-sterilized before use |  |  |  |  |
| AG2 | Sterilized items are transported and delivered in a dry and clean container |  |  |  |  |
| AG3 | Sterilized packages are allowed to cool down to room temperature before storage |  |  |  |  |
| AG4 | Separate area is allocated for storage of sterilized medical devices |  |  |  |  |
| AG5 | Sterilized packages are stored and distributed according to "the first one to enter is the first one to leave" |  |  |  |  |
| AG6 | The area for storing sterilized packages is a well-ventilated area that provides protection against dust, moisture, insects, and temperature and humidity extremes |  |  |  |  |