## Master
Table of Contents
Calculations
| Fixed and Operating Cost of In-house Maintenance | here
| Administrative, Preventive Maintenance and Corrective Maintenance Cost based on PLAMAHS data for Table 2, colum 3. | here
| NPV Calculation of Derpeciation and Maintenance Cost; Robustness Check for the Life Time Reduction Assumption | here
here
Share of Considered Devices in Total
| Calculation of Depreciation based on Device Value and Life Expectancy | here
| Colculation: Downtime for District Hospitals | here
| Average Number of Serives per Day | here
| Overhead Cost | here
Data
| Invoiced Program Cost | here
| Device Values | here
| Service Prices | here
| Device Lifecycles | here
| Currency Conversion | here
| Detailed Maintenance Component Cost In-house Maintenance Calculation | here
Notes: Sheets beginning with C_ are calucations, sheets beginning with D_ store data

## C_fullhouse
M <=
Fixed and Operating Cost of In-house Maintenance
|  | Assumed Discount Rate |  |  | #REF!
Year
|  |  | 10 years | 3 years | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Initial investments |  |  | 10 years: 910424; 3 years: ‚
| Construction WS |  | 10 years: #REF!; 3 years: #REF!; Year 0: 710424
| Tools |  | 10 years: #REF!; 3 years: #REF!; Year 0: 200000
Initial Admin
|  |  | #REF! | #REF! | 3090.352 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724
Operating cost |  |  | 10 years: #REF!; 3 years: #REF!
| Salaries |  | #REF! | #REF! |  | 600000 | 600000 | 600000 | 600000 | 600000 | 600000 | 600000 | 600000 | 600000 | 600000
| Spares & testing |  | #REF! | #REF! |  | 385000 | 385000 | 385000 | 385000 | 385000 | 385000 | 385000 | 385000 | 385000 | 385000
Depreciation rate: |  |  | #REF! | #REF! |  | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551 | 1960858.6551
[5 empty] | 913514.352 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275 | 3071722.7275
NPV | total |  | 10 years: #REF!; 3 years: #REF!
NPV | per year |  | 10 years: #REF!; 3 years: #REF!
H_fullhouse |  | for data
Notes: WS = Work Shop. We asume identical management cost for in-house and contracted-out maintenance (Table 2, column 2&3)
Note on C13: Michael Schleicher: adjust for lower BoQ of only 70%

## C_contracted_PLAMAHS
M <=
Administrative, Preventive Maintenance and Corrective Maintenance Cost based on PLAMAHS data for Table 2, colum 3.
Costing based on break-down of progam cost
In USD (2016)
|  | Assumed discount rate: |  |  |  | #REF!
|  |  | 10 years | 3 years | 3 yrs (w/o In. Inv.) | Year
| Contracted MTN |  |  |  |  | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Maintenance Cost
|  | Administrative cost: | #REF! | #REF! |  | 653576.5966 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724 | 125864.0724
|  | PM operating cost: | #REF! | #REF! | #REF! | 903151.6859 | 118378.7037 | 157286.1045 | 176908.4197 | 173965.0724 | 174406.5745 | 264371.9423 | 0 | 0 | 0 | 0
|  | CM spare cost: | #REF! | #REF! | #REF! |  | 9344.7184 | 50924.0522 | 29776.119 | 41551.6966 | 53327.2741 | 65090.5962 | 0 | 0 | 0 | 0
Notes: CM = Corrective Maintenance; PM = Preventive Maintenance; MTN = Maintenance; w/o In. Inv. = Without Initial Invoice
Note on B10: Michael Schleicher: note nicely adjusted for "per hospital", yet

## C_depreciation_mtn_NPV
M <=
NPV Calculation of Derpeciation and Maintenance Cost; Robustness Check for the Life Time Reduction Assumption
| for only five device types, in USD (2016) | [12 empty] | Discount Rate: | #REF!
|  |  | NPV | [11 empty] | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
No Maintenace |  |  | 20 years | 15 years | 10 years | 9 years | 8 years | 7 years | 6 years | 5 years | 4 years | 3 years | 2 years | 1 years
Depreciation rate: |  | Lifetime Reduction
|  | 39% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 0% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925
|  | 5% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605 | 37153111.3605
|  | 10% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027 | 39217173.1027
|  | 15% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382 | 41524065.6382
|  | 20% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406 | 44119319.7406
|  | 25% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233 | 47060607.7233
|  | 30% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035 | 50422079.7035
|  | 35% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192 | 54300701.2192
|  | 40% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541 | 58825759.6541
|  | 45% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863 | 64173555.9863
|  | 50% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849 | 70590911.5849
Downtime cost
| Apporach | 1 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276 | 399683.6276
| Approach | 2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574 | 1228457.5574
| Approach | 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852 | 40537779.2852
Total Cost |  | Lifetime Reduction
|  | 39% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 0% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 5% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 10% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 15% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 20% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 25% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 30% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 35% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 40% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 45% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  | 50% | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
|  |  | NPV | [11 empty] | #REF! 0: Year
Contracted-out Maintenance |  |  | 20 years | 15 years | 10 years | 9 years | 8 years | 7 years | 6 years | 5 years | 4 years | 3 years | 2 years | 1 years | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Maintenance Cost
| PM operating cost: |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 5109016.9467 | 4520704.8606 | 5562610.2006 | 6428785.5228 | 6659880.5689 | 6659880.5689 | 6659880.5689 | 6659880.5689 | 6659880.5689 | 6659880.5689 | 6659880.5689
| CM spare cost: |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 328826.5509 | 328826.5509 | 328826.5509 | 394591.8611 | 394591.8611 | 394591.8611 | 394591.8611 | 394591.8611 | 394591.8611 | 394591.8611
Depreciation rate: |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925 | 35295455.7925
Running cost total |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Downtime cost
| Apporach | 1 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441 | 490639.8441
| Approach | 2 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774 | 774635.6774
| Approach | 3 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! |  | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2252098.8492 | 2133567.3308
Fixed Cost |  |  | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611
Total Cost |  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Notes: For our main calcuation we assume that non-maintained devices have a 39% short life time, which is a value calibrated from the literature. As a robustness check, we calculate the deprection rate if maintenance has no effect or non-maintained devices have a 50% lower life time in 5% steps.
0
Note on A2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Calc of downtime ist nur für 35% der BoQ möglich
Note on A42: Michael Schleicher: note nicely adjusted for "per hospital", yet

## C_device_share
Share of Considered Devices in Total |  |  |  | M <= | Number of Devices from each maintenance Round
Equipment type selection: |  |  |  |  | BoQ | All Devices |  |  | Only Considered Device Types
1 | Laboratory equipment |  |  |  | PPM | Tot | Crit | NonCrit | Tot |  | Crit |  | NonCrit
| a | analyser | ANL | ANL | 1 | 776 | 81 | 695 | 507 | 65.34% | 68 | 83.95% | 439 | 63.17%
| b | centrifuge | CEN | CEN | 2 | 921 | 83 | 838 | 602 | 65.36% | 71 | 85.54% | 531 | 63.37%
| c | microscope | MIC | MIC | 3 | 1000 | 85 | 915 | 658 | 65.8% | 74 | 87.06% | 584 | 63.83%
2 | Cold chain equipment |  |  |  | 4 | 1075 | 87 | 988 | 714 | 66.42% | 76 | 87.36% | 638 | 64.57%
| a | refrigerator | CLD |  | 5 | 1108 | 92 | 1016 | 730 | 65.88% | 81 | 88.04% | 649 | 63.88%
| b | freezer | CLD |  | 6 | 1132 | 97 | 1035 | 749 | 66.17% | 86 | 88.66% | 663 | 64.06%
3 | Surgical equipment |  |  |  | 7 | 97 | 97 |  | 86 | 88.66% | 86 | 88.66%
| a | anaesthesia unit | ANS |  | 8 | 98 | 98 |  | 87 | 88.78% | 87 | 88.78%
| b | electrosurgical unit | ESU |  | 9 | 96 | 96 |  | 85 | 88.54% | 85 | 88.54%
| c | operating light | LGHT |  | Notes: The numbers are for all devices, including those not in district hopitals
| d | oxygen concentrator | OXG
| e | strelizing unit | STR |  | Summary Statistics and Share Calculation
| f | operating table | TAB | [5 empty] | DH only
4 | Monitoring |  |  |  | Average PPM-BoQ |  |  | 1002 | 1002
| a | ecg | ECG | ECG | Total maximum PPM-BoQ |  |  | 1132 | 1132
| b | monitor | MON |  | Selection maximum BoQ |  |  | 749 | 419
5 | Ultrasound imaging |  | ULS | ULS | Share of total BoQ |  |  | 0.66166 | 0.37014
6 | Xray imaging |  | XRAY | XRAY | Notes: DH = District Hospitals
Note on M20: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Unteschied aus „tatischen“ Erwägungen

## C_depreciation_yearly
Calculation of Depreciation based on Device Value and Life Expectancy
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | #REF! | 132984.1675 | #REF! | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | #REF! | 519391.3936 | 81268.1024 | #REF! | #REF!
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | #REF! | 8438.0817 | #REF! | 53 | 2.9444 | 3429852.7189 | 447218.3298 | #REF! | 190547.3733 | 24845.4628 | #REF! | #REF!
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | #REF! | 24301.6753 | #REF! | 41 | 2.2778 | 8490503.3344 | 996368.6865 | #REF! | 471694.6297 | 55353.8159 | #REF! | #REF!
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | #REF! | 20356.8143 | #REF! | 22 | 1.2222 | 3816329.4807 | 447849.9145 | #REF! | 212018.3045 | 24880.5508 | #REF! | #REF!
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | #REF! | 18534.2759 | #REF! | 22 | 1.2222 | 3822120.3448 | 407754.0708 | #REF! | 212340.0192 | 22653.0039 | #REF! | #REF!
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | #REF! | 124918.0502 | #REF! | 4 | 0.22222 | 4257930.368 | 499672.2009 | #REF! | 236551.6871 | 27759.5667 | #REF! | #REF!
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | #REF! | 47253.2575 | #REF! | 13 | 0.72222 | 4711193.9234 | 614292.3474 | #REF! | 261732.9957 | 34127.3526 | #REF! | #REF!
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | #REF! | 61971.7696 | #REF! | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | #REF! | 390199.2226 | 65414.6457 | #REF! | #REF!
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | #REF! | 17087.1154 | #REF! | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | #REF! | 388285.2135 | 56957.0514 | #REF! | #REF!
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | #REF! | 24815.091 | #REF! | 21 | 1.1667 | 5328804.3943 | 521116.9101 | #REF! | 296044.6886 | 28950.9394 | #REF! | #REF!
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | #REF! | 67986.8297 | #REF! | 18 | 1 | 15642347.1709 | 1223762.9338 | #REF! | 869019.2873 | 67986.8297 | #REF! | #REF!
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | #REF! | 47084.4959 | #REF! | 21 | 1.1667 | 6740631.3057 | 988774.413 | #REF! | 374479.517 | 54931.9118 | #REF! | #REF!
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | #REF! | 109695.062 | #REF! | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | #REF! | 794546.5942 | 103600.8919 | #REF! | #REF!
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | #REF! | 142772.3423 | #REF! | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | #REF! | 3379519.4505 | 396589.8397 | #REF! | #REF!
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | #REF! | 516205.3491 | #REF! | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | #REF! | 9188639.3709 | 1347869.5226 | #REF! | #REF!
Notes: MTN = Maintenance; Avg = Average, HF = Health Facility; purch = purchasing | [9 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | #REF! | 17785009.7475 | 2393189.4873 | #REF!
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | #REF! | Special Sel: | 1960858.6551 | #REF!
[11 empty] | Share Spec Sel: | 0.29773
Robustness Checks: Different Changes in Average Life Expectancy of Euqipment, Maintenance vs. Non-Maintenance Scenarios
Life exp. Change: |  |  | 0%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 7.5 | 132984.1675 | 132984.1675 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1462825.8426 | 519391.3936 | 81268.1024 | 81268.1024 | 0%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 9 | 8438.0817 | 8438.0817 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 447218.3298 | 190547.3733 | 24845.4628 | 24845.4628 | 0%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 10 | 24301.6753 | 24301.6753 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 996368.6865 | 471694.6297 | 55353.8159 | 55353.8159 | 0%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 10 | 20356.8143 | 20356.8143 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 447849.9145 | 212018.3045 | 24880.5508 | 24880.5508 | 0%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 11 | 18534.2759 | 18534.2759 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 407754.0708 | 212340.0192 | 22653.0039 | 22653.0039 | 0%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 10 | 124918.0502 | 124918.0502 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 499672.2009 | 236551.6871 | 27759.5667 | 27759.5667 | 0%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 9 | 47253.2575 | 47253.2575 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 614292.3474 | 261732.9957 | 34127.3526 | 34127.3526 | 0%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 7 | 61971.7696 | 61971.7696 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1177463.6218 | 390199.2226 | 65414.6457 | 65414.6457 | 0%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 8 | 17087.1154 | 17087.1154 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1025226.9259 | 388285.2135 | 56957.0514 | 56957.0514 | 0%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 12 | 24815.091 | 24815.091 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 521116.9101 | 296044.6886 | 28950.9394 | 28950.9394 | 0%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 15 | 67986.8297 | 67986.8297 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1223762.9338 | 869019.2873 | 67986.8297 | 67986.8297 | 0%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 8 | 47084.4959 | 47084.4959 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 988774.413 | 374479.517 | 54931.9118 | 54931.9118 | 0%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 9 | 109695.062 | 109695.062 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 1864816.0545 | 794546.5942 | 103600.8919 | 103600.8919 | 0%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 10 | 142772.3423 | 142772.3423 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 7138617.1137 | 3379519.4505 | 396589.8397 | 396589.8397 | 0%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 8 | 516205.3491 | 516205.3491 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 24261651.4068 | 9188639.3709 | 1347869.5226 | 1347869.5226 | 0%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 43077410.7721 | 17785009.7475 | 2393189.4873 | 2393189.4873
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 35295455.7925 | Special Sel: | 1960858.6551 | 1960858.6551
Life exp. Change: |  |  | 5%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 7.125 | 132984.1675 | 139983.3342 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1539816.6764 | 519391.3936 | 81268.1024 | 85545.3709 | 5%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 8.55 | 8438.0817 | 8882.1913 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 470756.1367 | 190547.3733 | 24845.4628 | 26153.1187 | 5%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 9.5 | 24301.6753 | 25580.7108 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1048809.1437 | 471694.6297 | 55353.8159 | 58267.1746 | 5%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 9.5 | 20356.8143 | 21428.2256 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 471420.9626 | 212018.3045 | 24880.5508 | 26190.0535 | 5%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 10.45 | 18534.2759 | 19509.7642 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 429214.8114 | 212340.0192 | 22653.0039 | 23845.2673 | 5%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 9.5 | 124918.0502 | 131492.6845 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 525970.7378 | 236551.6871 | 27759.5667 | 29220.5965 | 5%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 8.55 | 47253.2575 | 49740.271 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 646623.5236 | 261732.9957 | 34127.3526 | 35923.5291 | 5%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 6.65 | 61971.7696 | 65233.4417 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1239435.3914 | 390199.2226 | 65414.6457 | 68857.5217 | 5%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 7.6 | 17087.1154 | 17986.4373 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1079186.2378 | 388285.2135 | 56957.0514 | 59954.791 | 5%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 11.4 | 24815.091 | 26121.1484 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 548544.1159 | 296044.6886 | 28950.9394 | 30474.6731 | 5%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 14.25 | 67986.8297 | 71565.0838 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1288171.5092 | 869019.2873 | 67986.8297 | 71565.0838 | 5%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 7.6 | 47084.4959 | 49562.6272 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1040815.1716 | 374479.517 | 54931.9118 | 57823.0651 | 5%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 8.55 | 109695.062 | 115468.4863 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 1962964.2679 | 794546.5942 | 103600.8919 | 109053.5704 | 5%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 9.5 | 142772.3423 | 150286.6761 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 7514333.8039 | 3379519.4505 | 396589.8397 | 417462.9891 | 5%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 7.6 | 516205.3491 | 543374.0517 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 25538580.4282 | 9188639.3709 | 1347869.5226 | 1418810.0238 | 5%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 45344642.918 | 17785009.7475 | 2393189.4873 | 2519146.8288
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 37153111.3605 | Special Sel: | 1960858.6551 | 2064061.7422
Life exp. Change: |  |  | 10%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 6.75 | 132984.1675 | 147760.1861 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1625362.0473 | 519391.3936 | 81268.1024 | 90297.8915 | 10%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 8.1 | 8438.0817 | 9375.6463 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 496909.2554 | 190547.3733 | 24845.4628 | 27606.0697 | 10%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 9 | 24301.6753 | 27001.8614 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1107076.3183 | 471694.6297 | 55353.8159 | 61504.2399 | 10%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 9 | 20356.8143 | 22618.6825 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 497611.0161 | 212018.3045 | 24880.5508 | 27645.0564 | 10%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 9.9 | 18534.2759 | 20593.6399 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 453060.0787 | 212340.0192 | 22653.0039 | 25170.0044 | 10%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 9 | 124918.0502 | 138797.8336 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 555191.3344 | 236551.6871 | 27759.5667 | 30843.963 | 10%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 8.1 | 47253.2575 | 52503.6194 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 682547.0526 | 261732.9957 | 34127.3526 | 37919.2807 | 10%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 6.3 | 61971.7696 | 68857.5217 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1308292.9131 | 390199.2226 | 65414.6457 | 72682.9396 | 10%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 7.2 | 17087.1154 | 18985.6838 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1139141.0288 | 388285.2135 | 56957.0514 | 63285.6127 | 10%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 10.8 | 24815.091 | 27572.3233 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 579018.789 | 296044.6886 | 28950.9394 | 32167.7105 | 10%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 13.5 | 67986.8297 | 75540.9218 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1359736.5931 | 869019.2873 | 67986.8297 | 75540.9218 | 10%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 7.2 | 47084.4959 | 52316.1065 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1098638.2367 | 374479.517 | 54931.9118 | 61035.4576 | 10%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 8.1 | 109695.062 | 121883.4023 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2072017.8384 | 794546.5942 | 103600.8919 | 115112.1021 | 10%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 9 | 142772.3423 | 158635.9359 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 7931796.793 | 3379519.4505 | 396589.8397 | 440655.3774 | 10%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 7.2 | 516205.3491 | 573561.499 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 26957390.452 | 9188639.3709 | 1347869.5226 | 1497632.8029 | 10%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 47863789.7468 | 17785009.7475 | 2393189.4873 | 2659099.4304
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 39217173.1027 | Special Sel: | 1960858.6551 | 2178731.839
Life exp. Change: |  |  | 15%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 6.375 | 132984.1675 | 156451.9618 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1720971.5795 | 519391.3936 | 81268.1024 | 95609.5322 | 15%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 7.65 | 8438.0817 | 9927.1549 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 526139.2116 | 190547.3733 | 24845.4628 | 29229.9562 | 15%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 8.5 | 24301.6753 | 28590.2062 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1172198.4547 | 471694.6297 | 55353.8159 | 65122.1364 | 15%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 8.5 | 20356.8143 | 23949.1933 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 526882.2523 | 212018.3045 | 24880.5508 | 29271.2362 | 15%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 9.35 | 18534.2759 | 21805.0305 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 479710.6715 | 212340.0192 | 22653.0039 | 26650.5929 | 15%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 8.5 | 124918.0502 | 146962.412 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 587849.6481 | 236551.6871 | 27759.5667 | 32658.3138 | 15%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 7.65 | 47253.2575 | 55592.0676 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 722696.8793 | 261732.9957 | 34127.3526 | 40149.8266 | 15%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 5.95 | 61971.7696 | 72907.9642 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1385251.3198 | 390199.2226 | 65414.6457 | 76958.4067 | 15%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 6.8 | 17087.1154 | 20102.4887 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1206149.3246 | 388285.2135 | 56957.0514 | 67008.2958 | 15%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 10.2 | 24815.091 | 29194.2247 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 613078.7177 | 296044.6886 | 28950.9394 | 34059.9288 | 15%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 12.75 | 67986.8297 | 79984.5055 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1439721.0986 | 869019.2873 | 67986.8297 | 79984.5055 | 15%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 6.8 | 47084.4959 | 55393.5245 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1163264.0153 | 374479.517 | 54931.9118 | 64625.7786 | 15%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 7.65 | 109695.062 | 129053.0142 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2193901.2406 | 794546.5942 | 103600.8919 | 121883.4023 | 15%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 8.5 | 142772.3423 | 167967.4615 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 8398373.075 | 3379519.4505 | 396589.8397 | 466576.2819 | 15%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 6.8 | 516205.3491 | 607300.4107 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 28543119.3021 | 9188639.3709 | 1347869.5226 | 1585728.8501 | 15%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 50679306.7908 | 17785009.7475 | 2393189.4873 | 2815517.0439
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 41524065.6382 | Special Sel: | 1960858.6551 | 2306892.5355
Life exp. Change: |  |  | 20%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 6 | 132984.1675 | 166230.2094 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1828532.3032 | 519391.3936 | 81268.1024 | 101585.128 | 20%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 7.2 | 8438.0817 | 10547.6021 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 559022.9123 | 190547.3733 | 24845.4628 | 31056.8285 | 20%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 8 | 24301.6753 | 30377.0941 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1245460.8581 | 471694.6297 | 55353.8159 | 69192.2699 | 20%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 8 | 20356.8143 | 25446.0179 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 559812.3931 | 212018.3045 | 24880.5508 | 31100.6885 | 20%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 8.8 | 18534.2759 | 23167.8449 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 509692.5885 | 212340.0192 | 22653.0039 | 28316.2549 | 20%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 8 | 124918.0502 | 156147.5628 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 624590.2512 | 236551.6871 | 27759.5667 | 34699.4584 | 20%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 7.2 | 47253.2575 | 59066.5719 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 767865.4342 | 261732.9957 | 34127.3526 | 42659.1908 | 20%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 5.6 | 61971.7696 | 77464.712 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1471829.5273 | 390199.2226 | 65414.6457 | 81768.3071 | 20%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 6.4 | 17087.1154 | 21358.8943 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1281533.6574 | 388285.2135 | 56957.0514 | 71196.3143 | 20%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 9.6 | 24815.091 | 31018.8637 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 651396.1376 | 296044.6886 | 28950.9394 | 36188.6743 | 20%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 12 | 67986.8297 | 84983.5371 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1529703.6672 | 869019.2873 | 67986.8297 | 84983.5371 | 20%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 6.4 | 47084.4959 | 58855.6198 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1235968.0162 | 374479.517 | 54931.9118 | 68664.8898 | 20%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 7.2 | 109695.062 | 137118.8275 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2331020.0682 | 794546.5942 | 103600.8919 | 129501.1149 | 20%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 8 | 142772.3423 | 178465.4278 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 8923271.3922 | 3379519.4505 | 396589.8397 | 495737.2996 | 20%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 6.4 | 516205.3491 | 645256.6864 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 30327064.2585 | 9188639.3709 | 1347869.5226 | 1684836.9033 | 20%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 53846763.4652 | 17785009.7475 | 2393189.4873 | 2991486.8592
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 44119319.7406 | Special Sel: | 1960858.6551 | 2451073.3189
Life exp. Change: |  |  | 25%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 5.625 | 132984.1675 | 177312.2233 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 1950434.4568 | 519391.3936 | 81268.1024 | 108357.4698 | 25%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 6.75 | 8438.0817 | 11250.7756 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 596291.1064 | 190547.3733 | 24845.4628 | 33127.2837 | 25%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 7.5 | 24301.6753 | 32402.2337 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1328491.582 | 471694.6297 | 55353.8159 | 73805.0879 | 25%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 7.5 | 20356.8143 | 27142.4191 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 597133.2193 | 212018.3045 | 24880.5508 | 33174.0677 | 25%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 8.25 | 18534.2759 | 24712.3679 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 543672.0944 | 212340.0192 | 22653.0039 | 30204.0052 | 25%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 7.5 | 124918.0502 | 166557.4003 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 666229.6012 | 236551.6871 | 27759.5667 | 37012.7556 | 25%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 6.75 | 47253.2575 | 63004.3433 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 819056.4632 | 261732.9957 | 34127.3526 | 45503.1368 | 25%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 5.25 | 61971.7696 | 82629.0261 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1569951.4957 | 390199.2226 | 65414.6457 | 87219.5275 | 25%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 6 | 17087.1154 | 22782.8206 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1366969.2345 | 388285.2135 | 56957.0514 | 75942.7353 | 25%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 9 | 24815.091 | 33086.7879 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 694822.5468 | 296044.6886 | 28950.9394 | 38601.2526 | 25%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 11.25 | 67986.8297 | 90649.1062 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1631683.9117 | 869019.2873 | 67986.8297 | 90649.1062 | 25%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 6 | 47084.4959 | 62779.3278 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1318365.884 | 374479.517 | 54931.9118 | 73242.5491 | 25%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 6.75 | 109695.062 | 146260.0827 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2486421.406 | 794546.5942 | 103600.8919 | 138134.5226 | 25%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 7.5 | 142772.3423 | 190363.123 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 9518156.1516 | 3379519.4505 | 396589.8397 | 528786.4529 | 25%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 6 | 516205.3491 | 688273.7988 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 32348868.5424 | 9188639.3709 | 1347869.5226 | 1797159.3635 | 25%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 57436547.6962 | 17785009.7475 | 2393189.4873 | 3190919.3165
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 47060607.7233 | Special Sel: | 1960858.6551 | 2614478.2068
Life exp. Change: |  |  | 30%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 5.25 | 132984.1675 | 189977.3822 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 2089751.2037 | 519391.3936 | 81268.1024 | 116097.2891 | 30%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 6.3 | 8438.0817 | 12054.4024 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 638883.3283 | 190547.3733 | 24845.4628 | 35493.5182 | 30%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 7 | 24301.6753 | 34716.679 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1423383.8379 | 471694.6297 | 55353.8159 | 79076.8799 | 30%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 7 | 20356.8143 | 29081.1633 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 639785.5921 | 212018.3045 | 24880.5508 | 35543.644 | 30%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 7.7 | 18534.2759 | 26477.5371 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 582505.8155 | 212340.0192 | 22653.0039 | 32361.4342 | 30%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 7 | 124918.0502 | 178454.3575 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 713817.4299 | 236551.6871 | 27759.5667 | 39656.5239 | 30%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 6.3 | 47253.2575 | 67504.6536 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 877560.4963 | 261732.9957 | 34127.3526 | 48753.3609 | 30%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 4.9 | 61971.7696 | 88531.0994 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1682090.8883 | 390199.2226 | 65414.6457 | 93449.4938 | 30%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 5.6 | 17087.1154 | 24410.1649 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1464609.8942 | 388285.2135 | 56957.0514 | 81367.2163 | 30%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 8.4 | 24815.091 | 35450.1299 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 744452.7287 | 296044.6886 | 28950.9394 | 41358.4849 | 30%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 10.5 | 67986.8297 | 97124.0424 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1748232.7625 | 869019.2873 | 67986.8297 | 97124.0424 | 30%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 5.6 | 47084.4959 | 67263.5655 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1412534.8757 | 374479.517 | 54931.9118 | 78474.1598 | 30%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 6.3 | 109695.062 | 156707.2315 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2664022.935 | 794546.5942 | 103600.8919 | 148001.2742 | 30%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 7 | 142772.3423 | 203960.489 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 10198024.4482 | 3379519.4505 | 396589.8397 | 566556.9138 | 30%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 5.6 | 516205.3491 | 737436.213 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 34659502.0097 | 9188639.3709 | 1347869.5226 | 1925527.8894 | 30%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 61539158.2459 | 17785009.7475 | 2393189.4873 | 3418842.1248
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 50422079.7035 | Special Sel: | 1960858.6551 | 2801226.6502
Life exp. Change: |  |  | 35%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 4.875 | 132984.1675 | 204591.0269 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 2250501.2963 | 519391.3936 | 81268.1024 | 125027.8498 | 35%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 5.85 | 8438.0817 | 12981.6641 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 688028.1997 | 190547.3733 | 24845.4628 | 38223.7889 | 35%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 6.5 | 24301.6753 | 37387.1927 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1532874.9023 | 471694.6297 | 55353.8159 | 85159.7168 | 35%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 6.5 | 20356.8143 | 31318.1758 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 688999.8684 | 212018.3045 | 24880.5508 | 38277.7705 | 35%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 7.15 | 18534.2759 | 28514.2707 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 627313.9551 | 212340.0192 | 22653.0039 | 34850.7753 | 35%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 6.5 | 124918.0502 | 192181.6157 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 768726.463 | 236551.6871 | 27759.5667 | 42707.0257 | 35%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 5.85 | 47253.2575 | 72697.3192 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 945065.1498 | 261732.9957 | 34127.3526 | 52503.6194 | 35%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 4.55 | 61971.7696 | 95341.184 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1811482.4951 | 390199.2226 | 65414.6457 | 100637.9164 | 35%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 5.2 | 17087.1154 | 26287.8699 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1577272.1937 | 388285.2135 | 56957.0514 | 87626.233 | 35%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 7.8 | 24815.091 | 38177.063 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 801718.3232 | 296044.6886 | 28950.9394 | 44539.9068 | 35%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 9.75 | 67986.8297 | 104595.1225 | 18 | 1 | 15642347.1709 | 1223762.9338 | 1882712.2058 | 869019.2873 | 67986.8297 | 104595.1225 | 35%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 5.2 | 47084.4959 | 72437.6859 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1521191.4046 | 374479.517 | 54931.9118 | 84510.6336 | 35%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 5.85 | 109695.062 | 168761.6339 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 2868947.7762 | 794546.5942 | 103600.8919 | 159385.9876 | 35%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 6.5 | 142772.3423 | 219649.7573 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 10982487.8673 | 3379519.4505 | 396589.8397 | 610138.2148 | 35%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 5.2 | 516205.3491 | 794162.0755 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 37325617.5489 | 9188639.3709 | 1347869.5226 | 2073645.4194 | 35%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 66272939.6494 | 17785009.7475 | 2393189.4873 | 3681829.9805
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 54300701.2192 | Special Sel: | 1960858.6551 | 3016705.6233
Life exp. Change: |  |  | 40%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 4.5 | 132984.1675 | 221640.2792 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 2438043.071 | 519391.3936 | 81268.1024 | 135446.8373 | 40%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 5.4 | 8438.0817 | 14063.4695 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 745363.883 | 190547.3733 | 24845.4628 | 41409.1046 | 40%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 6 | 24301.6753 | 40502.7921 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1660614.4775 | 471694.6297 | 55353.8159 | 92256.3599 | 40%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 6 | 20356.8143 | 33928.0238 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 746416.5241 | 212018.3045 | 24880.5508 | 41467.5847 | 40%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 6.6 | 18534.2759 | 30890.4599 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 679590.118 | 212340.0192 | 22653.0039 | 37755.0066 | 40%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 6 | 124918.0502 | 208196.7504 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 832787.0015 | 236551.6871 | 27759.5667 | 46265.9445 | 40%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 5.4 | 47253.2575 | 78755.4292 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 1023820.579 | 261732.9957 | 34127.3526 | 56878.9211 | 40%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 4.2 | 61971.7696 | 103286.2826 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 1962439.3697 | 390199.2226 | 65414.6457 | 109024.4094 | 40%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 4.8 | 17087.1154 | 28478.5257 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1708711.5432 | 388285.2135 | 56957.0514 | 94928.4191 | 40%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 7.2 | 24815.091 | 41358.4849 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 868528.1835 | 296044.6886 | 28950.9394 | 48251.5657 | 40%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 9 | 67986.8297 | 113311.3828 | 18 | 1 | 15642347.1709 | 1223762.9338 | 2039604.8896 | 869019.2873 | 67986.8297 | 113311.3828 | 40%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 4.8 | 47084.4959 | 78474.1598 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1647957.355 | 374479.517 | 54931.9118 | 91553.1864 | 40%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 5.4 | 109695.062 | 182825.1034 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 3108026.7576 | 794546.5942 | 103600.8919 | 172668.1532 | 40%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 6 | 142772.3423 | 237953.9038 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 11897695.1896 | 3379519.4505 | 396589.8397 | 660983.0661 | 40%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 4.8 | 516205.3491 | 860342.2485 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 40436085.678 | 9188639.3709 | 1347869.5226 | 2246449.2043 | 40%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 71795684.6202 | 17785009.7475 | 2393189.4873 | 3988649.1456
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 58825759.6541 | Special Sel: | 1960858.6551 | 3268097.7586
Life exp. Change: |  |  | 45%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 4.125 | 132984.1675 | 241789.3955 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 2659683.3502 | 519391.3936 | 81268.1024 | 147760.1861 | 45%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 4.95 | 8438.0817 | 15341.9667 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 813124.236 | 190547.3733 | 24845.4628 | 45173.5687 | 45%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 5.5 | 24301.6753 | 44184.8641 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1811579.43 | 471694.6297 | 55353.8159 | 100643.3017 | 45%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 5.5 | 20356.8143 | 37012.3896 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 814272.5717 | 212018.3045 | 24880.5508 | 45237.3651 | 45%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 6.05 | 18534.2759 | 33698.6835 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 741371.0378 | 212340.0192 | 22653.0039 | 41187.2799 | 45%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 5.5 | 124918.0502 | 227123.7277 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 908494.9108 | 236551.6871 | 27759.5667 | 50471.9395 | 45%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 4.95 | 47253.2575 | 85915.0136 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 1116895.177 | 261732.9957 | 34127.3526 | 62049.7321 | 45%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 3.85 | 61971.7696 | 112675.9447 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 2140842.9487 | 390199.2226 | 65414.6457 | 118935.7194 | 45%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 4.4 | 17087.1154 | 31067.4826 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 1864048.9562 | 388285.2135 | 56957.0514 | 103558.2753 | 45%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 6.6 | 24815.091 | 45118.3472 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 947485.291 | 296044.6886 | 28950.9394 | 52638.0717 | 45%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 8.25 | 67986.8297 | 123612.4176 | 18 | 1 | 15642347.1709 | 1223762.9338 | 2225023.516 | 869019.2873 | 67986.8297 | 123612.4176 | 45%
4 | Monitoring | [5 empty] | 0
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 4.4 | 47084.4959 | 85608.1743 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1797771.66 | 374479.517 | 54931.9118 | 99876.2033 | 45%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 4.95 | 109695.062 | 199445.5673 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 3390574.6446 | 794546.5942 | 103600.8919 | 188365.258 | 45%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 5.5 | 142772.3423 | 259586.0769 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 12979303.8432 | 3379519.4505 | 396589.8397 | 721072.4357 | 45%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 4.4 | 516205.3491 | 938555.1801 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 44112093.4669 | 9188639.3709 | 1347869.5226 | 2450671.8593 | 45%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 78322565.0403 | 17785009.7475 | 2393189.4873 | 4351253.6133
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 64173555.9863 | Special Sel: | 1960858.6551 | 3565197.5548
Life exp. Change: |  |  | 50%
Equipment type: |  |  | M <= | Purchasing | Purchasing | Life Expectancy |  | Depreciation rate |  | BoQ | BoQ per HF | Accum purch. | Accum. deprec. rate |  | Avg purch value | Avg depr rate per HF |  | Life exp.
|  |  |  | value (NPR,14) | value (NPR,16) | w MTN | w\o MTN | w MTN | w\o MTN | (Cat2) | (Cat2) | value (NPR,14) | w MTN | w\o MTN | per HF (NPR,14) | w MTN | w\o MTN | change
1 | Laboratory equipment
| a | analyser | ANL | 849913.1895 | 997381.2563 | 7.5 | 3.75 | 132984.1675 | 265968.335 | 11 | 0.61111 | 9349045.0842 | 1462825.8426 | 2925651.6852 | 519391.3936 | 81268.1024 | 162536.2047 | 50%
| b | centrifuge | CEN | 64714.2022 | 75942.7353 | 9 | 4.5 | 8438.0817 | 16876.1634 | 53 | 2.9444 | 3429852.7189 | 447218.3298 | 894436.6596 | 190547.3733 | 24845.4628 | 49690.9255 | 50%
| c | microscope | MIC | 207085.4472 | 243016.7528 | 10 | 5 | 24301.6753 | 48603.3506 | 41 | 2.2778 | 8490503.3344 | 996368.6865 | 1992737.373 | 471694.6297 | 55353.8159 | 110707.6318 | 50%
2 | Cold chain equipment
| a | refrigerator | CLD | 173469.5219 | 203568.1429 | 10 | 5 | 20356.8143 | 40713.6286 | 22 | 1.2222 | 3816329.4807 | 447849.9145 | 895699.8289 | 212018.3045 | 24880.5508 | 49761.1016 | 50%
| b | freezer | CLD | 173732.7429 | 203877.0354 | 11 | 5.5 | 18534.2759 | 37068.5519 | 22 | 1.2222 | 3822120.3448 | 407754.0708 | 815508.1416 | 212340.0192 | 22653.0039 | 45306.0079 | 50%
3 | Surgical equipment
| a | anaesthesia unit | ANS | 1064482.592 | 1249180.5023 | 10 | 5 | 124918.0502 | 249836.1005 | 4 | 0.22222 | 4257930.368 | 499672.2009 | 999344.4019 | 236551.6871 | 27759.5667 | 55519.1334 | 50%
| b | electrosurgical unit | ESU | 362399.5326 | 425279.3174 | 9 | 4.5 | 47253.2575 | 94506.515 | 13 | 0.72222 | 4711193.9234 | 614292.3474 | 1228584.6948 | 261732.9957 | 34127.3526 | 68254.7053 | 50%
| c | operating light | LGHT | 369662.4215 | 433802.387 | 7 | 3.5 | 61971.7696 | 123943.5391 | 19 | 1.0556 | 7023586.0076 | 1177463.6218 | 2354927.2436 | 390199.2226 | 65414.6457 | 130829.2913 | 50%
| d | oxygen concentrator | OXG | 116485.564 | 136696.9235 | 8 | 4 | 17087.1154 | 34174.2309 | 60 | 3.3333 | 6989133.8423 | 1025226.9259 | 2050453.8518 | 388285.2135 | 56957.0514 | 113914.1029 | 50%
| e | strelizing unit | STR | 253752.5902 | 297781.0915 | 12 | 6 | 24815.091 | 49630.1819 | 21 | 1.1667 | 5328804.3943 | 521116.9101 | 1042233.8201 | 296044.6886 | 28950.9394 | 57901.8789 | 50%
| f | operating table | TAB | 869019.2873 | 1019802.4448 | 15 | 7.5 | 67986.8297 | 135973.6593 | 18 | 1 | 15642347.1709 | 1223762.9338 | 2447525.8676 | 869019.2873 | 67986.8297 | 135973.6593 | 50%
4 | Monitoring
| a | ecg | ECG | 320982.4431 | 376675.9669 | 8 | 4 | 47084.4959 | 94168.9917 | 21 | 1.1667 | 6740631.3057 | 988774.413 | 1977548.826 | 374479.517 | 54931.9118 | 109863.8237 | 50%
| b | monitor | MON | 841284.6292 | 987255.5583 | 9 | 4.5 | 109695.062 | 219390.1241 | 17 | 0.94444 | 14301838.6959 | 1864816.0545 | 3729632.1091 | 794546.5942 | 103600.8919 | 207201.7838 | 50%
5 | Ultrasound imaging |  | ULS | 1216627.0022 | 1427723.4227 | 10 | 5 | 142772.3423 | 285544.6845 | 50 | 2.7778 | 60831350.1092 | 7138617.1137 | 14277234.2275 | 3379519.4505 | 396589.8397 | 793179.6793 | 50%
6 | Xray imaging |  | XRAY | 3519053.3761 | 4129642.7926 | 8 | 4 | 516205.3491 | 1032410.6982 | 47 | 2.6111 | 165395508.6753 | 24261651.4068 | 48523302.8136 | 9188639.3709 | 1347869.5226 | 2695739.0452 | 50%
[10 empty] | 419 | 240 | 320130175.4557 | 43077410.7721 | 86154821.5443 | 17785009.7475 | 2393189.4873 | 4786378.9747
SOURCE |  |  |  | A_value estimate |  | C_lifecycle
[9 empty] | Spec BoQ | 223 | Special Sel: | 254236891.2277 | 35295455.7925 | 70590911.5849 | Special Sel: | 1960858.6551 | 3921717.3103
Note on K23: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on M27: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Share of Devices that are in District Hospitals of the Device Types Considered
Note on K52: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K77: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K102: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K127: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K152: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K177: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K202: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K227: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K252: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K277: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals
Note on K302: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Nur District Hospitals

## C_downtime_cost
Calculation: Downtime for District Hospitals
Assumed Downtime Share |  |  | 0.15 | Average annual CM downtime |  |  |  | Estimated downtime days based on date time | [5 empty] | Expected number of o\o days
M <= |  |  |  | at HF level, as reported in jobcards |  |  |  | passed between request and last inspection | [5 empty] | per year and per hospital |  |  |  | Price per | Price per |  |  | Average services lost per year and hospital | [5 empty] | Average returns lost per year and hospital (NPR, 2016)
|  |  |  | In hours |  | In days |  | # of CMs performed |  | # days since last visit |  | ass. share= | 0.15 | sh=36% | sh=2% | Cat2 | BoQ | Output | Output | Average services |  | DWNT 1 |  | DWNT 2 |  | DWNT 3 |  | DWNT 1 |  | DWNT 2 |  | DWNT 3
Equipment type: |  |  |  | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | BoQ | per HF | (NPR,16) | (I$PPP,16) | per day | per year | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN | w\o MTN | w MTN
1 | Laboratory equipment | [18 empty] | 33.6251 | 12273.1664 | 98.0662 | 51.9937 | 387.2151 | 200.0818 | 8591.2165 | 477.2898 | 9806.6229 | 5199.3671 | 38721.5136 | 20008.1762 | 859121.6457 | 47728.9803
| a | analyser | ANL | 52.8271 | 30.0833 | 2.2011 | 1.2535 | 0.42105 | 0.30769 | 236.9028 | 207.9533 | 14.9623 | 9.5978 | 80.3 | 4.4611 | 11 | 0.61111 | 100
| b | centrifuge | CEN | 82.0156 | 69.4464 | 3.4173 | 2.8936 | 0.52941 | 0.18919 | 147.8631 | 177.079 | 11.7421 | 5.0252 | 386.9 | 21.4944 | 53 | 2.9444 |  | 3.1251
| c | microscope | MIC | 75.1423 | 11.8021 | 3.1309 | 0.49175 | 0.29032 | 0.25397 | 180.0895 | 84.7361 | 7.8426 | 3.228 | 299.3 | 16.6278 | 41 | 2.2778
2 | Cold chain equipment
| a | refrigerator | CLD | [11 empty] | 8.9222 | 22 | 1.2222
| b | freezer | CLD | [11 empty] | 8.9222 | 22 | 1.2222
3 | Surgical equipment
| a | anaesthesia unit | ANS | [11 empty] | 1.6222 | 4 | 0.22222
| b | electrosurgical unit | ESU | [11 empty] | 5.2722 | 13 | 0.72222
| c | operating light | LGHT | [11 empty] | 7.7056 | 19 | 1.0556
| d | oxygen concentrator | OXG | [11 empty] | 24.3333 | 60 | 3.3333
| e | strelizing unit | STR | [11 empty] | 8.5167 | 21 | 1.1667
| f | operating table | TAB | [11 empty] | 7.3 | 18 | 1
4 | Monitoring
| a | ecg | ECG | 42.4656 | 12.225 | 1.7694 | 0.50938 | 0.21622 | 0.21951 | 159.2671 | 142.8099 | 5.1654 | 4.7023 | 153.3 | 8.5167 | 21 | 1.1667 | 200 | 6.2502 | 0.21963 | 80.1667 | 0.38862 | 0.11188 | 1.1345 | 1.0328 | 33.67 | 1.8706 | 77.7244 | 22.3753 | 226.9011 | 206.5567 | 6733.9999 | 374.1111
| b | monitor | MON | 49.75 | 21.1917 | 2.0729 | 0.88299 | 0.22581 | 0.15556 | 230 | 186.8507 | 7.7903 | 4.3598 | 124.1 | 6.8944 | 17 | 0.94444
5 | Ultrasound imaging |  | ULS | 61.3625 | 171.5349 | 2.5568 | 7.1473 | 0.035088 | 0.012346 | 180 | 150 | 0.94737 | 0.27778 | 365 | 20.2778 | 50 | 2.7778 | 430 | 13.438 | 5.8237 | 2125.6667 | 14.89 | 41.624 | 5.5172 | 1.6177 | 2125.6667 | 118.0926 | 6402.6917 | 17898.3082 | 2372.4096 | 695.614 | 914036.6993 | 50779.8166
6 | Xray imaging |  | XRAY | 110.5667 | 62.9346 | 4.6069 | 2.6223 | 0.61667 | 0.95775 | 244.7256 | 134.1074 | 22.6371 | 19.2661 | 343.1 | 19.0611 | 47 | 2.6111 | 262.5 | 8.2034 | 3.9269 | 1433.3333 | 18.0912 | 10.2975 | 88.8946 | 75.657 | 1347.3333 | 74.8518 | 4748.9414 | 2703.0991 | 23334.8366 | 19859.9519 | 353674.986 | 19648.6103
[6 empty] | 2.8222 | 2.2572 |  |  |  |  | 10.1553 | 6.6367 | 250.2857 | 11.3285 |  |  | 248.125 | [9 empty] | 21035.9804 | 25823.1497 | 64655.6609 | 40770.2988 | 2133567.3308 | 118531.5184
|  |  | Notes: | The time (h) between job request and job completed. This measure already reflects the differences in equipment-related BoQs, as it is directly informed by PLAMAHS |  |  |  | Time passed between a CM job request and the last time the MC looked at the device. We assume a that for 10% of the time elapsed the device was o\o, such that for any CM activity we obtain a number of downtime days. | [5 empty] | Assume the same overall share of operational or o\o devices for each device. Given the avg number of devices per hospital calculate the sum of operational days by multiplying share*365*#devices |  |  |  | B_service_prices |  |  | CM: Corrective Maintenance; DWNT: Downtime Scenario; MTN: Maintenance
Note on I4: Michael Schleicher: average number at the hospital level
Note on G5: Michael Schleicher: Average downtime in first ppm round
Note on H5: Michael Schleicher: average downtime in the last ppm rounds

## C_services_daily
Average Number of Serives per Day |  |  |  | M <=
|  |  | 2012/13 |  |  |  |  | 2013/14
|  |  |  | Services per day |  |  |  |  | Services per day
Region | District | PopTile | Population | Lab | ECG | Xray | US | Population | Lab | ECG | Xray | US
FW | bajura | 1 | 140226 | 18.189 | 0 | 1.9534 | 0.060274 | 142345 | 32.1343 | 0 | 1.4082 | 0.46027
FW | darchula | 1 | 135344 | 32.4219 | 0.4 | 7.8329 | 1.3973
MW | dolpa | 1 | 37962 | 4.5233 | 0 | 2.1397 | 0.48767 | 38002 | 4.4 | 0 | 4.6264 | 2.6438
MW | humla | 1 | 52857 | 4.6055 | 0 | 0 | 2.1362 | 52940 | 2.8986 | 0 | 0 | 2.0575
MW | kalikot | 1 | 148255 | 22.2384 | 0.36368 | 2.7205 | 0.83836 | 146335 | 30.2301 | 0.22375 | 0.49041 | 1.3863
MW | mugu | 1 | 57334 | 2.4575 | 0 | 0.73699 | 1.6247 | 59057 | 6.8219 | 0 | 1.3151 | 3.2082
FW | bajhang | 2 | 201278 | 33.148 | 0.13973 | 5.3973 | 6.6603 | 202733 | 41.0411 | 0.052055 | 6.8493 | 6.1205
FW | doti | 2 | 212575 | 35.6603 | 0 | 4.6164 | 4.2411 | 207263 | 47.1808 | 0.79452 | 6.7562 | 6.3753
MW | jajarkot | 2 | 179323 | 9.337 | 0 | 1.0904 | 0.019178 | 180621 | 10.3863 | 0.09863 | 2.1699 | 2.5699
MW | pyuthan | 2 | 238912 | 19.9178 | 0 | 3.0438 | 0 | 230853 | 28.7699 | 0 | 3.7589 | 0.61096
MW | rolpa | 2 | 229856 | 53.548 | 0.49589 | 8.0247 | 0.87123 | 224318 | 39.0603 | 0.12877 | 8.137 | 2.6192
MW | rukum | 2 | 214606 | 40.8657 | 0.18082 | 7.7945 | 0.78904 | 212982 | 35.3123 | 0.24384 | 7.2712 | 5.2658
FW | achham | 3 | 262459 | 60.5836 | 0 | 5.1014 | 1.4027 | 263382 | 61.9753 | 0 | 6.1233 | 1.1644
FW | baitadi | 3 | 254988 | 52.2301 | 0 | 5.5507 | 4.3973 | 254727 | 73.9617 | 0.0054795 | 5.2849 | 6.4
MW | bardia | 3 | 434300 | 28.5096 | 0 | 3.4685 | 0 | 435353 | 44.0685 | 0 | 3.5041 | 0.43288
MW | dailekh | 3 | 270431 | 54.6521 | 0.0082192 | 8.0301 | 5.7781 | 271195 | 53.7589 | 0.31781 | 7.9699 | 8.0548
MW | salyan | 3 | 248631 | 32.4959 | 0.10137 | 1.7479 | 3.6027 | 269711 | 35.5206 | 0.021918 | 2.8685 | 2.5918
Mean values, middle-sized HFs: |  |  |  | 32.0795 | 0.13607 | 4.9945 | 2.0968 |  | 33.6251 | 0.21963 | 5.8237 | 3.9269
Mean values, middle-sized HFs, per year: |  |  |  | 11709.0003 | 49.6667 | 1822.9999 | 765.3333 |  | 12273.1664 | 80.1667 | 2125.6667 | 1433.3333

## C_overhead_cost
Calculation of Overhead Cost (incl. NPV)
M <= |  |  |  |  | Original price obtained: |  | Effective price | [5 empty] | FC & MC Totals |  | Share | Cost
|  |  | share | quant | ref period | price tot | original | rep year | USD in repY | NPR in repY | NPR in 2016 | USD in 2016 | NPR in 2016 |  |  | In NPR 2016
Start-up cost
| IT-System |  | 0.7 | 1 | 3 years | 55000 | 38500 | 2014 | 51079.249 | 4982993.5728 | 5798062.5886 | 53993.8219 |  |  | 1 | 5798062.5886
| Car |  | 0.5 | 1 | 3 years | 30000 | 15000 | 2014 | 19901.0061 | 1941426.0673 | 2278282.0576 | 21216.2518 |  |  | 1 | 2278282.0576
| Rooms |  | 1 | 40 | 3 years | 5500 | 220000 | 2016 | 2048.7259 |  | 220000 | 2048.7259 |  |  | 1 | 220000
| RIA |  | 1 | 1 | 3 years | 50000 | 50000 | 2010 | 66221.2228 | 4851523.0002 | 8043070.2682 | 74900.2096 |  |  | 1 | 8043070.2682
Running cost | [12 empty] | 16339414.9143
Staff
|  | RTM | 1 | 2 | p.m. | 35000 | 840000 | 2016 | 7822.4079 |  |  | 7822.4079 |  |  | 1 | 840000
|  | PAM Engineer | 1 | 2 | p.m. | 35000 | 840000 | 2016 | 7822.4079 |  |  | 7822.4079 |  |  | 1 | 840000
|  | PAM Chief | 0.5 | 1 | p.m. | 70000 | 420000 | 2016 | 3911.2039 |  |  | 3911.2039 |  |  | 1 | 420000
| IT-System |  | 0.3 | 1 | 3 years | 55000 | 5500 | 2014 | 7297.0356 | 711856.2247 | 835370.0878 | 7779.2923 |  |  | 1 | 835370.0878
| Car |  | 0.5 | 1 | p.m. | 30000 | 180000 | 2014 | 1845.1288 | 180000 | 211231.7215 | 1967.0722 |  |  | 1 | 211231.7215
3146601.8093
| Consultant |  | 0.75 |  |  | 110000 | 82500 | 2014 | 109455.5336 | 10677843.3702 | 12530551.3167 | 116689.3847 |  |  | 1 | 10677843.3702
[13 empty] | 116689.3847 |  | Notes: Assumed share of overhead cost occuring if only the hospitals considered were maintained (district hospitals.
[11 empty] | w\o consult | 181461.3933
[11 empty] | w. consult | 298150.778
Notes: FC = Fixed Cost; OC = Operating Cost. Share: Assumed share of sverhead cost that would occur if only devices considered were maintained.
Overhead cost NPV calculation
#REF! |  |  | NPV (1 Yrs) | NPV (2 Yrs) | NPV (3 Yrs) | NPV (4 Yrs) | NPV (5 Yrs) | NPV (6 Yrs) | NPV (7 Yrs) | NPV (8 Yrs) | NPV (9 Yrs) | NPV (10 Yrs) | 15yrs | 20yrs | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Start-up cost
| IT-System |  | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 5798062.5886 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Car |  | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 2278282.0576 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| Rooms |  | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 220000 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| RIA |  | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 8043070.2682 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Start-up cost total (incl. mobilization)
|  |  | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 21448431.8611 | 15yrs: 21448431.8611; 20yrs: 21448431.8611
Running cost
Staff
|  | RTM | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000
|  | PAM Engineer | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000 | 840000
|  | PAM Chief | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 420000 | 420000 | 420000 | 420000 | 420000 | 420000 | 420000 | 420000 | 420000 | 420000
| IT-System |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878 | 835370.0878
| Car |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 0 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215 | 211231.7215
Running cost total
|  |  | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | 15yrs: #REF!; 20yrs: #REF!
Note on K3: Michael Schleicher: the year to which the reported value refers to
Note on N3: Michael Schleicher: inflation rate over reference and reported year
Note on O3: Michael Schleicher: value in €uro, based on reported year
Note on P3: Michael Schleicher: inflation rate over reference and reported year
Note on G7: Michael Schleicher: in square foot
Note on B15: Michael Schleicher: Looking at the mileage counter, you can estimate for the project car about 30,000 km / year for supervision and project monitoring, a good part of this on rough roads.
Note on P18: Michael Schleicher: 110,000 Euro are stated to account for PLAMAHS development and training, of which I deducted 25 p.p because of overlaps with existing PAM structure
Note on B38: Michael Schleicher: Looking at the mileage counter, you can estimate for the project car about 30,000 km / year for supervision and project monitoring, a good part of this on rough roads.

## D_invoiced_cost
Maintenance cost calculation |  |  | M <=
Original program cost calculation (accum over all facilities in a PPM round) | [8 empty] | Plamahs maintenance cost (accum. Over all cat2 facilities in a PPM round)
in Million NPR |  | Program |  | PPM abz | Invoiced | Spares |  |  | Maintenance | Total CM | Invoiced CM | MTN fees as | Approx. PM | CM amount
PPM |  | Total | PPM total | sanction | spread | Invoiced | Total |  | Round | amount | amount | per contract | fees | + PM fees
0 | MP | 22.1 | 0 | 0 |  | 0 | 0 |  | 1 | 30166 | 0 | 4769100 | 850500 | 880666
1 |  | 9.8059 | 9.8059 | 9.7776 | 0.7112 | 0 | 0.032813 |  | 2 | 121132 | 64771 | 4703100 | 1066140 | 1187272
2 |  | 9.8059 | 9.8059 | 9.7776 | 0.7112 | 0 | 0.032813 |  | 3 | 205745 | 117595 | 4771100 | 1162980 | 1368725
3 |  | 11.0366 | 10.4522 | 10.4239 | 0.7112 | 0.58438 | 1.0267 |  | 4 | 618753.6 | 524038 | 5055100 | 1383600 | 2002354
4 |  | 14.5825 | 13.6665 | 13.6382 | 0.7112 | 0.91599 | 1.2997 |  | 5 | 315092.8 | 201140 | 5137100 | 1414920 | 1730013
5 |  | 16.4202 | 15.1464 | 15.1181 | 0.7112 | 1.2738 | 1.591 |  | 6 | 167004.9 | 134325 | 5057100 | 1449360 | 1616365
6 |  | 14.2121 | 12.7191 | 12.6908 | 0.7112 | 1.493 | 1.9169 |  | Notes: CM = Corrective Maintenance; PM = Preventive Maintenance; MTN = Maitenance
N.A. | Sanction | -0.16991 | 0 | 0 |  | 0 | 0
|  | 97.7934 |  | 71.4262 |  | 4.2672 | after 3 years
|  | 1068703.2757 |  | 86.4256 |  | 5.1206 | betw year 4 and 6
|  |  |  | 115.2342 |  | 6.8275 | betw year 7 and 10
Notes: PPM = Preventive Maintenance Round
Break-down cost for total Bill of Quantity (Illustration Purposes Only)
Year
1 | share | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
| PPM | 22.1 | 19.5551 | 24.0621 | 27.8089 | 28.8085 | 28.8085 | 28.8085 | 28.8085 | 28.8085 | 28.8085 | 28.8085
| Spares inv | 0 | 1.4224 | 1.4224 | 1.4224 | 1.7069 | 1.7069 | 1.7069 | 1.7069 | 1.7069 | 1.7069 | 1.7069
Spares tot |  | 0 | 0.065625 | 2.3264 | 3.5079 | 3.5079 | 3.5079 | 3.5079 | 3.5079 | 3.5079 | 3.5079 | 3.5079
Notes: Numbers in Million NPR
Break-down cost for main Bill of Quantity selectionion (66%)
Year
0.66166 | share | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
Program cost data (Basis for Maintenance Cost Calculation Table 2, Column 2)
| PPM | 14.6227 | 12.9389 | 15.9209 | 18.4001 | 19.0615 | 19.0615 | 19.0615 | 19.0615 | 19.0615 | 19.0615 | 19.0615
| Spares inv | 0 | 0.94115 | 0.94115 | 0.94115 | 1.1294 | 1.1294 | 1.1294 | 1.1294 | 1.1294 | 1.1294 | 1.1294
Plamahs cost data (Basis for Maintenance Ciost Calculation Table 2, column 3)
| Approx. PM | 0 | 1.9166 | 2.5466 | 2.8643 |  |  | 4.2804
| Total CM | 0 | 0.1513 | 0.8245 | 0.4821 | 0.67275 | 0.86341 | 1.0539
Notes: Numbers in Million NPR
Note on A29: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Share of the number of devices in of the types we consider in our analysis divided by the total number of devices in the maintenance scheme
Note on B31: Michael Schleicher: so far not nicely calculated per HF
Note on G35: Michael Schleicher: assume increase of 2.19 between end of t3 and t6. Then split evenly into three steps

## D_device_values
M <=
| estimates of installed value for replacement equipment of good quality, 2014 price level | [13 empty] | Original Excel: | RIA final v6.xlsx
[10 empty] | 1 | Laboratory equipment |  |  |  | Original Sheet: | value estimate
|  |  |  | quantity | cost in 1.000 Euro | [5 empty] | a | analyser | ANL | 6.5667
|  |  | 2014 RIA: Equipment List in E, C, W | found | unit | total |  |  |  |  | b | centrifuge | CEN | 0.5
| Imaging | 1 | X-Ray unit, stationary 100mA | 9 | 22 | 198 | 111 | 0.081081 | 1.7838 |  | c | microscope | MIC | 1.6
|  | 2 | X-Ray unit, stationary 300mA | 14 | 34 | 476 |  | 0.12613 | 4.2883 | 2 | Cold chain equipment
|  | 3 | X-Ray unit, stationary 500mA | 19 | 50 | 950 |  | 0.17117 | 8.5586 |  | a | refrigerator | CLD | 1.3403
|  | 4 | X-Ray unit, C-arm, image intensifier | 5 | 68 | 340 |  | 0.045045 | 3.0631 |  | b | freezer | CLD | 1.3423
|  | 5 | X-Ray unit, mobile | 61 | 16 | 976 |  | 0.54955 | 8.7928 | 3 | Surgical equipment
|  | 6 | X-Ray unit, dental, panoramic | 3 | 26 | 78 |  | 0.027027 | 0.7027 |  | a | anaesthesia unit | ANS | 8.2245
|  | 7 | X-Ray unit, dental | 10 | 5.5 | 55 |  |  |  |  | b | electrosurgical unit | ESU | 2.8
|  | 8 | X-Ray film processor, automatic | 12 | 5 | 60 |  |  |  |  | c | operating light | LGHT | 2.8561
|  | 9 | X-Ray film dryer, automatic | 1 | 1.5 | 1.5 |  |  |  |  | d | oxygen concentrator | OXG | 0.9
|  | 10 | Scanner, ultrasonic, portable | 3 | 5.5 | 16.5 | 65 | 0.046154 | 0.25385 |  | e | strelizing unit | STR | 1.9606
|  | 11 | Scanner, ultrasonic, diagnostic | 53 | 8.5 | 450.5 |  | 0.81538 | 6.9308 |  | f | operating table | TAB | 6.7143
|  | 12 | Scanner, ultrasonic, Doppler | 9 | 16 | 144 |  | 0.13846 | 2.2154 | 4 | Monitoring
|  | 13 | Printer for ultrasound scanner | 63 | 0.3 | 18.9 |  |  |  |  | a | ecg | ECG | 2.48
| Diagnostic | 14 | Electrocardiograph, multi channel | 76 | 2.8 | 212.8 |  |  |  |  | b | monitor | MON | 6.5
|  | 15 | Electrocardiograph, single channel | 6 | 1.2 | 7.2 |  |  |  | 5 | Ultrasound imaging |  | ULS | 9.4
|  | 16 | Electroencephalograph | 2 | 3 | 6 |  |  |  | 6 | Xray imaging |  | XRAY | 27.1892
|  | 17 | ENT Treatment Unit | 3 | 8.5 | 25.5
|  | 18 | Endoscope ENT, with Monitor | 2 | 5 | 10
|  | 19 | Endoscope | 25 | 3.5 | 87.5
|  | 20 | Bronchoscope | 1 | 3.5 | 3.5
|  | 21 | Colonoscope | 3 | 4 | 12
|  | 22 | Laparoscope | 4 | 3.5 | 14
|  | 23 | Insufflator for laparoscope | 4 | 4 | 16
|  | 24 | Colposcope | 1 | 3.5 | 3.5
|  | 25 | Monitor, multiparameter, centralized | 1 | 8.5 | 8.5
|  | 26 | Monitor, multiparameters | 117 | 6.5 | 760.5
|  | 27 | Ultrasonic foetal heart monitor | 6 | 0.6 | 3.6
|  | 28 | Audiometer | 9 | 2.4 | 21.6
|  | 29 | Cardio Tocographic (CTG) Machine | 2 | 2.8 | 5.6
|  | 30 | Tympanometer | 2 | 1.6 | 3.2
|  | 31 | Pulmonary function test unit | 1 | 3.5 | 3.5
| Laboratory | 32 | Microscope, binocular | 164 | 1.6 | 262.4
|  | 33 | Centrifuge, electric | 134 | 0.5 | 67
|  | 34 | Centrifuge, haematocrit | 5 | 0.5 | 2.5
|  | 35 | Colorimeter | 94 | 0.6 | 56.4
|  | 36 | Analyser, biochemistry | 37 | 7.5 | 277.5 | 60 | 0.61667 | 4.625
|  | 37 | Analyser, blood gas | 2 | 7 | 14 |  | 0.033333 | 0.23333
|  | 38 | Analyser, electrolyte | 9 | 4 | 36 |  | 0.15 | 0.6
|  | 39 | Analyser, haematology | 7 | 4.5 | 31.5 |  | 0.11667 | 0.525
|  | 40 | CD4 Count Machine | 7 | 6 | 42
|  | 41 | Analyser, immunoassay | 3 | 7 | 21 |  | 0.05 | 0.35
|  | 42 | Analyzer, MTB/RIF | 2 | 7 | 14 |  | 0.033333 | 0.23333
|  | 43 | Haemoglobinometer | 2 | 2 | 4
|  | 44 | ELISA micro-plate reader | 2 | 3.5 | 7
|  | 45 | Coagulometer | 3 | 0.6 | 1.8
|  | 46 | Bilirubinometer | 1 | 0.4 | 0.4
|  | 47 | Laboratory incubator | 79 | 2.4 | 189.6
|  | 48 | Bio-safety cabinet | 2 | 5 | 10
|  | 49 | Drying oven | 51 | 1.2 | 61.2
|  | 50 | Shaker / mixer | 30 | 0.3 | 9
|  | 51 | Tissue processor | 3 | 2.5 | 7.5
| Monitoring | 52 | Infant incubator | 26 | 8.5 | 221
|  | 53 | Infant incubator, transport | 3 | 5.5 | 16.5
|  | 54 | Infant warmer, with sensor | 15 | 1 | 15
|  | 55 | Ventilator, intensive care | 23 | 9 | 207
|  | 56 | Ventilator, emergencies & transport | 8 | 7.5 | 60
|  | 57 | Infusion pump | 8 | 0.3 | 2.4
|  | 58 | Air compressor, ICU | 1 | 0.6 | 0.6
| Other | 59 | Suction apparatus, electric | 305 | 0.4 | 122
|  | 60 | Oxygen concentrator | 108 | 0.9 | 97.2
|  | 61 | Dental unit complete | 32 | 9 | 288
|  | 62 | Dental chair, electrical | 1 | 2.5 | 2.5
|  | 63 | Resin curing light, dental | 6 | 0.1 | 0.6
|  | 64 | Ultrasonic scaler | 9 | 0.1 | 0.9
|  | 65 | Sterilizing unit, dry heat | 26 | 0.7 | 18.2 | 71 | 0.3662 | 0.25634
|  | 66 | Sterilizing unit, non-vacuum, basic | 27 | 1 | 27 |  | 0.38028 | 0.38028
|  | 67 | Sterilizing unit, non-vacuum, automatic | 4 | 3 | 12 |  | 0.056338 | 0.16901
|  | 68 | Sterilizing unit, vacuum, basic | 5 | 2 | 10 |  | 0.070423 | 0.14085
|  | 69 | Sterilizing unit, vacuum, automatic | 9 | 8 | 72 |  | 0.12676 | 1.0141
|  | 70 | Dialysis machine | 10 | 9.5 | 95
|  | 71 | Dialysis water treatment unit | 1 | 10 | 10
|  | 72 | Lithotripter, intracorporeal, with camera | 1 | 16 | 16
|  | 73 | Diathermy unit, short wave | 5 | 2 | 10
|  | 74 | Electro Convulsive Therapy (ECT) | 2 | 1.2 | 2.4
|  | 75 | Interferential therapy apparatus | 1 | 1.8 | 1.8
|  | 76 | Ultrasonic therapy unit | 5 | 1.4 | 7
|  | 77 | Muscle stimulator | 2 | 0.8 | 1.6
|  | 78 | Syringe pump | 12 | 0.2 | 2.4
|  | 79 | Digital traction unit | 1 | 1.8 | 1.8
|  | 80 | Lens edging machine, automatic | 1 | 0.3 | 0.3
| Theatre | 81 | Anaesthesia apparatus, Boyle's | 14 | 7 | 98 | 49 | 0.28571 | 2
|  | 82 | Anaesthesia apparatus, EMO | 4 | 6.5 | 26 |  | 0.081633 | 0.53061
|  | 83 | Anaesthesia apparatus, Ventilator | 31 | 9 | 279 |  | 0.63265 | 5.6939
|  | 84 | Defibrillator | 12 | 2.6 | 31.2
|  | 85 | Surgical Drill | 3 | 0.2 | 0.6
|  | 86 | Electrosurgical unit, general purpose | 60 | 2.8 | 168
|  | 87 | Electrosurgical unit, ophthalmology | 1 | 3 | 3
|  | 88 | Operating light, ceiling mounted | 37 | 6 | 222 | 139 | 0.26619 | 1.5971
|  | 89 | Operating light, mobile | 80 | 1.5 | 120 |  | 0.57554 | 0.86331
|  | 90 | Operating light, mobile, battery backup | 22 | 2.5 | 55 |  | 0.15827 | 0.39568
|  | 91 | Ophtalmic microscope | 6 | 3.5 | 21
|  | 92 | Operating microscope | 7 | 6.5 | 45.5
|  | 93 | Pulse oximeter, with monitor | 6 | 0.3 | 1.8
|  | 94 | Operating table, electric | 6 | 8.5 | 51 | 56 | 0.10714 | 0.91071
|  | 95 | Operating table, hydraulic | 50 | 6.5 | 325 |  | 0.89286 | 5.8036
| Cooling | 96 | Air conditioner, window type | 9 | 0.2 | 1.8
| Cold Chain | 97 | Air conditioning unit, split type | 129 | 0.4 | 51.6
|  | 98 | Cold room with cooling system | 20 | 28 | 560
|  | 99 | Freezer, electric, for cold chain | 308 | 1.3 | 400.4 | 312 | 0.98718 | 1.2833
|  | 100 | Freezer, kerosene/electric | 4 | 1.6 | 6.4 |  | 0.012821 | 0.020513
|  | 101 | Freezer, solar powered | 2 | 6 | 12 |  | 0.0064103 | 0.038462
|  | 102 | Refrigerator, electric, for cold chain | 316 | 1.3 | 410.8 | 365 | 0.86575 | 1.1255
|  | 103 | Refrigerator, kerosene/electric | 49 | 1.6 | 78.4 |  | 0.13425 | 0.21479
|  | 104 | Mortuary refrigerator, 6 body | 3 | 14 | 42
| Electricity | 105 | Voltage stabilizer, up to 2 kVA | 306 | 0.1 | 30.6
| Supply & | 106 | Voltage stabilizer, larger than 2 kVA | 81 | 0.5 | 40.5
| Protection | 107 | UPS, up to 2 kVA | 20 | 0.2 | 4
|  | 108 | UPS, larger than 2 kVA | 6 | 2 | 12
|  | 109 | Inverter/charger, up to 2 kVA | 38 | 0.3 | 11.4
|  | 110 | Inverter/charger, larger than 2 kVA | 5 | 2 | 10
|  | 111 | Generator set, up to 5.5 kVA | 96 | 2.5 | 240
|  | 112 | Generator set, 6 to 15 kVA | 30 | 7.5 | 225
|  | 113 | Generator set, 15 to 60 kVA | 33 | 12.5 | 412.5
|  | 114 | Generator set, 100 to 320 kVA | 4 | 30 | 120
|  |  | total items | 3619
|  |  | estimated ex-factory value | total | mill Euro | 11.5214
|  |  |  | insurance, freight, customs & clearing, administration, distribution, installation | 25% | 14.4017
|  |  |  | 134 | NPR/Euro | excl VAT
|  |  |  | estimated installed value for replacement equipment of good quality | mill NPR | 1929.8345
[5 empty] | Crore | 192.9834
[6 empty] | 1930 | target
| E & C & W |  | total items E & C & W | 3619
|  |  | of which Un-commisssioned | 128
|  |  | items to be considered for outsourcing | 3491
|  |  | estimates for Inventory as reported | E & C & W | MW & FW | total
|  |  | equipment quantities in the regions | 3490 | 1470 | 4960
inclusive of all suction apparatus but excluding un-commissioned items
| nation-wide |  | equipment value excl VAT Mill NPR | 1840 | 590 | 2430
243 Crore
|  |  | proposed annual budget for MC / year | excl VAT in Million NPR |  | 177
17.7 Crore
|  |  |  | as % of equipment value |  | 7%
| MW & FW |  | the exact cost of the MC contract 2011 - 2014 for MW & FW will be known by June 2014 after the final invoices have been verified; from the current information the cost will be equivalent to about 6% of equipment value
| Note 1: |  | a 10 % rate for outsourcing is well justified considering the logistic challenges of Nepal and is well within the internationally accepted range, which is up to 12 % for medical equipment
| Note 2: |  | added value: inventory up-dating is part of the MC package
Note on B2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Diese Daten kommen vom Consultant, daher schwer einzuschätzen

## D_service_prices
M <= |  |  |  |  | Original Excel: | Renumeration Rates per Service - Health Insurance.xlsx
|  | S.N. | Investigation | Rate in NPR | Original Sheet: | Tabelle 1
| X-Ray: | 3 | Plain X-Ray (10x12/14x17) | 150
|  | 4 | CR X-Ray Small (8x10/10x12) | 200
|  | 5 | X-Ray Medium (11x14) | 300
|  | 6 | X-Ray Large (14x17) | 400
|  | 7 | X-Ray Digital (double/three/four exposure) | 400/500/600
|  | 8 | Dental X-Ray | 100
| Ultrasound | 1 | USG A Scan (eye) and soft tissue - Neck, Breast etc. Scan | 360
|  | 2 | USG (abdomen/pelvis) Scan | 500
|  | 18 | USG guided FNAC | 850
|  | 66 | USG B Scan | 500
| Electrocardiogram | 9 | Electrocardiogram ECG | 200
| Colonoscopy | 17 | Colonoscopy | 2500
| Hematology | 1 | TC/DC/Hb | 35/35/35
|  | 2 | Absolute count (Eosinophil/Neutrophil/Lymphocyte/Monocyte/Basophil) per cell type | 50
|  | 3 | APTT (activated partial thromboplastine time) | 290
|  | 4 | CT/BT | 70
|  | 5 | Mean Corpuscular Hemoglobin (MCH) | 30
|  | 6 | Mean Corpuscular Hemoglobin Concentration (MCHC) | 30
|  | 7 | Mean Corpuscular Volume (MCV) | 30
|  | 8 | Packed Cell Volume (PCV) | 50
|  | 9 | Peripheral smear/RBC morphology comment | 100
|  | 10 | Platelet count | 50
|  | 11 | PT/INR | 180
|  | 12 | RBC count | 35
|  | 13 | Reticulocyte count | 125

## D_device_lifecycles
M <= |  |  | GIZ Life Expectancy | [14 empty] | LIC (from GIZ) |  |  | HIC |  | Original Excel: | Equipment list-PLAMAHS-10thOct2018-BR.xlsx
Category | Description | eqp_many | PQ_bad1 | PQ_bad2 | PQwell1 | PQ_well2 | GQ_bad1 | GQ_bad2 | GQ_well1 | GQ_well2 | LE2 | LE3 | P2 |  |  |  |  | w\o MTN | w. MTN | decrease | w. MTN | w\o MTN | Original Sheet: | lifecycle
AC | Air conditioner | YES | 3 | 3 | 5 | 7 | 5 | 6 | 10 | 12 | 10 | 10 |  | 1 | Laboratory equipment
ANS | Anaesthesia apparatus | YES | 2 | 5 | 5 | 10 | 5 | 10 | 10 | 15 | 10 | 10 |  |  | a | analyser | ANL |  |  |  | 7.5 | #REF!
ANL | Analyser | YES | [8 empty] | 8 | 7 |  |  | b | centrifuge | CEN | 6 | 10 | 0.4 | 9 | #REF!
ADM | Audiometer | NO | [8 empty] | 10 | 10 |  |  | c | microscope | MIC | 6.5 | 12 | 0.45833 | 10 | #REF!
BLN | Balance, electronic, laboratory | NO | [8 empty] | 7 | 7 |  | 2 | Cold chain equipment
BIR | Bilirubinometer | NO | [8 empty] | 8 | 10 |  |  | a | refrigerator | CLD | 6.5 | 10.5 | 0.38095 | 10 | #REF!
BSC | Bio-safety cabinet | NO | [8 empty] | - | - |  |  | b | freezer | CLD |  |  |  | 11 | #REF!
BRO | Bronchoscope | NO | [8 empty] | 8 | 7 |  | 3 | Surgical equipment
CTG | Cardio Tocographic (CTG) Machine | NO | [8 empty] | 8 | 6 |  |  | a | anaesthesia unit | ANS | 5 | 10 | 0.5 | 10 | #REF!
CAU | Cautery apparatus, eye clinic | NO | [8 empty] | 10 | - |  |  | b | electrosurgical unit | ESU |  |  |  | 9 | #REF!
CDF | CD4 Count Machine | NO | [8 empty] | - | - |  |  | c | operating light | LGHT |  |  |  | 7 | #REF!
CEN | Centrifuge | YES | 3 | 4 | 7 | 8 | 6 | 9 | 10 | 12 | 8 | 10 |  |  | d | oxygen concentrator | OXG |  |  |  | 8 | #REF!
COS | Change over switch, automatic | NO | [8 empty] | - | - |  |  | e | strelizing unit | STR | 6.5 | 8.5 | 0.23529 | 12 | #REF!
CGLM | Coagulometer | NO | [8 empty] | 5 | 5 |  |  | f | operating table | TAB |  |  |  | 15 | #REF!
CLDR | Cold room with complete cooling system | NO | [8 empty] | - |  |  | 4 | Monitoring
CLNS | Colonoscope | NO | [8 empty] | 5 | 5 |  |  | a | ecg | ECG |  |  |  | 8 | #REF!
CMR | Colorimeter | YES | [8 empty] | 12 | - |  |  | b | monitor | MON |  |  |  | 9 | #REF!
CLPS | Colposcope | NO | [8 empty] | 7 | 15 |  | 5 | Ultrasound imaging |  | ULS |  |  |  | 10 | #REF!
CMPR | Compressor, Air, ICU | NO | [8 empty] | 15 | 15 |  | 6 | Xray imaging |  | XRAY |  |  |  | 8 | #REF!
CNV | Converter, solar power | NO | [8 empty] | - | - | [5 empty] | #REF! |  | Reduction
CNT | Counter, blood cell, electronic | NO | [8 empty] | 10 | 7
DEF | Defibrillator | NO | [8 empty] | 8 | 7
DCH | Dental chair, electrical | NO | [8 empty] | 10 | -
DUC | Dental unit complete | NO | [8 empty] | 10 | 10
DLSM | Dialysis machine | NO | [8 empty] | 8 | 10
DLSW | Dialysis water treatment unit | NO | [8 empty] | - | -
DTU | Diathermy unit, short wave | NO | [8 empty] | 10 | 15
DRL | Drill, surgical | NO | [8 empty] | 10 | 10
ECT | Electro Convulsive Therapy (ECT) | NO | [8 empty] | - | 15
ECG | Electrocardiograph | YES | [8 empty] | 8 | 10
EEPG | Electroencephalograph | NO | [8 empty] | 10 | 10
ESU | Electrosurgical unit | YES | [8 empty] | 8 | 10
ELI | ELISA micro-plate reader | NO | [8 empty] | - | -
END | Endoscope | NO | [8 empty] | 10 | -
ENT | ENT Treatment Unit | NO | [8 empty] | 12 | 12
FOT | Fiber optic tube, colonoscope | NO | [8 empty] | - | -
FRZ | Freezer | YES | [8 empty] | 11 | 12
GNR | Generator | YES | 2 | 6 | 5 | 10 | 6 | 15 | 10 | 20 | 10 | -
HAEM | Haemoglobinometer | NO | [8 empty] | 10 | 10
INC | Incubator | YES | [8 empty] | 10 | 15
INS | Insufflator | NO | [8 empty] | 10 | 20
ITA | Interferential therapy apparatus | NO | [8 empty] | - | -
INV | Inverter | NO | [8 empty] | - | -
LAP | Laparoscope | NO | [8 empty] | 10 | 10
LEM | Lens edging machine, automatic, optician | NO | [8 empty] | - | 12
LGHT | Light | YES | [8 empty] | 10 | 7
LITH | Lithotripter, intracorporeal pneumatic, with camera | NO | [8 empty] | 5 | 10
MIC | Microscope | YES | 3 | 6 | 5 | 10 | 6 | 10 | 10 | 20 | 12 | 7
MIX | Mixer, Vortex | NO | [8 empty] | 10 | 8
MON | Monitor | YES | [8 empty] | 9 | -
OVN | Oven | YES | 2 | 6 | 5 | 8 | 6 | 10 | 10 | 15 | 10 | 8
OXM | Oximeter | YES | [8 empty] | 10 | 10
OXG | Oxygen concentrator | YES | [8 empty] | 8 | 10
PTM | Photometer, flame | NO | [8 empty] | 9 | 8
PIP | Pipette, digital | NO | [8 empty] | 10 | 10
PRN | Printer, thermal, Ultrasound scanner | NO | [8 empty] | 8 | 7
PFT | Pulmonary function test unit | NO | [8 empty] | 10 | 10
PMP | Pump | NO | [8 empty] | 5 | 10
REF | Refrigerator | YES | 3 | 5 | 4 | 8 | 2 | 10 | 10 | 17 | 12 | 10
REG | Regulator, charging, solar | NO | [8 empty] | 10 | -
ULS | Scaler, ultrasonic | YES | [8 empty] | 10 | 10 |  |  | 6.25
SHK | Shaker / mixer | YES | [8 empty] | 10 | 10
STB | Stabilizer | NO | [8 empty] | - | -
STR | Sterilizing unit | YES | 3 | 6 | 5 | 12 | 6 | 10 | 8 | 14 | 12 | 15
STM | Stimulator, muscle | NO | [8 empty] | 10 | 7
SUC | Suction apparatus, electric | NO | 1 | 3 | 5 | 7 | 5 | 8 | 10 | 15 | 10 | -
TAB | Table, operating | YES | [8 empty] | 15 | 15
TUU | Therapy unit, ultrasonic | NO | [8 empty] | 7 | -
TISP | Tissue processor | NO | [8 empty] | 10 | 10
TRCT | Traction unit, digital | NO | [8 empty] | 10 | 10
TYMP | Tympanometer | NO | [8 empty] | 10 | -
UPS | UPS | NO | [8 empty] | - | -
VNT | Ventilator | YES | [8 empty] | 10 | 7
WRM | Warmer, infant, with sensor | NO | [8 empty] | 10 | 10
WBH | Water bath heater and stirrer | NO | [8 empty] | 10 | 10
XRAY | X-Ray | YES | [8 empty] | 8 | 8
Note on L2: Author: ohiocea_medical-equipment-life-expectancy-list.pdf
Note on M2: Author: BEAG_Life Span of Biomedical Devices - Guidance Paper Final.pdf
Note on L40: Author: specific generator
Note on L54: Author: pulse=5y

## D_currency_conversion
M <=
Exchange Rates: Nepalese Rupee per USD, Euro per USD & Nepalese Consumer Price Index
| 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017
CPI# | 82.3343 | 91.4691 | 100 | 109.2271 | 119.5597 | 130.3681 | 141.2723 | 152.3889 | 165.7844 | 171.1375
ER NPR* | 69.7617 | 77.5734 | 73.2624 | 74.02 | 85.1967 | 92.9933 | 97.5542 | 102.4051 | 107.3838 | 104.5119
ER EUR* | 0.68267 | 0.71984 | 0.75504 | 0.71936 | 0.77829 | 0.75316 | 0.75373 | 0.90166 | 0.90404 | -
Notes: #Consumper price index Nepal from World Development Indicators. *Official exchange rate refers to the exchange rate determined by national authorities or to the rate determined in the legally sanctioned exchange market. It is calculated as an annual average based on monthly averages (local currency units per USD).
Annual Percent Change in Nepalese CPI
i |  | 0.11095 | 0.093265 | 0.092271 | 0.094598 | 0.090402 | 0.083642 | 0.078689 | 0.087903 | 0.032289
i+1 |  | 1.1109 | 1.0933 | 1.0923 | 1.0946 | 1.0904 | 1.0836 | 1.0787 | 1.0879 | 1.0323
Notes: The change in CPI (i) is calculated as: (CPI current year - CPI previous year)/CPI previous year
Exchange Rates: Nepalese Rupee per I$, Euro per I$
| 2008 | 2009 | 2010 | 2011 | 2012 | 2013 | 2014 | 2015 | 2016 | 2017
ER NPR* | 17.7704 | 20.4416 | 23.2667 | 25.2546 | 26.0044 | 26.7999 | 28.1527 | 29.9425 | 31.9988 | 31.235
ER EUR* | 0.80615 | 0.79442 | 0.79232 | 0.78042 | 0.77412 | 0.75549 | 0.75199 | 0.75523 | 0.72972 | -
Notes: *Official exchange rate refers to the exchange rate determined by national authorities or to the rate determined in the legally sanctioned exchange market. It is calculated as an annual average based on monthly averages (local currency units per international dollar).

## D_fullhouse
0-30 Beds Hospital: |  | Bio-Medical Engineer(BME)-0; Bio-Medical Equipment Technician(BMET)-1; Tools-NRS 200,000; Spares stock-NRS 350,000; Space-130 sq. ft.; Room-1 | [11 empty] | M <=
31-99 Beds Hospital: |  | Bio-Medical Engineer(BME)-1; Bio-Medical Equipment Technician(BMET)-2; Tools-NRS 250,000; Spares stock-NRS 400,000; Space-194 sq. ft.; Room-1
100-249 Beds Hospital: |  | Bio-Medical Engineer(BME)-2; Bio-Medical Equipment Technician(BMET)-3; Tools-NRS 300,000; Spares stock-NRS 450,000; Space-258 sq. ft.; Room-2
250 and more Beds Hospital: |  | Bio-Medical Engineer(BME)-3; Bio-Medical Equipment Technician(BMET)-4; Tools-NRS 350,000; Spares stock-NRS 500,000; Space-333 sq. ft.; Room-2
Requirements estimation
| SN | Facility | Facility type | Current BoQ (active + passive), before 2nd PPM | No. of Beds | BME | salary+misc(NRS) | BMET | salary+misc(NRS) | Electrician & Plumber | salary+misc (NRS) | Total maintenance staff salary per Year(NRS) | Tools(NRS) | Spares(NRS) | Cal/testing tools | Space(sq m) | Space(sq ft) | Construction cost (5500 per sq ft) | Rooms
x | 1 | Achham District Hospital | DH | 37 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 3 | Baitadi District Hospital | DH | 37 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 4 | Bajhang District Hospital | DH | 36 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 5 | Bajura District Hospital | DH | 58 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 7 | Dailekh District Hospital | DH | 64 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 8 | Darchula District Hospital | DH | 58 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 9 | Dolpa District Hospital | DH | 30 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 10 | Doti District Hospital | DH | 39 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 15 | Humla District Hospital | DH | 43 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 16 | Jajarkot District Hospital | DH | 42 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 18 | Kalikot District Hospital | DH | 34 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 26 | Mugu District Hospital | DH | 38 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 30 | Pyuthan District Hospital | DH | 52 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 33 | Rolpa District Hospital | DH | 45 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
x | 34 | Rukum District Hospital | DH | 35 | 15 |  |  | 1 | 50000 |  |  | 600000 | 200000 | 350000 | 200000 | 12 | 129.168 | 710424 | 1
[12 empty] | 9000000 | 3000000 | 5250000 | 3000000 |  |  | 10656360