MRM Optimisation

Multiple Reaction Monitoring (MRM) optimisation was performed using the automated function in TSQ Quantiva Tune 2.0 by direct infusion of individual standards (25-50µg.mL-1). The optimised collision energy and RF lens voltage for the optimised MRM transitions for each analyte are shown in the table below.

| Compound | Ion | Precursor (m/z) | Product (m/z) | Collision Energy (V) | RF Lens (V) |
| --- | --- | --- | --- | --- | --- |
| DL Homocysteine | Quan | 136.039 | 46.986 | 28.809 | 42.854 |
| DL Homocysteine | Qual | 136.039 | 55.986 | 17.736 | 42.854 |
| Homocysteine-d4 | Quan | 140.039 | 48.929 | 29.517 | 41.618 |
| Homocysteine-d4 | Qual | 140.039 | 58.986 | 18.949 | 41.618 |
| Methionine | Qual | 150.07 | 104.000 | 10.253 | 44.337 |
| Methionine | Quan | 150.07 | 133.000 | 10.253 | 44.337 |
| Methionine-(methyl 13C3) d3 | Qual | 154.283 | 107.986 | 10.253 | 43.843 |
| Methionine-(methyl 13C3) d3 | Quan | 154.283 | 137.000 | 10.253 | 43.843 |
| Pyridoxine | Quan | 170.117 | 134.040 | 20.618 | 58.180 |
| Pyridoxine | Qual | 170.117 | 152.040 | 12.427 | 58.180 |
| Pyridoxine-d2 | Quan | 172.07 | 136.058 | 21.933 | 55.213 |
| Pyridoxine-d2 | Qual | 172.07 | 154.058 | 13.691 | 55.213 |
| SAH | Qual | 385.191 | 88.000 | 37.556 | 74.000 |
| SAH | Quan | 385.191 | 136.040 | 20.011 | 74.000 |
| SAH-d4 | Qual | 389.217 | 92.000 | 39.528 | 71.034 |
| SAH-d4 | Quan | 389.217 | 138.071 | 19.961 | 71.034 |
| SAM | Qual | 399.23 | 136.04 | 26.18 | 75.483 |
| SAM | Quan | 399.23 | 250.111 | 14.904 | 75.483 |
| Folic Acid | Qual | 442.222 | 176.040 | 37.910 | 67.000 |
| Folic Acid | Quan | 442.222 | 295.040 | 15.966 | 67.000 |
| folic acid-(glutamic acid-13C5,15N) | Qual | 448.183 | 176.375 | 39.124 | 65.348 |
| folic acid-(glutamic acid-13C5,15N) | Quan | 448.183 | 295.129 | 18.697 | 65.348 |
| 5MTHF | Qual | 460.27 | 194.111 | 32.551 | 77.708 |
| 5MTHF | Quan | 460.27 | 313.111 | 19.303 | 77.708 |
| Vitamin B12 | Quan | 678.461 | 147.111 | 39.022 | 112.562 |
| Vitamin B12 | Qual | 678.461 | 359.111 | 22.539 | 112.562 |