[bookmark: _Toc97211901]Figure S6. Summary presentations of study sample sizes (n) and numbers of data pairs compared (N), accuracy root mean square (Arms), mean bias (SD) and limits of agreement (LoA) of pulse oximeters for the subgroup of non-Black, non-White ethnic groups
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Note: 
· The Chi2 test for heterogeneity in the primary analysis suggested a Q(df = 8) = 56.20, with P value < 0.0001. 
· Tau2 between the 3 studies = 0 (95% CI 0 to 2.34); Tau2 between the 9 comparisons = 0.23 (0.07 to 0.90).
· The estimated overall I2 for the primary analysis = 47.95%, of which about 0% is due to between-studies heterogeneity, and 47.95% due to within-study heterogeneity. 
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Study Pulse oximeters Sample (data pairs) Arms, % Mean bias (SD), % LoA, % Mean bias (95% ClI), %
Valbuena 2021 Unspecified 243 (243) 5.80 -0.30 (5.80) -11.80 to 11.30 — -0.30 (-1.07 to 0.47)
Valbuena 2021 Unspecified 238 (238) 6.50 0.80 (6.40) -12.10 to 13.70 — 0.80 (0.03 to 1.57)
Wiles 2021 Nellcor (GE B1x5 M/P) 34 (1241) 1.13 0.33 (1.09) -2.47 t0 1.80 = 0.33 (-0.04 to 0.70)
Feiner 2007 Masimo Radical (clip) 7 (154) 2.32 0.99 (2.10) -3.13to 5.11 —_— 0.99 (-0.70 to 2.68)
Feiner 2007 Nellcor N-595 (clip) 7 (154) 1.71 0.64 (1.59) -2.48t0 3.76 —_— 0.64 (-1.05 to 2.33)
Feiner 2007 Nonin 9700 (clip) 7 (154) 1.92 0.03 (1.92) -3.731t0 3.79 —_— 0.03 (-1.66 to 1.72)
Feiner 2007 Masimo Radical (adhesive) 7 (154) 1.99 -0.34 (1.96) -4.18 to 3.50 —_—— -0.34 (-2.03 to 1.35)
Feiner 2007 Nellcor N-595 (OxiMax, adhesive) 7 (154) 2.48 0.83 (2.34) -3.76 to 5.42 —_— 0.83 (-0.86 to 2.52)
Feiner 2007 Nonin 9700 (adhesive) 7 (154) 2.69 0.51 (2.64) -4.66 to 5.68 —_— 0.51 (-1.18 to 2.20)
Primary analysis 3 studies (9 comparisons) 522 (2646) 1.58 0.31 (1.55) -2.72 10 3.35 - 0.31 (0.09 to 0.54)
Sensitivity (RoB) 2 studies (7 comparisons) 41 (2165) 1.23 0.36 (1.18) -1.95 t0 2.66 — 0.36 (-0.24 to 0.95)
Sensitivity (Trans) 2 studies (8 comparisons) 488 (1405) 2.07 0.30 (2.04) -3.70 to 4.31 B — 0.30 (-0.80 to 1.40)
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