# Supplementary files

The impact of therapeutic-dose induced intestinal enrofloxacin concentrations in healthy pigs on fecal Escherichia coli populations

Joren De Smet1, Filip Boyen2, Siska Croubels1, Geertrui Rasschaert3, Freddy Haesebrouck2, Robin Temmerman1, Patrick De Backer1, Mathias Devreese1*

1Department of Pharmacology, Toxicology and Biochemistry, Faculty of Veterinary Medicine, Ghent University, Salisburylaan 133, 9820 Merelbeke, Belgium

2Department of Pathology, Bacteriology and Avian Diseases, Faculty of Veterinary Medicine, Ghent University, Salisburylaan 133, 9820 Merelbeke, Belgium

3Technology and Food Science Unit, Flanders Research Institute for Agriculture, Fisheries and Food, Brusselsesteenweg 370, 9090 Melle, Belgium

* Corresponding Author: Mathias Devreese e-mail: Mathias.Devreese@UGent.be tel: 32 (0)92647347

Experimental design

# Figure A: Flow-chart of the experimental setup for the enrofloxacin (ENR) administrations with the dose per kg of bodyweight (BW).

# Chromatographic separation and instrumentational parameters

| Table A1. Gradient elution scheme for chromatographic separation with mobile phases: 0.1% acetic acid in H2O (A) and acetonitrile (ACN) (B). Liquid chromatography (LC) was performed on a Surveyer ThermoFisher Scientific system (Breda, The Netherlands). Flow rate was 400 µL/mL. |  |
| --- | --- |
| Time (min) | Gradient |
| 0 – 4 | 85% A, 15%B |
| 4 – 4.5 | Linear to 10% A, 90%B |
| 4.5 – 7 | 10% A, 90% B |
| 7 – 7.5 | Linear to 85% A, 15% B |
| 7.5 – 10 | 85% A, 15% B |

| Table A2. Liquid chromatography (LC) and mass spectrometry (MS) instrumentational parameters obtained after direct infusion of enrofloxacin (ENR) and internal standard (IS) ENR-d5 working solutions of 0.01 µg/mL in an H2O/acetonitrile (ACN) mixture (80/20 v/v). |  |
| --- | --- |
| LC or MS parameter | Optimized value |
| Column oven temperature | 45°C |
| Autosampler tray temperature | 5°C |
| Spray voltage | 3,800 V |
| Vaporizer temperature | 300°C |
| Sheath gas pressure | 33 au (arbitrary units) |
| Ion sweep gas pressure | 2.0 au |
| Auxiliary gas pressure | 15 au |
| Capillary temperature | 300°C |
| Collision pressure | 1.5 mTorr |

Validation parameters LC-MS/MS method

| Table B1. Results of major validation parameters: linearity (with correlation coefficient r and goodness-of-fit g), limit of quantification (LOQ, n=6, mean value given ± standard deviation (SD)), limit of detection (LOD, n=6, mean value given ± SD) and signal suppression or enhancement (SSE, n=6, mean value given ± SD) for enrofloxacin (ENR) with calibrator range in feces: 0.050 – 25 µg/g and plasma: 0.035 – 25 µg/mL. Experimental samples always fitted within range. |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Matrix | r | g | LOQ (µg/g or µg/mL) | LOD (µg/g or µg/mL) | SSE (%) |
| Feces | 0.9983 | 5.07 | 0.050 ± 0.00096 | 0.0057 | 52.65 ± 4.61 |
| Plasma | 0.9978 | 5.59 | 0.035 ± 0.0027 | 0.0020 | 66.25 ± 5.83 |

| Table B2. Validation results for the parameters accuracy and precision based on different theoretical concentration levels. |  |  |  |  |
| --- | --- | --- | --- | --- |
| Matrix | Theoretical concentration (µg/mL or µg/g) | Mean concentration ± SD (µg/mL of µg/g) | Precision RSD (%) | Accuracy (%) |
| Faeces | 0.050a | 0.0508 ± 0.000964 | 1.9 | 1.7 |
|  | 0.250a | 0.251 ± 0.00537 | 2.1 | 0.4 |
|  | 25.00a | 25.142 ± 1.0778 | 4.3 | 0.6 |
|  | 0.250b | 0.249 ± 0.00529 | 2.1 | -0.4 |
|  | 25.00b | 25.474 ± 1.163 | 4.6 | 1.9 |
| Plasma | 0.035a | 0.0363 ± 0.00271 | 7.5 | 3.7 |
|  | 0.250a | 0.251 ± 0.00545 | 2.2 | 0.2 |
|  | 25.00a | 25.910 ± 1.228 | 4.7 | 3.6 |
|  | 0.250b | 0.243 ± 0.0135 | 5.5 | -2.6 |
|  | 25.00b | 25.925 ± 0.690 | 2.7 | 3.7 |
| a Within-run accuracy and precision (n=6) b Between-run accuracy and precision (n=6) SD: standard deviation; RSD: relative standard deviation. Acceptance criteria: accuracy: > 0.01 µg/mL or µg/g: −20% to +10%; within-run precision: RSDmax: ≥ 0.01 µg/mL or µg/g and < 0.1 µg/g or µg/mL: 15%, ≥ 0.1 µg/mL or µg/g: 10%; between-run precision: RSDmax = 2(1-0.5logC) (with C (10-9) the concentration at which the samples were fortified). |  |  |  |  |