Supplementary Table 1. In vitro aerosol characterisation

| Device | Volumetric median diameter | Geometric standard deviation | Delivered dose |
| --- | --- | --- | --- |
| Phase 1: device comparison |  |  |  |
| A-VMN* | 2.82 μm | 1.75 | 26.99% |
| JN | 2.88 μm | 1.95 | 9.88% |
| pMDI | 2.36 μm | 1.77 | 14.23% |

*Aerosol characteristics of A-VMN were used for both B1 and B2 simulations to isolate the effect of the ventilation flow profile

A-VMN, Aerogen vibrating mesh nebuliser; JN, jet nebuliser; pMDI, pressurised metered dose inhaler

Supplementary Fig. 1. Flow profile of B1, a sinusoidal (VCV-like) waveform and B2, a rectangular pressure (PCV-like) waveform

PCV, pressure-controlled ventilation; VCV, volume-controlled ventilation

Supplementary Table 2. Numerical deposition values

| A. Device comparison |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Region / metric (% or ratio) | A-VMN |  | JN |  | pMDI |
| Right upper lobe deposition | 3.5 |  | 1.8 |  | 1.8 |
| Right middle lobe deposition | 1.8 |  | 1.0 |  | 1.0 |
| Right lower lobe deposition | 7.4 |  | 3.5 |  | 4.5 |
| Left upper lobe deposition | 4.5 |  | 1.6 |  | 2.4 |
| Left lower lobe deposition | 6.0 |  | 2.3 |  | 3.2 |
| Left lung central deposition | 0.8 |  | 0.4 |  | 0.4 |
| Left lung peripheral deposition | 9.8 |  | 3.4 |  | 5.3 |
| Right lung central deposition | 1.1 |  | 0.5 |  | 0.5 |
| Right lung peripheral deposition | 11.7 |  | 4.7 |  | 6.2 |
| Thoracic airway deposition | 5.3 |  | 2.3 |  | 2.1 |
| Alveolar tissue deposition | 18.8 |  | 6.5 |  | 10.6 |
| Exhaled fraction | 2.8 |  | 1.0 |  | 1.5 |
| Left lung C/P ratio | 0.08 |  | 0.11 |  | 0.07 |
| Right lung C/P ratio | 0.09 |  | 0.11 |  | 0.09 |
| B. Flow profile comparison |  |  |  |  |  |
| Region / metric (% or ratio) |  | B1 |  | B2 |  |
| Right upper lobe deposition |  | 3.6 |  | 2.3 |  |
| Right middle lobe deposition |  | 1.8 |  | 1.4 |  |
| Right lower lobe deposition |  | 7.4 |  | 4.9 |  |
| Left upper lobe deposition |  | 4.5 |  | 3.2 |  |
| Left lower lobe deposition |  | 6.0 |  | 4.0 |  |
| Left lung central deposition |  | 0.8 |  | 0.5 |  |
| Left lung peripheral deposition |  | 9.8 |  | 6.7 |  |
| Right lung central deposition |  | 1.1 |  | 0.8 |  |
| Right lung peripheral deposition |  | 11.7 |  | 7.9 |  |
| Thoracic airway deposition |  | 5.3 |  | 4.1 |  |
| Alveolar tissue deposition |  | 18.8 |  | 12.1 |  |
| Exhaled fraction |  | 2.8 |  | 10.4 |  |
| Left lung C/P ratio |  | 0.08 |  | 0.08 |  |
| Right lung C/P ratio |  | 0.09 |  | 0.10 |  |

A-VMN, Aerogen vibrating mesh nebuliser; B1, sinusoidal flow waveform; B2, rectangular pressure-driven flow waveform; C/P, central-to-peripheral; JN, jet nebuliser; pMDI, pressurised metered dose inhaler