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	No.
	Task
	Errors Observed

	1.
	Assembles necessary materials
	- Forgot materials (lancet, pipette, buffer, swab, cotton)

	2.
	Read RDT expiration date
	- Did not check expiration date

	3.
	Remove contents of test packet 
	- Opened packet with teeth

	4.
	Write patient's name on cassette 
	- Did not write patient’s name on cassette

	5.
	Identify patient’s details and date on the RDT cassette
	- Did not write date on cassette

	6.
	Explain procedure to patient
	- Did not specify that they were testing for malaria
- Did not explain procedure to patient


	7.
	Wear gloves
	- Attempted to conduct RDT without wearing gloves

	8.
	Select 4th finger from the thumb of the left hand for blood collection 
	- Selected the wrong finger
- Selected the wrong hand

	9.
	Clean finger with alcohol swab and allow it to dry
	- Excessive wiping
- Cleaned with cotton instead of alcohol

	10.
	Prick finger firmly with sterile lancet
	- Placed lancet down after uncapping

	11.
	Discard lancet in sharps bin immediately after pricking finger
	- Placed used lancet on mat
- Placed used lancet in waste bag (non-sharps waste container)

	12.
	Do not squeeze finger excessively
	- Squeezed finger excessively to obtain blood sample

	13.
	Collect an adequate volume of blood with pipette
	- Did not wipe away first drop of blood
- Collected too little blood
- Collected too much blood

	14.
	Dispense blood in correct well
	- Dispensed collected blood in buffer well

	15.
	Discards the pipette in the sharps box
	- Placed used pipette in waste bag (non-sharps waste container)

	16.
	Dispose of gloves & cotton wool in non-sharps container
	- Did not remove used gloves
- Unintentionally smeared blood on other item (buffer, pen, phone, mat)

	17.
	Dispense correct volume of buffer
	- Dispensed too much buffer
- Missed buffer well while trying to dispense

	18.
	Wait for twenty (20) minutes
	- Attempted to read results too soon
- Did not set timer

	19.
	Read results correctly
	- Misinterpreted test result (faint-positive)

	20.
	Verify internal test control
	- Did not verify control strip







