Supplement table 3: Diagnostic performance of index tests to detect Plasmodium infection among parturient women using placental blood microscopy as reference in Majang Zone of Gambella Region, Southwest Ethiopia, November 2022-March 2023 (n= 180).

| Index test methods |  | Placental blood microscopy from parturient women as a reference test method (PM= 180) |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | TP | FP | FN | TN | Sensitivity (95% CI) | Specificity (95% CI) | PPV (95% CI) | NPV (95% CI) | Accuracy (95% CI) | Kappa (95% CI) | AUC (95% CI) | p |
| 1. RDT | Peripheral blood | 26 | 6 | 6 | 142 | 81.3% (75.5-86.9) | 95.9% (93.1-98.8) | 81.3% (75.5-86.9) | 95.9% (93.1-98.8) | 93.3% (89.6-96.2) | 0.772 (0.528-0.830) | 0.863 (0.722-1.005) | < 0.001 |
|  | Placental blood | 25 | 4 | 7 | 144 | 78.1% (72.1-84.1) | 97.3% (94.9-99.6) | 86.2% (81.1-91.2) | 95.4% (92.3-98.5) | 93.9% (89.3-96.9) | 0.783 (0.659-0.906) | 0.825 (0.684-0.982) | < 0.001 |
|  | Cord blood | 5 | 0 | 27 | 148 | 15.6% (10.3-20.9) | 100.0% (100.0-100.0) | 100.0% (100.0-100.0) | 84.6% (79.3-89.8) | 85.0% (79.9-89.8) | 0.233 (0.062-0.403) | 0.538 (0.353-0.724) | 0.684 |
|  | Overall | 28 | 7 | 4 | 141 | 87.5% (82.6-92.3) | 95.3% (92.2-98.3) | 80.0% (74.1-85.8) | 97.2% (94.9-99.6) | 93.9% (89.3-96.6) | 0.798 (0.684-0.911) | 0.891 (0.769-1.013) | < 0.001 |
| 2. Microscopy | Peripheral blood | 31 | 7 | 1 | 141 | 96.9% (94.3-99.4) | 95.3% (92.2-98.3) | 81.6% (75.9-87.2) | 99.3% (98.1-100.5) | 95.6% (90.3-97.6) | 0.842 (0.742-0.942) | 0.940 (0.849-1.031) | < 0.001 |
|  | Cord blood | 7 | 0 | 25 | 148 | 21.9% (15.8-27.9) | 100.0% (100.0-100.0) | 100.0% (100.0-100.0) | 85.5% (80.4-90.7) | 86.1% (80.2-90.8) | 0.401 (0.191-0.610) | 0.538 (0.353-0.724) | 0.684 |
|  | Overall | 32 | 7 | 0 | 141 | 100.0% (100.0-100.0) | 95.3% (92.1-98.3) | 82.1% (76.5-87.6) | 100.0% (100.0-100.0) | 96.1% (92.1-98.4) | 0.877 (0.788-0.965) | 0.979 (0.944-1.014) | < 0.001 |
| 3. Microscopy & RDT Combined | Peripheral blood | 31 | 8 | 1 | 140 | 96.9% (94.3-99.4) | 94.6% (91.3-97.9) | 79.5% (73.6-85.4) | 99.3% (98.1-100.5) | 95.0% (90.7-97.2) | 0.842 (0.742-0.942) | 0.930 (0.836-1.023) | < 0.001 |
|  | Placental blood | 31 | 4 | 1 | 144 | 96.9% (94.3-99.4) | 97.3% (94.9-99.6) | 88.6% (83.9-93.2) | 99.3% (98.1-100.5) | 97.2% (96.3-99.1) | 0.908 (0.830-0.986) | 0.979 (0.944-1.014) | < 0.001 |
|  | Cord blood | 7 | 0 | 25 | 148 | 21.9% (15.8-27.9) | 100.0% (100.0-100.0) | 100.0% (100.0-100.0) | 85.5% (80.4-90.7) | 86.1% (80.2-90.8) | 0.315 (0.132-0.497) | 0.538 (0.353-0.724) | 0.684 |
| 4. qPCR | Peripheral blood | 12 | 5 | 1 | 42 | 92.3% (86.8-99.1) | 89.4% (84.5-95.2) | 70.6% (59.0- 82.1) | 97.7% (93.9-101.5) | 90.0% (79.5-96.2) | 0.735 (0.539-0.931) | 0.908 (0.809-1.007) | < 0.001 |
|  | Placental blood | 13 | 3 | 0 | 44 | 100.0% (100.0-100.0) | 93.6% (87.4-99.8) | 81.3% (71.3-91.1) | 100.0% (100.0-100.0) | 95.0% (86.1-98.9) | 0.864 (0.715-1.013) | 0.968 (0.925-1.011) | < 0.001 |
|  | Cord blood | 2 | 1 | 11 | 46 | 15.4% (6.2-24.5) | 97.9% (94.6-101.5) | 66.7% (54.7-78.6) | 80.7% (70.7-90.7) | 80.0% (67.6-89.2) | 0.184 (-0.080-0.448) | 0.566 (0.378-0.754) | 0.489 |
|  | Overall | 13 | 6 | 0 | 41 | 100.0% (100.0-100.0) | 87.2% (78.8-95.7) | 68.4% (56.6-80.2) | 100.0% (100.0-100.0) | 90.0% (79.5-96.2) | 0.748 (0.562-0.934) | 0.936 (0.876-0.997) | < 0.001 |
| 5. Histopathology of placental biopsy |  | 30 | 14 | 2 | 134 | 93.8% (90.2-97.3) | 90.5% (86.2-94.8) | 68.2% (61.4-74.9) | 98.5% (96.7-100.3) | 91.2% (85.9-94.8) | 0.735 (0.613-0.856) | 0.898 (0.796-0.999) | < 0.001 |