# Supplementary materials

Supplementary Table 1. Primer and probe sequences for species-specific qPCR assays

| Species | Primers and probes | Sequences (5’ -> 3’) |
| --- | --- | --- |
| P. falciparum | Pf18S forward primer | GTAATTGGAATGATAGGAATTTACAAGGT |
|  | Pf18S reverse primer | TCAACTACGAACGTTTTAACTGCAAC |
|  | Pf18S probe | FAM-AACAATTGGAGGGCAAG-MGBNFQ |
| P. vivax | Pv18S forward primer | AGCAGCCGCGGTAATTCCA |
|  | Pv18S reverse primer | ATGCGCACAAAGTCGATACGAAG |
|  | Pv18S probe | TEXASRED-AGCAACGCTTCTAGCTTAATCCACAT-BHQ2 |

Supplementary Figure 1. Standard curves for P. falciparum and P. vivax qPCR multiplex experiments with the equation and R-squared values

## Average Ct values for P. falciparum and P. vivax plasmids

Supplementary Table 2. Average Ct values of standards for both P. vivax and P. falciparum experiments

| Plasmid serial dilution (10^n) | P. vivax averages | P. falciparum averages |
| --- | --- | --- |
| 3 | 37.18 | 37.69 |
| 4 | 33.59 | 34.13 |
| 5 | 30.54 | 31.08 |
| 6 | 27.35 | 27.88 |
| 7 | 23.66 | 24.24 |

Summary of average slope and efficiency in qPCR analysis

The slope, efficiency, and R² values obtained from the average of all qPCR experiments are presented in the table below.

Supplementary Table 3. Slope, efficiency, and R² values averaged across all qPCR experiments

| Target | Average slope | Average efficiency (%) | R2 |
| --- | --- | --- | --- |
| Pf18S DNA | -3.32 | 100.3 | 0.9996 |
| Pv18S DNA | -3.33 | 99.7 | 0.9994 |

Supplementary Figure 2. ROC curves for the diagnostic performance of BIOCREDIT Ag Pf/Pv (pLDH/pLDH) RDT and microscopy against qPCR (reference standard)

The curves compare the BIOCREDIT Ag Pf/Pv (pLDH/pLDH) RDT (yellow) and facility microscopist (blue), with area under the curve values and 95% confidence intervals annotated. Panel (a) shows ROC curves for P. falciparum, and panel (b) for P. vivax.

Supplementary Table 4. A 2x2 contingency table for the BIOCREDIT Ag Pf/Pv (pLDH/pLDH) RDT and facility microscopy versus 18S rRNA qPCR (reference standard)

| Test | TP | FP | TN | FN | Total qPCR+ | Total qPCR- |
| --- | --- | --- | --- | --- | --- | --- |
| RDT vs qPCR (Any Plasmodium) | 437 | 30 | 683 | 50 | 487 | 713 |
| RDT vs qPCR (P. falciparum) | 296 | 24 | 833 | 47 | 343 | 857 |
| RDT vs qPCR (P. vivax) | 156 | 3 | 1017 | 24 | 180 | 1020 |
| Microscopy vs qPCR (Any Plasmodium) | 291 | 26 | 687 | 196 | 487 | 713 |
| Microscopy vs qPCR (P. falciparum) | 197 | 19 | 838 | 146 | 343 | 857 |
| Microscopy vs qPCR (P. vivax) | 115 | 10 | 1010 | 65 | 180 | 1020 |

TP: True positive, FP: False positive, TN: True negative, FN: False negative

Supplementary Table 5. Descriptive statistics of qPCR-detected parasite densities by species

| Species | N Positives | Median (copies/µL) | Mean (copies/µL) | Q1 | Q3 | Max (copies/µL) |
| --- | --- | --- | --- | --- | --- | --- |
| P. falciparum | 343 | 364 | 27,781 | 18 | 10,456 | 1,017,020 |
| P. vivax | 180 | 173 | 9,114 | 2 | 6,767 | 150,505 |
| Mixed infections | 36 | Pf: 302 | Pf: 8,245 | 48 | 1,856 | 126,974 |
|  |  | Pv: 165 | Pv: 4,872 | 32 | 1,234 | 51,945 |

Q1: 25th percentile; Q3: 75th percentile

Supplementary Table 6. Sensitivity of the BIOCREDIT Ag Pf/Pv (pLDH/pLDH) RDT and microscopy by parasite density as measured by copies/µL of the 18S rRNA qPCR, P. falciparum (n=343) and P. vivax (n=180)

| Species | Method | <1,000 copies/µL n, % (95% CI) | 1,000–4,999 copies/µL n, % (95% CI) | ≥5,000 copies/µL n % (95% CI) | Spearman's rho | Logistic OR |
| --- | --- | --- | --- | --- | --- | --- |
| P. falciparum | RDT | 144, 67.4 (59.1–74.9) | 53, 100 (93.3–100) | 146, 100 (97.5–100) | 0.564 (p<0.001) | 121.26 (31.21–695.78) |
| P. falciparum | Facility microscopy | 144, 78.5 (71.2–84.6) | 53, 96.6 (88.3–99.6) | 146, 99.2 (95.7–100) | 0.53 (p<0.001) | 10.77 (5.88–22.55) |
| P. vivax | RDT | 66, 68.2 (55.6–79.1) | 32, 98.4 (89.7–100) | 82, 100 (95.6–100) | 0.469 (p<0.001) | 6.84 (3.5–16.35) |
| P. vivax | Facility microscopy | 66, 72 (60.9–81.3) | 32, 90.6(75–98) | 82, 100 (91.6–100) | 0.446 (p<0.001) | 5.52 (3.01–12) |

n: Number of qPCR-confirmed positives in stratum; % (95% CI): Sensitivity percentage (95% confidence interval); Spearman's rho: Correlation between density and detection (p-value); Logistic OR: Odds ratio of detection per log₁₀ density increase (95% CI); RDT: BIOCREDIT Ag Pf/Pv (pLDH/pLDH) RDT.