## Sheet1
Phantom ID |  | Magnet Characteristics | [7 empty] | Sequence Characteristics | [6 empty] | CoV Per Field-Dependent Tube Without Temperature Adjustment∞
| Temp ± SD (°C) | Vendor | Tesla | Name | YOM | Software at Study Start [Upgraded to: Date] | Bore* (cm) | Gradient performance§ | Coil channels A¦P^ | Sequence Name | WIP# [Changed to WIP#: Date] | Count per session$ | Scan Frequency≅ | TE | NI (HR¶) | FA | I | H | G | F / M | E / L | D / K | C | B / J | A
15E001ae | 22.8 ± 0.4 | Siemens | 1.5 | Aera | 2015 | VE11 | 70 | 45/200 | 18 + 32 | MOLLI Myomaps Pre | Myomaps Product | 2 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.73969 | 1.3282 | 0.74109 | 2.1325 | 2.8739 | 1.7051 | 0.38182 | 2.2156 | 0.84552
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 2 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.18916 | 0.56038 | 0.11769 | 0.14191 | 0.79303 | 0.28938 | 0.22856 | 1.0065 | 0.23719
15E001in | 17.7 ± 1.2 | Philips | 1.5 | Ingenia | 2013 | R4.1.3SP2 | 70 | 33/200 | 16 ¦ 12 | MOLLI 3s(3s)5s | / | 2 | 4/12 | / | 896 (67) | 50 | 0.62391 | 1.4131 | 1.2807 | 1.0084 | 1.902 | 0.72048 | 0.90756 | 2.0172 | 0.46195
[10 empty] | MOLLI 3b(3s)3b(3s)5b | / | 2 | 4/12 | / | 896 (67) | 50 | 0.37975 | 0.45164 | 0.92218 | 0.38487 | 0.6171 | 0.28349 | 0.71554 | 0.67978 | 0.37249
[10 empty] | Sub-optimal MOLLI 5b(1b)1b(1b)1b in the absence of iterative/data dropping steps in map creation as per ShMOLLI | / | 2 | 4/12 | / | 896 (67) | 50 | Stubborn Artefact | Stubborn Artefact | 1.1311 | 0.48182 | 0.78527 | 0.22082 | Stubborn Artefact | 0.56628 | 0.71521
[10 empty] | SASHA | / | 2 | 4/12 | / | 896 (67) | 50 | 0.64775 | 0.72198 | 1.4905 | 0.44941 | 0.90909 | 0.27566 | 0.65767 | 0.71176 | 0.66343
15E002 | 19.5 ± 0.5 | Siemens | 1.5 | Aera | 2014 | VE11 | 70 | 45/200 | 18 ¦ 12 | MOLLI Myomaps Pre | Myomaps Product | 2 | 3/52 | 1.22 | 900 (66.7) | 35 | 0.34453 | 0.46815 | 0.73876 | 0.77132 | 0.80349 | 0.85658 | 0.49879 | 1.1188 | 0.56291
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 1 | 3/52 | 1.22 | 900 (66.7) | 35 | 0.19478 | 0.3741 | 0.54681 | 0.56252 | 0.58697 | 0.64866 | 2.4502 | 0.86823 | 0.66667
15E003 | 21.7 ± 1.4 | Siemens | 1.5 | Avanto | 2009 | VB17A | 60 | 45/200 | 08 ¦ 12 | MOLLI 5s(3s)3s | WIP448B | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.85652 | 0.35087 | 0.78774 | 0.4274 | 0.95761 | 0.19092 | 0.6738 | 1.1342 | 0.68018
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP448B | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.60149 | 0.36914 | 0.85159 | 0.49885 | 0.6936 | 0.10661 | 0.58007 | 0.92158 | 0.72566
15E005 | 22.0 ± 0.0 | Siemens | 1.5 | Avanto | 2007 | VB17 | 60 | 45/200 | 06 ¦ 12 | MOLLI 5b(3s)3b | WIP448 | 1 | 4/12 | 1.04 | 900 (66.7) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP448 | 1 | 4/12 | 1.04 | 900 (66.7) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
15E006 | 22.0 ± 0.0 | Siemens | 1.5 | Aera | 2014 | VD13 | 70 | 45/200 | 18 ¦ 32 | MOLLI 5s(3s)3s | WIP780B | 1 | 1/12 | 1.12 | 800 (75) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP780B | 1 | 1/12 | 1.12 | 800 (75) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
[10 empty] | SASHA Pre-contrast | WIP780B | 1 | 1/12 | 1.12 | 800 (75) | 70 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
[10 empty] | SASHA Post-contrast | WIP780B | 1 | 1/12 | 1.12 | 800 (75) | 70 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
15E007 | 21.5 ± 0.7 | Siemens | 1.5 | Espree | 2009 | VB17A | 70 | 33/170 | 16 ¦ 16 | MOLLI 5s(3s)3s | WIP448 | 1 | 4/12 | 1.17 | 1000 (60) | 35 | 0.43718 | 0.47659 | 0.56795 | 0.66521 | 1.4408 | 0.58476 | 0.98864 | 1.5304 | 0.53902
15E008 | 21.4 ± 0.9 | Philips | 1.5 | Ingenia | 2014 | R5.1.7SP2 | 70 | 33/200 | 16 ¦ 12 | MOLLI 5s(3s)3s | / | 1 | 3/52 | 1.20 | 896 (67) | 20 | 0.33649 | 0.55681 | 0.70979 | 1.1426 | 2.1591 | 1.7578 | 0.64186 | 3.157 | 1.0777
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 1 | 3/53 | 1.20 | 896 (67) | 20 | 0.4333 | 0.6171 | 0.64789 | 1.5755 | 2.4759 | 1.4146 | 0.3267 | Stubborn Artefact | 1.0777
[10 empty] | MOLLI 3s(3s)3s(3s)5s | / | 1 | 3/52 | 1.15 | 896 (67) | 20 | 0.32077 | 0.41347 | 0.89475 | 1.083 | 1.882 | 0.75242 | 0.32732 | 2.4418 | 0.8276
15E009 | 21.0 ± 2.1 | Philips | 1.5 | Ingenia | 2014 | R5.1.7SP2 | 70 | 33/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 2 | 2/52 | 0.84 | 896 (67) | 35 | 3.4183 | 3.4791 | 3.4869 | 3.811 | 4.3993 | 3.604 | 3.5276 | 4.571 | 3.4895
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 2 | 2/52 | 0.84 | 896 (67) | 35 | 1.4826 | 1.4556 | 2.0913 | 1.9208 | 2.6294 | 1.2407 | 1.3596 | 2.72 | 2.1014
15E010 | 21.6 ± 1.0 | Philips | 1.5 | Ingenia | 2012 | R5.2.0SP2 | 70 | 33/200 | 18 ¦ 12 | MOLLI 5s(3s)3s | CardiacQuant | 2 | 3/52 | 0.84 | 896 (67) | 35 | 0.6199 | 1.0754 | 0.80125 | 0.46184 | 0.61972 | 0.55221 | 1.264 | 0.60899 | 1.0547
[10 empty] | MOLLI 4s(1s)3s(1s)2s | CardiacQuant | 2 | 3/52 | 0.84 | 896 (67) | 35 | 0.52754 | 1.1256 | 0.89003 | 1.6514 | 2.5943 | 1.1307 | 0.41934 | 2.5773 | 0.77578
15E011 | 20.9 ± 1.0 | Siemens | 1.5 | Aera | 2012 | VD13A | 70 | 45/200 | 18 ¦ 12 | MOLLI 5b(3s)3b | WIP780B | 2 | 3/52 | 1.12 | 900 (66.7) | 35 | 0.88 | 0.86 | 0.76 | 0.86 | 1.1 | 0.97 | 1.02 | 1.18 | 1.01
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780B | 1 | 3/52 | 1.12 | 900 (66.7) | 35 | 0.58 | 0.34 | 0.87 | 0.82 | 1.33 | 0.76 | 0.56 | 1.55 | 1.11
[10 empty] | ShMOLLI | WIP780B | 2 | 3/52 | 1.07 | 900 (66.7) | 35 | 0.85304 | 0.36043 | 0.59364 | 2.4611 | 3.2462 | 1.0238 | 0.4108 | 3.4218 | 0.66935
15E013 | 22.6 ± 1.4 | Philips | 1.5 | Achieva | 2015 | R5.2.0SP2 | 60 | 40/200 | 06 ¦ 12 | MOLLI 5s(3s)3s | CardiacQuant | 1 | 2/52 | 1.08 | 896 (67) | 35 | 0.41101 | 0.58548 | 0.60048 | 1.0259 | 1.8377 | 0.70292 | 1.1937 | 2.2026 | 0.67222
[10 empty] | MOLLI 4s(1s)3s(1s)2s | CardiacQuant | 1 | 2/52 | 1.08 | 896 (67) | 35 | 0.37501 | 0.37233 | 0.80244 | 0.96776 | 1.4927 | 0.51432 | 0.25128 | 1.6669 | 1.078
15E015 | 20.9 ± 1.0 | Siemens | 1.5 | Avanto | 2008 | VB17A | 60 | 45/200 | 16 ¦ 16 | MOLLI 3b(3s)3b(3s)5b | WIP448 | 1 | 1/52 | 1.06 | 900 (66.7) | 35 | 0.54285 | 0.43604 | 0.55915 | 0.64256 | 0.78321 | 0.54741 | 0.4834 | 0.87215 | 0.75644
[10 empty] | ShMOLLI | / | 1 | 1/52 | 1.05 | 900 (66.7) | 35 | 0.73166 | 0.82637 | 0.51106 | 1.0977 | 1.499 | 0.59606 | 0.61731 | 1.5855 | 0.59621
15E017 | 21.3 ± 1.1 | Siemens | 1.5 | Aera | 2012 | VD13A | 70 | 45/200 | 18 ¦ 12 | MOLLI 5b(3s)3b | WIP780B | 1 | 1/52 | 1.12 | 900 (66.7) | 35 | 0.78723 | 0.66272 | 1.0019 | 0.7758 | 1.0887 | 0.77795 | 0.76573 | 1.2315 | 1.0122
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780B | 1 | 1/52 | 1.12 | 900 (66.7) | 35 | 0.64373 | 0.64162 | 0.89348 | 0.70474 | 0.87236 | 0.53316 | 0.67379 | 0.99451 | 0.89677
15E018 | 21.3 ± 0.6 | Siemens | 1.5 | Aera | 2013 | VE11A | 70 | 45/200 | 06 ¦ 12 | MOLLI 5s(3s)3s | WIP1041B | 1 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.26207 | 0.14891 | 0.31955 | 1.267 | 1.7988 | 0.66488 | 0.25423 | 1.4651 | 0.6176
[10 empty] | ShMOLLI | WIP1041B | 1 | 2/12 | 1.07 | 900 (66.7) | 35 | 0.13814 | 0.3432 | 0.20034 | 0.58746 | 0.66662 | 0.4289 | 0.26551 | 0.85912 | 0.12157
[10 empty] | SASHA Pre-contrast | WIP1041B | 1 | 2/12 | 1.12 | 900 (66.7) | 70 | 0.51205 | 0.72166 | 0.099622 | 0.35408 | 0.21178 | 0.55956 | 0.32016 | 0.70313 | 0.71325
15E019 | 22.6 ± 0.5 | Siemens | 1.5 | Aera | 2005 | VE11 | 60 | 45/200 | 18 ¦ 12 | MOLLI Myomaps Pre | Myomaps Product | 1 | 2/12 | 1.13 | 900 (66.7) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
15E020 | 22.0 ± 0.0 | Siemens | 1.5 | Avanto | 2010 | VB17 | 55 | 45/200 | 08 ¦ 12 | ShMOLLI | WIP448C | 1 | 2/12 | 1.05 | 900 (66.7) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
15E021 | 22.3 ± 0.6 | Siemens | 1.5 | Avanto | 2005 | VB17A | 60 | 45/200 | 18 ¦ 08 | MOLLI 5b(3s)3b | WIP448 | 3 | 2/12 | 1.13 | 900 (66.7) | 35 | 0.34477 | 0.57563 | 0.20921 | 0.080052 | 0.19382 | 0.66663 | 0.16919 | 0.632 | 1.5746
[10 empty] | MOLLI 4b(1s)2b(1s)2b | WIP448 | 1 | 2/12 | 1.13 | 900 (66.7) | 35 | 0.15447 | 0.19329 | 0.30528 | 0.16746 | 0.54427 | 0.15831 | 0.11062 | 0.15614 | 0.14344
15E022 | 21.3 ± 0.7 | Siemens | 1.5 | Avanto | 2008 | VB17A | 60 | 45/200 | 06 ¦ 12 | MOLLI 3b(3s)3b(3s)5b | WIP448 | 3 | 2/12 | 1.06 | 900 (66.7) | 35 | 0.4501 | 0.66163 | 0.62236 | 1.0507 | 1.3323 | 0.87314 | 0.61003 | 1.517 | 0.76395
T-TIME#
15E024 | 22.1 ± 0.5 | Siemens | 1.5 | Aera | 2012 | VD13A | 70 | 45/200 | 18 ¦ 09 | MOLLI 5b(3s)3b | WIP780 | 2 | 2/52 | 1.12 | 900 (66.7) | 35 | 0.65432 | 0.66644 | 0.4805 | 1.3274 | 1.6147 | 0.80002 | 0.98547 | 1.65 | 0.61316
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780 | 2 | 2/52 | 1.12 | 900 (66.7) | 35 | 0.36152 | 0.48103 | 0.55878 | 1.0169 | 1.1065 | 0.73598 | 0.49002 | 1.2933 | 0.45199
15E025 | 21.0 ± 0.0 | GE | 1.5 | Discovery MR450 | 2012 | DV24.0 R01 1344.a | 60 | 50/200 | 32 Cardiac | SMART1Map Native | / | 10 | 2/12 | 1.19 | 600 (100) | 65 | 2.9194 | 0.58667 | 1.5687 | 0.92976 | 1.1209 | 0.52619 | 4.21 | 1.9754 | 16.2494
[10 empty] | SMART1Map Post-contrast | / | 5 | 2/12 | 1.19 | 600 (100) | 65 | 5.1901 | 1.437 | 8.4559 | 3.5303 | 5.2619 | 1.961 | 4.2168 | 5.5125 | 18.4248
15E026 | 21.4 ± 0.7 | Siemens | 1.5 | Aera | 2011 | VE11A | 70 | 45/200 | 18 ¦ 12 | MOLLI Myomaps Pre | Myomaps Product | 3 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.6647 | 0.59117 | 0.59651 | 0.82256 | 0.83987 | 0.64448 | 0.80016 | 1.0624 | 0.64872
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 3 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.47424 | 0.56284 | 0.7101 | 0.57151 | 0.64419 | 0.52838 | 0.64218 | 0.82871 | 0.69836
15E027 | 22.0 ± 1.3 | GE | 1.5 | Optima MR450w | 2014 | DV25.0_R01_1451.a | 70 | 44/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | / | 1 | 2/12 | 1.11 | 1000 (60) | 35 | 4.5463 | 4.1363 | 10.5225 | 12.5944 | 7.7483 | 9.617 | 3.7889 | 7.3363 | 6.3535
[10 empty] | Sub-optimal MOLLI 5b(1b)1b(1b)1b in the absence of iterative/data dropping steps in map creation as per ShMOLLI | / | 1 | 2/12 | 1.11 | 1000 (60) | 35 | 1.6929 | 0.85488 | 0.60712 | 1.4406 | 1.8237 | 1.103 | 0.41736 | 1.3249 | 1.6716
15E028 | 23.5 ± 0.6 | Siemens | 1.5 | Avanto | 2006 | VD13B | 60 | 45/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | WIP780 | 3 | 2/12 | 1.40 | 900 (66.7) | 35 | 3.6071 | 4.1877 | 2.5893 | 8.9188 | 9.5367 | 6.3212 | 2.4058 | 2.7512 | 0.698
15E029 | 22.2 ± 1.1 | Philips | 1.5 | Achieva | 2008 | R5.1.8SP2 | 60 | 40/200 | 16 ¦ 16 | MOLLI 5b(3s)3b(3s)2b | / | 1 | 1/12 | 0.91 | 896 (67) | 35 | 0.37967 | 2.2162 | 0.61977 | 8.1108 | 7.4826 | 4.6933 | 0.87065 | 7.6185 | 0.5516
[10 empty] | MOLLI 4b(1s)3b(1s)2b | / | 1 | 1/12 | 0.91 | 896 (67) | 35 | 0.44158 | 0.98763 | 0.47174 | 1.5445 | 2.152 | 1.1331 | 0.44143 | 2.276 | 0.42513
[10 empty] | MOLLI 3b(3s)3b(3s)5b | / | 1 | 1/12 | 0.91 | 896 (67) | 35 | 0.29016 | 0.76726 | 0.38993 | 0.7212 | 1.094 | 0.66316 | 0.33338 | 1.7498 | 0.47234
[10 empty] | SASHA | / | 1 | 1/12 | 0.91 | 896 (67) | 35 | 1.5006 | 1.0192 | 1.5921 | 1.7425 | 0.87271 | 1.4772 | 1.624 | 1.0679 | 1.2497
15E030 | 24.0 ± 0.0 | Siemens | 1.5 | Aera | 2014 | VD13A | 70 | 45/200 | 18 ¦ 09 | Long IR | / | 1 | 4/12 | 21 | / | 180
15E031 | 21.6 ± 1.9 | Siemens | 1.5 | Avanto | 2005 | VB17A [VE11: 20.08.2016] | 60 | 45/200 | 06 ¦ 06 | MOLLI 5b(3s)3b | WIP448B | 2 | 3/52 | 1.03 | 900 (66.7) | 35 | 0.97243 | 0.79277 | 1.1403 | 0.62433 | 1.0012 | 0.53159 | 0.85196 | 1.1192 | 1.4318
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP448B | 2 | 3/52 | 1.12 | 900 (66.7) | 35 | 0.90311 | 0.6675 | 1.1371 | 0.47943 | 0.73804 | 0.42181 | 0.7622 | 1.0586 | 1.4412
15E032 | 21.0 ± 1.2 | Siemens | 1.5 | Avanto | 2006 | VB17A | 60 | 45/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | WIP448B | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.77952 | 0.49117 | 0.52902 | 1.4578 | 2.0996 | 1.2237 | 1.0056 | 2.5132 | 0.62772
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP448B | 1 | 1/12 | 1.00 | 900 (66.7) | 35 | 0.55487 | 0.58834 | 0.63941 | 1.3014 | 1.663 | 1.0745 | 0.62459 | 2.1644 | 0.72377
[10 empty] | ShMOLLI | WIP448C | 1 | 1/12 | 1.05 | 900 (66.7) | 35 | 0.44945 | 0.58813 | 0.7237 | 1.2499 | 1.6169 | 0.97956 | 0.47154 | 2.8653 | 0.71572
15E033ae1 | 21.0 ± 0.0 | Siemens | 1.5 | Aera | 2014 | VD13A | 70 | 45/200 | 18 ¦ 09 | MOLLI 5b(3s)3b | WIP780B | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.87519 | 0.62364 | 0.40142 | 0.89609 | 0.76833 | 0.65741 | 0.91185 | 0.82668 | 0.54943
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780B | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.40801 | 0.54729 | 0.45811 | 0.53752 | 0.64635 | 0.57566 | 0.41727 | 0.67775 | 0.88025
[10 empty] | MOLLI 5s(3s)3s | WIP780B | 1 | 1/12 | 1.01 | 900 (66.7) | 35 | 0.64613 | 0.52079 | 0.43406 | 1.2158 | 0.8249 | 0.65251 | 1.1488 | 1.0087 | 0.41108
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP780B | 1 | 1/12 | 1.01 | 900 (66.7) | 35 | 0.44249 | 0.60023 | 0.39902 | 0.43362 | 0.59368 | 0.86944 | 0.47947 | 0.84986 | 0.39563
[10 empty] | ShMOLLI | WIP780B | 1 | 1/12 | 1.07 | 900 (66.7) | 35 | 0.43346 | 0.57292 | 0.54932 | 0.56268 | 0.75494 | 0.71803 | 0.50794 | 0.86345 | 0.48653
[10 empty] | SASHA HR < 90 | WIP780B | 1 | 1/12 | 1.48 | 900 (66.7) | 70 | 0.3387 | 0.5177 | 0.56294 | 0.41627 | 0.69716 | 0.49967 | 0.43812 | 0.71822 | 0.45444
[10 empty] | SASHA HR > 90 | WIP780B | 1 | 1/12 | 1.32 | 900 (66.7) | 70 | 0.74865 | 0.65371 | 0.81029 | 0.9314 | 0.99385 | 0.88918 | 0.61235 | 0.94383 | 1.007
15E033ae2 | 21.0 ± 0.0 | Siemens | 1.5 | Aera | 2015 | VE11A | 70 | 45/200 | 18 ¦ 09 | MOLLI Myomaps Pre HR < 90 | Myomaps Product | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.68774 | 1.0191 | 0.12745 | 0.77366 | 1.3697 | 1.0357 | 0.26681 | 5.8945 | 5.8945
[10 empty] | MOLLI Myomaps Pre HR > 90 | Myomaps Product | 1 | 1/12 | 1.20 | 900 (66.7) | 35 | 0.65784 | 1.1103 | 0.12731 | 0.67479 | 1.4939 | 0.83049 | 0.28242 | 5.6667 | 5.6667
[10 empty] | MOLLI Myomaps Post HR < 90 | Myomaps Product | 1 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.32442 | 0.71698 | 0.11105 | 0.58367 | 1.0009 | 1.1639 | 0.10528 | 4.4899 | 4.4899
[10 empty] | MOLLI Myomaps Post HR > 90 | Myomaps Product | 1 | 1/12 | 1.01 | 900 (66.7) | 35 | 0.34201 | 0.9403 | 0.083126 | 0.56738 | 0.80172 | 0.89025 | 0.14228 | 4.2463 | 4.2463
[10 empty] | MOLLI 5s(3s)3s | WIP1041 | 1 | 1/12 | 1.13 | 900 (66.7) | 35 | 0.77409 | 0.53111 | 0.14986 | 0.90422 | 1.1419 | 1.3058 | 0.49374 | 5.7297 | 5.7297
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP1041 | 1 | 1/12 | 1.13 | 900 (66.7) | 35 | 0.30106 | 1.2272 | 0.16448 | 0.67866 | 1.3967 | 1.1749 | 0.30362 | 5.3254 | 5.3254
[10 empty] | ShMOLLI | WIP1048 | 1 | 1/12 | 1.07 | 900 (66.7) | 35 | 0.26459 | 0.82739 | 0.091302 | 0.55969 | 1.1943 | 1.0952 | 0.12066 | 6.185 | 6.185
[10 empty] | SASHA Pre-contrast | WIP1041 | 1 | 1/12 | 1.13 | 900 (66.7) | 70 | 0.65057 | 0.96556 | 0.26178 | 0.41991 | 1.1672 | 1.0843 | 0.23252 | 5.1806 | 5.1806
[10 empty] | SASHA Post-contrast | WIP1041 | 1 | 1/12 | 1.13 | 900 (66.7) | 70 | 0.81444 | 1.0489 | 0.2115 | 0.46273 | 1.1871 | 1.298 | 0.28617 | 4.7858 | 4.7858
15E034 | 20.0 ± 0.0 | Siemens | 1.5 | Avanto | 2009 | VB17 | 70 | 45/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | WIP448 | 1 | 4/12 | 1.17 | 900 (66.7) | 35 | 0.99 | 1.51 | 1.07 | 1.12 | 1.59 | 1.49 | 1.12 | 1.35 | 1.43
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP448 | 1 | 4/12 | 1.17 | 900 (66.7) | 35 | 1.07 | 1.03 | 1.47 | 0.78 | 0.99 | 0.99 | 0.92 | 1.51 | 1.11
[10 empty] | ShMOLLI | WIP448 | 1 | 4/12 | 1.05 | 900 (66.7) | 35 | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘ | ⌘
15E035 | 24.0 ± 0.0 | Siemens | 1.5 | Avanto | 2007 | VD13B | 60 | 45/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | WIP780B | 3 | 2/12 | 1.11 | 900 (66.7) | 35 | 0.44319 | 0.1772 | 0.22217 | 0.63119 | 0.9298 | 0.23798 | 0.47673 | 1.3102 | 0.37943
[10 empty] | ShMOLLI | WIP780B | 3 | 2/12 | 1.07 | 900 (66.7) | 35 | 0.24153 | 3.0453 | 0.312 | 2.7569 | 4.1988 | 1.135 | 0.21441 | 5.0842 | 0.21534
15E036 | 18.7 ± 1.2 | Siemens | 1.5 | Avanto | 2005 | VB17A | 60 | 45/200 | 06 ¦ 12 | MOLLI 3b(3s)3b(3s)5b | WIP448 | 1 | 4/12 | 1.04 | 900 (66.7) | 35 | 0.26433 | 0.14216 | 0.75328 | 1.0321 | 1.312 | 0.64662 | 0.15653 | 1.3941 | 0.97367
[10 empty] | MOLLI 5b(3s)3b | WIP448 | 1 | 4/12 | 1.04 | 900 (66.7) | 35 | 0.81776 | 1.1282 | 0.70179 | 0.31797 | 1.2385 | 0.090879 | 0.46582 | 1.215 | 1.208
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP448 | 1 | 4/12 | 1.04 | 900 (66.7) | 35 | 0.43045 | 0.38318 | 0.57822 | 0.53624 | 1.0856 | 0.14776 | 0.4346 | 1.034 | 0.6855
15E036mob | 19.0 ± 0.0 | Siemens | 1.5 | Avanto | 2015 | VD13A | 60 | 45/200 | 12 ¦ 08 | MOLLI 5b(3s)3b | WIP780C | 1 | 4/12 | 1.13 | 900 (66.7) | 35 | 0.33296 | 0.23312 | 0.098662 | 0.13453 | 0.14765 | 0.010598 | 0.022231 | 0.37868 | 0.40124
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780C | 1 | 4/12 | 1.13 | 900 (66.7) | 35 | 0.15048 | 0.14346 | 0.19594 | 0.17239 | 0.39385 | 0.32514 | 0.058277 | 0.19028 | 0.33605
[10 empty] | SASHA | WIP780C | 1 | 4/12 | 1.13 | 900 (66.7) | 70 | 0.14388 | 0.15819 | 0.084891 | 0.42522 | 0.62051 | 0.35106 | 0.23248 | 0.29658 | 0.012459
15E037 | 22.6 ± 0.5 | Siemens | 1.5 | Aera | 2009 | VE11 | 70 | 45/200 | 18 ¦ 12 | MOLLI Myomaps Pre | Myomaps Product | 3 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.38817 | 0.32196 | 0.37118 | 0.70981 | 0.97287 | 0.57739 | 0.54102 | 1.1365 | 0.30957
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 1 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.34196 | 0.33293 | 0.41124 | 0.58311 | 0.86566 | 0.50952 | 0.35084 | 0.94072 | 0.5663
15E038in | 22.6 ± 0.7 | Philips | 1.5 | Achieva XR | 2014 | R5.1.7SP2 | 70 | 40/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 2 | 1/12 | 0.84 | 896 (67) | 35 | 0.31858 | 0.79224 | 0.45533 | 0.43231 | 0.58158 | 0.471 | 0.3415 | 0.48375 | 0.59236
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 2 | 1/12 | 0.84 | 896 (67) | 35 | 0.28145 | 0.49006 | 0.35346 | 0.34666 | 0.5011 | 0.87369 | 0.35216 | 0.38621 | 0.87554
15E039 | 20.1 ± 1.2 | Philips | 1.5 | Achieva | 2005 | R3.2.2 | 60 | 40/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | / | 3 | 2/52 | 1.30 | 896 (67) | 35 | 0.94306 | 2.7983 | 1.3137 | 1.2247 | 1.8549 | 0.62533 | 1.0503 | 2.0809 | 0.89463
[10 empty] | ShMOLLI | / | 3 | 2/52 | 1.36 | 896 (67) | 35 | 0.83043 | 1.129 | 0.87109 | 0.84231 | 1.4894 | 0.52925 | 0.87432 | 1.6793 | 0.92137
15E040 | 21.0 ± 1.4 | Siemens | 1.5 | Aera | 2012 | VD13A | 70 | 45/200 | 18 ¦ 08 | MOLLI 5s(3s)3s | WIP780B | 3 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.34459 | 0.37642 | 0.53163 | 0.71916 | 0.92648 | 0.42273 | 0.43668 | 1.4808 | 0.65389
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP780B | 3 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.58542 | 0.48272 | 0.73659 | 0.71296 | 1.1036 | 0.16991 | 0.52663 | 1.215 | 1.0886
[10 empty] | SASHA Pre-contrast | WIP780B | 3 | 4/12 | 1.12 | 900 (66.7) | 70 | 0.29337 | 0.42405 | 2.035 | 0.92713 | 1.5805 | 0.59776 | 0.70743 | 1.7028 | 2.5581
[10 empty] | SASHA Post-contrast | WIP780B | 3 | 4/12 | 1.12 | 900 (66.7) | 70 | 0.2395 | 0.69257 | 1.0824 | 0.87977 | 1.3068 | 0.76317 | 0.32952 | 1.611 | 0.76691
30E001ac | 22.7 ± 0.6 | Philips | 3 | Achieva TX | 2007 | R3.2.3 | 60 | 40/200 | 16 ¦ 16 | MOLLI 5b(3s)3b | / | 2 | 4/12 | / | 896 (67) | 50 | 3.5157 | 2.9457 | 5.9672 | 7.374 | 8.1339 | 6.0086 | 1.8588 | 6.2414 | 2.27
[10 empty] | MOLLI 3b(3s)3b(3s)5b | / | 2 | 4/12 | / | 896 (67) | 50 | 1.1962 | 0.37407 | 1.8464 | 0.12508 | 0.49898 | 0.10433 | 0.95979 | 1.3395 | 0.52615
[10 empty] | MOLLI 3b(3s)5b | / | 2 | 4/12 | / | 896 (67) | 50 | 0.98486 | 0.47133 | 1.5618 | 0.0075557 | 0.49102 | 0.047036 | 1.2859 | 0.78228 | 0.45803
[10 empty] | Sub-optimal MOLLI 5b(1b)1b(1b)1b in the absence of iterative/data dropping steps in map creation as per ShMOLLI | / | 2 | 4/12 | / | 896 (67) | 50 | 1.0723 | 0.58703 | 0.33985 | 0.28284 | 0.68644 | 0.026137 | 1.5255 | 2.389 | Stubborn Artefact
[10 empty] | SASHA | / | 2 | 4/12 | / | 896 (67) | 50 | 1.0596 | 1.7414 | 1.4939 | 6.0024 | 1.7185 | 0.84549 | 0.072435 | 10.3433 | 0.89863
30E001pt | 21.5 ± 0.7 | Siemens | 3 | Biograph mMR | 2013 | VB20P | 60 | 45/200 | 06 ¦ 24 | MOLLI 5b(3s)3b | WIP448B | 2 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.043186 | 0.025038 | 0.034739 | 0.16704 | 0.36926 | 0.61207 | 0.11641 | 0.48225 | 0.58275
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP448B | 2 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.051951 | 0.021963 | 0.073798 | 0.16164 | 0.30611 | 0.48544 | 0.10597 | 0.72423 | 0.20932
30E002 | 23.0 ± 0.0 | Philips | 3 | Ingenia | 2012 | R5.1.8.2 | 70 | 33/200 | 16 ¦ 16l | MOLLI 3b(3s)3b(3s)5b | / | 1 | 2/52 | 1.35 | 896 (67) | 30 | 1.5242 | 0.48876 | 1.7122 | 1.3311 | 1.5911 | 0.87985 | 0.62876 | 2.4087 | 0.98955
30E003in | 22.1 ± 0.7 | Philips | 3 | Ingenia | 2011 | R5.1.7.2 | 70 | 33/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 2 | 2/52 | 0.98 | 896 (67) | 20 | 0.67958 | 0.61858 | 0.48387 | 1.0318 | 1.6059 | 1.1833 | 0.60273 | 1.674 | 0.58605
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 2 | 2/52 | 0.98 | 896 (67) | 20 | 0.17155 | 0.14325 | 0.45909 | 1.1792 | 1.6221 | 0.28005 | 0.50042 | 1.6549 | Stubborn Artefact
30E003sk | 19.2 ± 1.1 | Siemens | 3 | Skyra | 2014 | VE11 | 70 | 45/200 | 18 ¦ 12 | MOLLI Myomaps Pre | Myomaps Product | 1 | 3/52 | 1.22 | 900 (66.7) | 35 | 0.51025 | 0.49636 | 0.28987 | 1.2883 | 1.4831 | 0.97467 | 0.46997 | 2.1528 | 0.37254
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 1 | 3/52 | 1.22 | 900 (66.7) | 35 | 0.23969 | 0.29911 | 0.35509 | 0.98188 | 1.1139 | 0.66794 | 0.30831 | 1.7358 | 0.35217
30E005 | 20.0 ± 1.3 | Siemens | 3 | Skyra | 2012 | VE11 | 70 | 45/200 | 16 ¦ 16 | MOLLI Myomaps Pre | Myomaps Product | 2 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.19145 | 0.20966 | 0.26453 | 1.402 | 2.157 | 0.71542 | 0.76075 | 3.4556 | 0.4215
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 2 | 2/12 | 1.12 | 900 (66.7) | 35 | 0.091195 | 0.14239 | 0.19796 | 1.8269 | 2.4891 | 0.81453 | 0.056638 | 4.6864 | 0.33522
30E006 | 21.1 ± 0.8 | Philips | 3 | Achieva TX | 2015 | R5.1.7.2 | 60 | 40/200 | 06 ¦ 12 | MOLLI 5s(3s)3s | WIP BTFE 3T MK-Native SENSE | 1 | 3/52 | 0.95 | 896 (67) | 20 | 0.60529 | 1.0218 | 0.28096 | 4.1952 | 4.8734 | 2.0813 | 0.5745 | 5.9715 | 0.73862
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP BTFE 3T MK-Enhanced SENSE | 1 | 3/52 | 0.95 | 896 (67) | 20 | 0.52124 | 0.53799 | 0.42616 | 1.3204 | 1.3824 | 1.1068 | 0.54994 | 1.8391 | 0.7325
30E007 | 21.0 ± 2.4 | Siemens | 3 | Verio | 2009 | VB17 | 70 | 45/200 | 06 ¦ 12 | MOLLI 3b(3s)3b(3s)5b | WIP448 | 3 | 1/12 | 1.35 | 900 (66.7) | 35 | 1.2294 | 1.156 | 3.2839 | 2.5815 | 3.0354 | 1.7647 | 0.81494 | 4.0395 | 3.9106
30E008 | 21.6 ± 0.9 | Philips | 3 | Achieva TX | 2012 | R5.1.8.2 | 70 | 40/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.30287 | 0.66323 | 0.49941 | 6.0357 | 4.7932 | 4.0826 | 0.3937 | 5.171 | 0.79091
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.17516 | 0.43889 | 0.8123 | 2.9124 | 3.7771 | 4.3273 | 0.36815 | Stubborn Artefact | 1.3593
[10 empty] | MOLLI 3b(3s)3b(3s)5b | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.27346 | 0.46863 | 0.78602 | 1.2949 | 2.0064 | 0.94141 | 0.28213 | 4.0147 | 1.3829
30E009 | 23.0 ± 0.0 | Philips | 3 | Achieva TX | 2011 | R3.2.3.2 | 60 | 40/200 | 06 ¦ 12 | MOLLI 5b(3s)3b | / | 4 | 4/12 | 0.89 | 896 (67) | 35 | 0.2756 | 0.56364 | 2.3299 | 0.71474 | 1.1857 | 0.41871 | 0.49279 | 1.5068 | 2.3488
[10 empty] | MOLLI 4b(1s)3b(1s)2b | / | 3 | 4/12 | 0.89 | 896 (67) | 35 | 0.055401 | 0.59241 | 2.4528 | 0.47866 | 0.58966 | 0.78547 | 0.20324 | 1.2246 | 2.3241
30E010 | 23.0 ± 0.0 | Siemens | 3 | Skyra | 2014 | VE11 | 70 | 45/200 | 16 ¦ 16 | MOLLI Myomaps Pre | Myomaps Product | 1 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.59639 | 0.5989 | 0.36445 | 0.83019 | 0.93295 | 0.94331 | 0.82889 | 1.3638 | 0.38635
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 1 | 4/12 | 1.12 | 900 (66.7) | 35 | 0.28554 | 0.21806 | 0.13302 | 0.87625 | 1.3042 | 0.60974 | 0.68805 | 2.9771 | 0.21285
30E012 | 23.0 ± 0.9 | GE | 3 | Discovery MR750w | 2014 | DV25.0_R01_1451.a [DV25.0_R02_1549.b: 24.02.2016] | 70 | 44/200 | Body 30 Small | SMART1Map Native | / | 1 | 1/52 | 2.02 | 600(100) | 45 | 4.9334 | 1.182 | 6.5294 | 2.9023 | 4.2327 | 2.0763 | 5.1057 | 6.2922 | 9.7829
30E014 | 21.5 ± 0.8 | Siemens | 3 | Prisma | 2015 | VE11 | 60 | 80/200 | 18 ¦ 09 | MOLLI Myomaps Pre | Myomaps Product | 3 | 1/12 | 1.12 | 900 (66.7) | 35 | 1.0006 | 1.051 | 1.6356 | 1.792 | 2.1473 | 1.5439 | 1.5728 | 3.1143 | 2.0428
[10 empty] | MOLLI Myomaps Post | Myomaps Product | 2 | 1/12 | 1.12 | 900 (66.7) | 35 | 0.27259 | 0.34125 | 0.33458 | 1.1519 | 1.6645 | 0.79944 | 0.45764 | 3.3026 | 0.24735
30E015 | 21.9 ± 0.3 | Philips | 3 | Ingenia | 2011 | R5.1.8.2 | 70 | 33/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.3064 | 0.56865 | 0.7191 | 4.6856 | 6.8557 | 4.5493 | 0.34724 | 4.0018 | 0.99137
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.23774 | 0.35535 | 0.69716 | 1.7886 | 2.6844 | 1.2184 | 0.28429 | 3.9024 | 1.2409
[10 empty] | MOLLI 3b(3s)3b(3s)5b | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 0.24159 | 0.30221 | 0.50914 | 4.0082 | 5.9736 | 2.307 | 0.26491 | 8.4393 | 0.96988
[10 empty] | SASHA | / | 1 | 1/12 | 0.91 | 896 (67) | 20 | 1.7733 | 1.2496 | 1.8397 | 1.5213 | 5.2397 | 3.7011 | 1.4752 | 4.1017 | 3.0223
30E016 | 23.3 ± 0.6 | Philips | 3 | Achieva TX | 2012 | R5.1.7.2 | 60 | 40/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 2 | 4/12 | 0.96 | 896 (67) | 35 | 8.5772 | 4.7957 | 1.8956 | 0.7589 | 3.6904 | 4.8066 | 0.61797 | 5.8167 | 5.5223
[10 empty] | MOLLI 5b(3s)3b | / | 3 | 4/12 | 0.96 | 896 (67) | 35 | 10.0731 | 5.5417 | 3.6313 | 1.8204 | 0.80529 | 4.7973 | 0.66267 | 1.8361 | 6.4412
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 1 | 4/12 | 1.02 | 896 (67) | 35 | 8.5065 | 5.4222 | 4.5862 | 2.0179 | 5.2733 | 6.7256 | 1.0677 | 9.3135 | 4.5643
[10 empty] | Sub-optimal MOLLI 5b(1b)1b(1b)1b in the absence of iterative/data dropping steps in map creation as per ShMOLLI | / | 3 | 4/12 | 0.96 | 896 (67) | 35 | 8.4356 | Stubborn Artefact | 2.7767 | 5.3748 | 8.8177 | 1.3972 | 5.7318 | 11.5068 | 2.3457
30E017 | 20.4 ± 0.7 | Siemens | 3 | Skyra PTX | 2011 | VD13C [VE11: 03.10.2016] | 70 | 45/200 | 12 ¦ 08 | MOLLI 5b(3s)3b | WIP780B [Myomaps Product: 03.10.2016] | 2 | 3/52 | 1.12 | 900 (66.7) | 35 | 0.27835 | 0.40865 | 0.27649 | 1.2194 | 1.4254 | 0.8421 | 0.35625 | 1.9667 | 0.27101
[10 empty] | MOLLI 4b(1s)3b(1s)2b | WIP780B [Myomaps Product: 03.10.2016] | 2 | 3/52 | 1.12 | 900 (66.7) | 35 | 0.27852 | 0.42935 | 0.31976 | 1.0214 | 1.1806 | 0.74691 | 0.3008 | 1.8443 | 0.26021
30E019 | 19.3 ± 2.0 | Philips | 3 | Achieva TX | 2010 | R5.1.7.2 | 60 | 40/200 | 16 ¦ 16 | MOLLI 3s(3s)5s | / | 2 | 2/52 | 0.89 | 896 (67) | 20 | 0.57458 | 0.85637 | 0.57675 | 2.696 | 3.2519 | 2.0694 | 0.88906 | 4.276 | 0.80241
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 2 | 2/52 | 0.89 | 896 (67) | 20 | 0.54075 | 0.82093 | 0.54383 | 2.4907 | 2.8013 | 1.9492 | 0.76367 | 4.5833 | 0.57497
30E020 | 20.3 ± 0.8 | Siemens | 3 | Skyra | 2012 | VD13A | 70 | 45/200 | 18 ¦ 09 | MOLLI 5b(3s)3b | WIP780B | 1 | 2/52 | 1.07 | 900 (66.7) | 35 | 0.32526 | 0.47915 | 0.47844 | 1.2877 | 1.5461 | 1.07 | 0.67534 | 2.0537 | 0.34934
[10 empty] | ShMOLLI | WIP780B | 1 | 2/52 | 1.07 | 900 (66.7) | 35 | 0.42511 | 0.46247 | 0.33111 | 1.4123 | 1.6222 | 1.1053 | 0.46921 | 2.2302 | 0.15404
[10 empty] | SASHA | WIP780B | 1 | 2/52 | 1.07 | 900 (66.7) | 70 | 3.6057 | 0.69929 | 0.58508 | 3.7217 | 1.8564 | 1.1135 | 10.9564 | 1.8321 | 0.92687
PTB 30E021 | Range: 19 – 24 | Siemens | 3 | Verio | 2010 | VB17A | 70 | 45/200
30E023 | 23.0 ± 0.0 | Philips | 3 | Ingenia | 2010 | R5.1.7.2 | 70 | 33/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | / | 2 | 1/12 | 1.14 | 896 (67) | 35 | 0.23044 | 0.11197 | 0.75734 | 0.44096 | 0.30012 | 0.80913 | 0.20371 | 0.45113 | 0.59062
[10 empty] | MOLLI 4s(1s)3s(1s)2s | / | 2 | 1/12 | 1.14 | 896 (67) | 35 | 0.29884 | 0.21386 | 0.85631 | 0.27375 | 0.26368 | 1.4645 | 0.16715 | 0.40695 | 0.39761
30E024 | 21.3 ± 1.4 | Siemens | 3 | Skyra | 2012 | VD13A | 70 | 45/200 | 16 ¦ 16 | MOLLI 5s(3s)3s | WIP780B | 3 | 2/12 | 1.12 | 900 (66.7) | 20 | 0.13035 | 0.28667 | 0.33092 | 1.0221 | 1.283 | 0.90865 | 0.1961 | 1.5574 | 0.49595
[10 empty] | MOLLI 4s(1s)3s(1s)2s | WIP780B | 3 | 2/12 | 1.12 | 900 (66.7) | 20 | 0.17032 | 0.19566 | 0.29165 | 0.85309 | 0.91244 | 0.63851 | 0.1728 | 1.4622 | 0.77209
[10 empty] | SASHA Pre-contrast | WIP780B | 3 | 2/12 | 1.26 | 900 (66.7) | 50 | 0.45348 | 0.45033 | 0.85926 | 1.2345 | 1.5143 | 0.92704 | 0.42096 | 1.738 | 0.73597
[10 empty] | SASHA Post-contrast | WIP780B | 3 | 2/12 | 1.26 | 900 (66.7) | 50 | 0.38974 | 0.50992 | 0.24833 | 1.2068 | 1.4559 | 1.0624 | 0.23461 | 1.9573 | 0.41874
30E025t7≈ |  | Siemens | 7 | Magnetom 7 | 2013 | VB17B | 60 | 72/200
30E026 | 22.0 ± 1.4 | Siemens | 3 | Skyra | 2010 | VD13A | 70 | 45/200 | 18 ¦ 12 | MOLLI 5s(3s)3s | WIP780B | 1 | 4/12 | 1.08 | 900 (66.7) | 35 | 0.29391 | 0.22133 | 0.27188 | 1.4553 | 2.3612 | 1.7604 | 0.40781 | 3.0357 | 0.4079
[10 empty] | ShMOLLI | WIP780B | 1 | 4/12 | 1.07 | 900 (66.7) | 35 | 0.63285 | 0.20514 | 0.42149 | 1.1703 | 1.975 | 0.71326 | 0.20052 | 3.2162 | 0.31292
30E027 | 23.0 ± 0.0 | Siemens | 3 | Prisma | 2014 | VD13D | 60 | 80/200 | 16 ¦ 16 | ShMOLLI | WIP780C | 1 | 2/52 | 1.07 | 900 (66.7) | 35 | 0.26904 | 0.33472 | 0.32608 | 0.77993 | 0.88516 | 0.64225 | 0.25785 | 1.5147 | 0.22204
[10 empty] | SASHA HR < 90 | WIP780C | 1 | 2/52 | 1.26 | 900 (66.7) | 50 | 0.3192 | 0.55878 | 0.41152 | 0.85624 | 1.1191 | 0.85384 | 0.35047 | 1.3137 | 0.28975
NIST US~ |  | Agilent | 1.5 | Varian | 2013 | VnmrJ 4 | 14 | 300/475 | 140 mm quadrature birdcage coil
NIST US~ |  | Agilent | 1.5 | Varian | 2013 | VnmrJ 4 | 14 | 300/475 | 140 mm quadrature birdcage coil
Of 73 magnets at 69 centers that originally signed up to the T1MES program, serial data from 8 magnets were not returned on account of local staffing issues, 3 phantoms were donated to national metrology institutes, 2 were donated to support ongoing clinical trials (#, ≈), and submitted data from 4 magnets (all 1.5T, ⌘) did not meet the minimum 1 year study requirement.
# This phantom is a gift to support the ongoing 'T-TIME' study. It is being scanned across multiple UK centers.
≈ This phantom is a gift to support ongoing T1 mapping research at 7 Tesla.
* Inner diameter i.e. around patient.
§ Maximum gradient performances as returned on the T1MES registration forms by each site. These values are subject to many modifying conditions.
^ Refers to the anterior and posterior coil hardware used for phantom scans (that we specified, should be the same as that used by centers in clinical practice). Coil data does not refer to the number of coil groups enabled by centers for the scan.
$ Summarises the number of times a sequence was repeated within the same scan session. Where count = 1, the sequence was only run once (no intra-scacn repeats); where count > 1, the sequence was run multiple times in the same scan session.
≅ Scan frequency denoted as X/52 or X/12; e.g. 2/52 = one scan every 2 weeks; 4/12 = one scan every quarter year.
¶ Centers were instructed to set RR interval to 900ms where possible.
¦¦ TR and TE were extracted from uploaded Digital Imaging and Communications in Medicine (DICOM) images.
~ Loose tubes only for 1.5T and 3T.
∞ Not counting intra-study repeats of a given T1 mapping sequence. Tubes in 1.5 T phantom are: I, H, G, F, E, D, C, B, A. Tubes in 3 T phantom are: I, H, G, M, L, K, C, J, A.
⌘ Insufficient serial data for computing reliable CoV from a site or for a given sequence (<1 year).
Stubborn artefact that recurred in >25% of the serial submissions in spite of corrective actions.
A¦P = anterior ¦ posterior number of coil channels; CoV = coefficient of variation; FA = flip angle; IR = inversion recovery; HR = heart rate; NI = nominal interval; SE = spin echo; T = Tesla; TE = echo train length; Temp = temperature; TR = repetition time; YOM = year of manufacture.