Appendix 1- Tables for parameters of MyoMapNet and imaging sequences, phantom results, and in-vivo T1 precision

Table S1. Notation, input, training, prediction, and application of each model

| MyoMapNet | Input T1-weighted Signals | Training Dataset | Predication and Application | Evaluation |
| --- | --- | --- | --- | --- |
| 4, PreGd | 4 | MOLLI5(3)3 | Native T1, ECV with MyoMapNet 4, PostGd | Existing MOLLI5(3)3 data, prospectively phantom and native in-vivo data by LL4 |
| 4, PostGd | 4 | MOLLI4(1)3(1)2 | Post-contrast T1 and ECV with MyoMapNet 5, PreGd or MyoMapNet 4, PreGd | Existing MOLLI5(3)3 data, prospectively phantom and native in-vivo data by LL4 |
| 4, Pre+PostGd | 4 | MOLLI5(3)3 and MOLLI4(1)3(1)2 | Native and post-contrast T1, ECV | Existing MOLLI5(3)3 data, prospectively phantom and native in-vivo data by LL4 |
| 5, PreGd | 5 | MOLLI5(3)3 | Native T1 , ECV with MyoMapNet 4, PostGd | Only existing MOLLI5(3)3 data, |

Table S2. Imaging parameters for all sequences used in this study

| Parameter | IR-SE | CPMG-SE | MOLLI5(3)3 | MOLLI4(1)3(1)2 | LL4 |
| --- | --- | --- | --- | --- | --- |
| Readout Sequence | spin echo | spin echo | bSSFP | bSSFP | bSSFP |
| Field of View (R×P mm3) | 200×200 | 200×200 | 360×325 | 360×325 | 360×325 |
| Voxel Size (mm3) | 0.8×0.8×8 | 0.8×0.8×8 | 1.7×1.7×8 | 1.7×1.7×8 | 1.7×1.7×8 |
| Slice | 1 | 1 | 3 | 3 | 1 |
| Flip Angle (º) | 90 | 90 | 35 | 35 | 35 |
| TR | 10 s | 10 s | 2.5 ms | 2.5 ms | 2.5 ms |
| TE (ms) | 8 | 8.5 | 1.02 | 1.02 | 1.02 |
| Bandwidth (Hz/Pixel) | 250 | 250 | 1093 | 1093 | 1093 |
| GRAPPA Acceleration | - | - | R=2 | R=2 | R=2 |
| Partial Fourier | - | - | 7/8 | 7/8 | 7/8 |
| Lines of k-space per Shot | 1 | 1 | 75 | 75 | 75 |
| Acquisition Window (ms) | - | - | 198 | 198 | 198 |
| T1- or T2- weighted Images | 14 | 8 | 8 | 9 | 4 |
| Inversion-Recovery/Echo Time (ms) | 100, 200, 300, 400, 500, 600, 700, 800, 900, 1000, 1500, 2000, 2500, 3000 | 8.5, 17.0, 25.5, 34.0, 42.5, 51.0, 59.5, 68.0 | 108,188 | 108,188,268 | 108 |

* IR-SE: Inversion-Recovery Spin-Echo; CPMG-SE: Carr-Purcell-Meiboom-Gill Spin-Echo; MOLLI: Modified Look-Locker inversion recovery. bSSFP: balanced Steady-State Free-Precision; R×P represents readout, phase; GRAPPA: GeneRalized Autocalibrating Partial Parallel Acquisition.

Table S3. Mean, standard deviation (SD) and coefficient of variation (CV) of T1 of each phantom estimated by MOLLI and three MyoMapNet models

|  | IR-SE | CPMG-SE | MOLLI |  |  |  | MyoMapNet |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | 5(3)3 |  | 4(1)3(1)2 |  | 4, PreGd |  | 4, Pre+PostGd |  | 4, PostGd |  |
| # | T1 (ms) | T2 (ms) | T1 (ms) | CV(%) | T1 (ms) | CV(%) | T1 (ms) | CV% | T1 (ms) | CV% | T1 (ms) | CV% |
| 1 | 310 | 146 | 304±4 | 1.3±0.17 | 303±3.4 | 1.1±0.02 | 294±5.1 | 1.7±0.03 | 303±4.8 | 1.6±0.04 | 311±3.4 | 1.1±0.02 |
| 2 | 353 | 45 | 313±15 | 4.8±0.42 | 319±4.1 | 1.3±0.02 | 302±5.7 | 1.9±0.07 | 311±5.2 | 1.7±0.05 | 317±3.8 | 1.2±0.04 |
| 3 | 415 | 166 | 403±7 | 1.7±0.04 | 404±6.3 | 1.5±0.03 | 394±8.2 | 2.1±0.03 | 410±8.6 | 2.1±0.04 | 408±9.1 | 2.2±0.03 |
| 4 | 497 | 45 | 453±4 | 0.9±0.02 | 455±3.7 | 0.8±0.02 | 445±5.7 | 1.3±0.03 | 454±4.3 | 1.0±0.03 | 453±4.4 | 1.0±0.04 |
| 5 | 515 | 170 | 502±4.7 | 0.9±0.03 | 502±4.3 | 0.9±0.02 | 514±5.9 | 1.1±0.07 | 505±5.2 | 1.0±0.06 | 506±5 | 1.0±0.04 |
| 6 | 620 | 45 | 552±4.9 | 0.9±0.03 | 554±4.4 | 0.8±0.02 | 559±5.1 | 0.9±0.04 | 548±5.1 | 0.9±0.04 | 548±4.9 | 0.9±0.05 |
| 7 | 724 | 42 | 620±6.3 | 1.0±0.01 | 621±5.7 | 0.9±0.02 | 624±6.1 | 1.0±0.02 | 612±6.1 | 1.0±0.02 | 610±5.7 | 1.0±0.02 |
| 8 | 922 | 41 | 714±5.5 | 0.8±0.03 | - | - | 715±5.5 | 0.8±0.04 | 706±6.4 | 0.9±0.05 | - | - |
| 9 | 942 | 49 | 814±4.4 | 0.5±0.02 | - | - | 812±4.8 | 0.6±0.02 | 811±5.5 | 0.7±0.03 | - | - |
| 10 | 1177 | 43 | 986±6 | 0.6±0.02 | - | - | 982±6.6 | 0.7±0.05 | 981±7 | 0.7±0.06 | - | - |
| 11 | 1419 | 247 | 1367±10.3 | 0.8±0.03 | - | - | 1367±11.1 | 0.8±0.06 | 1365±12.4 | 0.9±0.06 | - | - |
| 12 | 1688 | 227 | 1563±12.9 | 0.8±0.02 | - | - | 1573±13.5 | 0.9±0.05 | 1573±13 | 0.9±0.05 | - | - |
| µ±σ | - | - | - | 1.3±1.2 | - | 1.0±0.3 | - | 1.2±0.5 | - | 1.1±0.4 | - | 1.2±0.5 |

*IR-SE: inversion-recovery spin-echo sequence; CPMG-SE: Carr-Purcell-Meiboom-Gill spin-echo; MOLLI: Modified Look-Locker inversion recovery;

** Vial #1-7 for post-contrast myocardium and blood; Vial #8-12 for pre-contrast myocardium and blood.

***T1, SD and CV are calculated by averaging the corresponding results of each measurement across all repetitions.

****All P>0.05 when compared to MOLLI5(3)3 or MOLLI4(1)3(1)2

Table S4. Standard deviation (SD) and coefficient of variation (CV) of T1 of existing data estimated by MOLLI and MyoMapNet

|  |  | Myocardium |  | Blood |  |
| --- | --- | --- | --- | --- | --- |
|  |  | SD (ms) | CV (%) | SD (ms) | CV (%) |
| Native T1 | MyoMapNet4, Pre | 61.6±12.3 | 5.2±1.1 | 37.7±10.4 | 2.0±0.6 |
|  | MyoMapNet4, Pre+PostGd | 62.0±12.5 | 5.2±1.1 | 37.8±10.5 | 2.0±0.6 |
|  | MyoMapNet5, Pre | 61.4±12.2 | 5.2±1.1 | 36.8±10.2 | 2.0±0.5 |
|  | MOLLI5(3)3 | 59.7±12.6 | 5.0±1.1 | 33.9±9.8 | 1.8±0.5 |
| Post-Contrast T1 | MyoMapNet4, PostGd | 33.6±12.3 | 5.9±2.4 | 15.6±5.2 | 3.7±1.2 |
|  | MyoMapNet4, Pre+PostGd | 34.3±11.9 | 6.1±2.3 | 15.5±5.3 | 3.7±1.2 |
|  | MOLLI4(1)3(1)2 | 29.4±9.5 | 5.2±1.8 | 9.2±2.5 | 2.2±0.6 |

*All P<0.05 when compared to MOLLI5(3)3 or MOLLI4(1)3(1)2

Table S5. Standard deviation (SD) and coefficient of variation (CV) of T1 for prospectively collected data by MOLLI and LL4

|  |  | Myocardium |  | Blood |  |
| --- | --- | --- | --- | --- | --- |
|  |  | SD (ms) | CV (%) | SD (ms) | CV (%) |
| Native T1 | MyoMapNet4, Pre | 55.6±10.9 | 4.7±0.9 | 34.1±11.6§ | 1.9±0.7§ |
|  | MyoMapNet4, Pre+PostGd | 57.0±13.4 | 4.8±0.9 | 34.3±14.8§ | 1.9±0.8 |
|  | MOLLI5(3)3 | 55.5±12.1 | 4.7±1.0 | 28.6±12.8 | 1.6±0.7 |
| Post-Contrast T1 | MyoMapNet4, PostGd | 28.8±11.3 | 4.2±2.1 | 14.7±6.4§ | 2.7±1.4§ |
|  | MyoMapNet4, Pre+PostGd | 29.6±9.1§ | 4.3±1.8§ | 15.0±6.7§ | 2.7±1.4§ |
|  | MOLLI4(1)3(1)2 | 26.1±9.1 | 3.8±1.4 | 9.4±7.9 | 1.7±1.4 |

§p-value < 0.05 when compared to MOLLI5(3)3 or MOLLI4(1)3(1)2