Fig S1. Rigid registration of CT-CBCT and sCT-CBCT images for a representative H&N case. Panels (a1–a3) show the CT-CBCT registration in the axial, coronal, and sagittal planes, respectively. Panels (b1–b3) display the corresponding sCT-CBCT registration. The delineated structures displays PTV, brainstem, and spinal cord.

Fig S2. Rigid registration of CT-CBCT and sCT-CBCT images for a representative H&N case. Panels (a1–a3) show the CT-CBCT registration in the axial, coronal, and sagittal planes, respectively. Panels (b1–b3) display the corresponding sCT-CBCT registration. The delineated structures are displayed: PTV, femoral heads, and rectum.

Table S1 Overview of clinical evaluation criteria with results obtained in this study.

| sCT Evaluation | Parameters (Guideline or Reference) | Criteria* | Results |  |
| --- | --- | --- | --- | --- |
|  |  |  | H&N | Pelvis |
| Geometric Accuracy | DSC (TG-132 [25]) | ≥ 0.8-0.9 | Boby 0.99 ± 0.00 Soft Tissue 0.96 ± 0.01 Bone 0.87 ± 0.03 Air 0.82 ± 0.04 | Body 1.00 ± 0.00 Soft Tissue 0.98 ± 0.01 Bone 0.88 ± 0.02 |
| Image Intensity | ME | - | Body: -8.94 ± 9.50 Soft tissue: 7.75 ± 7.34 Bone: -121.08 ± 34.45 Air: 118.65 ± 41.65 | Body: -8.74 ± 6.26 Soft tissue: -1.25 ± 4.57 Bone: -71.31 ± 26.46 |
|  | MAE [32] | H&N: Body 65-131 HU Pelvis: Body 27-65 HU | Body 67.30 ± 6.94 Soft Tissue 29.91 ± 3.30 Bone 203.57 ± 26.79 Air 211.20 ± 34.51 | Body 33.64 ± 3.91 Soft Tissue 24.17 ± 3.51 Bone 112.92 ± 15.24 |
| Dose Difference | DVH Difference [1, 35] | ≤ 2% | ≤ 0.76 ± 0.55% | ≤ 0.35 ± 0.27% |
|  | Gamma Pass Rate (TG-218 [28]) | 3%/2 mm ≥ 95% 2%/2 mm ≥ 90% | 3%/2 mm: 96.34 ± 1.31% 2%/2 mm: 95.06 ± 1.40% | 3%/2 mm: 99.11 ± 0.76% 2%/2 mm: 98.72 ± 1.03% |
| Treatment Plan QA | Global Gamma Pass Rate 3%/2 mm (TG-218 [28]) 2%/2 mm [42] | 3%/2 mm ≥ 95% 2%/2 mm ≥ 90% | 3%/2 mm: 98.83 ± 3.13% 2%/2 mm: 97.36 ± 5.11% | 3%/2 mm: 99.31 ± 0.77% 2%/2 mm: 94.71 ± 7.63% |
|  | sCT-CT Gamma Pass Rate Difference | - | 3%/ 2 mm: 0.01 ± 0.04% 2%/2 mm: 0.05 ± 0.06% | 3%/2 mm: 0.05 ± 0.08% 2%/2 mm: 0.12 ± 0.16% |
| Patient Set-up Verification | CT-CBCT and sCT-CBCT differences | - | ≤ 0.13 mm in translations ≤ 0.06° in rotations | ≤ 0.17 mm in translations ≤ 0.13° in rotations |

*As there are no standardized guidelines for sCT evaluation, criteria were adopted based on existing clinical recommendations and relevant publications.

Table S2 Overview of key geometric and dosimetric metrics from prior studies.

| First author, year [Ref] | Region | No. of patients | sCT Generated Software | Geometric Accuracy | Image Intensity Accuracy |  | Dosimetric Accuracy |  | Treatment Plan QA Evaluation | Patient Verification |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | DSC | ME (HU) | MAE (HU) | %DVH Difference | Gamma Pass Rate (%) | sCT Plan- CT Plan difference | CBCT-sCT vs CBCT-CT |
| This study | H&N Pelvis | 10 H&N 9 Prostate 9 Rectum 5 Cervix | MRI Planner v.2.3 | H&N Body: 0.99 ± 0.00 Soft tissue: 00.96 ± 0.01 Bone: 0.87 ± 0.03 Air: 0.82 ± 0.04 Pelvis Body: 1.00 ± 0.00 Soft tissue: 0.98 ±0.01 Bone: 0.88 ± 0.02 | H&N: Body: -8.94 ± 9.50 Soft tissue: 7.75 ± 7.34 Bone: -121.08 ± 34.45 Air: 118.65 ± 41.65 Pelvis Body: -8.74 ± 6.26 Soft tissue: -1.25 ± 4.57 Bone: -71.31 ± 26.46 | H&N Body: 67.30 ± 6.94 Soft tissue: 29.91 ± 3.30 Bone: 203.57 ± 26.79 Air: 211.20 ± 34.51 Pelvis Body: 33.64 ± 3.91 Soft tissue: 24.17 ± 3.51 Bone: 112.92 ± 15.24 | ≤0.76% for H&N ≤0.35% for pelvis | H&N 3%/2 mm: 96.34 ± 1.31 2%/2 mm: 95.06 ± 1.40 1%/1 mm: 92.09 ± 2.02 Pelvis 3%/2 mm: 99.11 ± 0.76 2%/2 mm: 98.72 ± 1.02 1%/1 mm: 97.58 ± 1.73 | H&N 3%/2 mm: 0.01 ± 0.04% 2%/2 mm: 0.05 ± 0.06% Pelvis 3%/2 mm: 0.05 ± 0.08% 2%/2 mm: 0.12 ± 0.16% | H&N Translation: ≤ 0.13 mm Rotation: ≤ 0.06° Pelvis Translation: ≤ 0.17 mm Rotation: ≤ 0.13° |
| Palmér et al., 2017 [9] | H&N | 44 | MRI Planner v.2.3 | Body: 0.98 ± 0.05 Soft tissue: 0.95 ± 0.05 Bone: 0.80 ± 0.07 Air: 0.81 ± 0.10 | Body: -5 ± 10 Soft tissue: -1 ± 7 Bone: -62 ± 28 Air: 107 ± 75 | Body: 67 ± 14 Soft tissue: 38 ± 6 Bone: 195 ± 27 Air: 198 ± 68 | < 0.3% | 2%/ 1 mm: 99.4% | - | - |
| Lerner et al., 2021 [20] | Brain | 20 | MRI Planner v.2.3 | - | Body: -5.6 ± 4.6 Soft tissue: -1.3 ± 2.0 Bone: -46.3 ± 19.4 | Body: 62.2 ± 4.1 Soft tissue: 9.5 ± 0.7 Bone: 176.5 ± 18.8 | < 0.2% | 3%/ 3 mm: 100 ± 0.00 2%/ 2 mm: 99.8 ± 0.2 1%/ 1 mm: 99.1 ± 0.6 | - | Translation: -0.3 to 0.5 mm Rotation: -0.2 to 0.2° |
| Cronholm et al., 2020 [19] | Prostate | 34 | MRI Planner v.2.3 | - | - | Body: 39.6 ± 6.7 Soft tissue: 24.2 ± 3.0 Bone: 176.5 ± 18.8 | Mean in Target: -0.02% | 2%/2 mm: 99.94% | - | Translation: 0.02 ± 0.07 mm Rotation: 0.05 ± 0.29° |
| Singhrao et al., 2023 [3] | H&N | 14 | MRI Planner v.2.2 | - | - | - | Target: ≤ 2.3% OARs: ≤ 0.7% | 3%/ 3 mm: 95.5 ± 2% 2%/2 mm: 92.7 ± 2.1% | - | - |
| Han et al., 2017 [10] | Prostate | 18 | In House DCNN: U-Net | - | Body: -3.1 ± 17.3 | Body: 84.8 ± 17.3 | - | - | - | - |
| Wang et al., 2019 [13] | H&N | 33 | In House DCNN: U-Net | - | Body: -6 ± 13 Soft tissue: -48 ± 10 Bone: 247 ± 44 | Body: 131 ± 24 Soft tissue: 97 ± 13 Bone: 357 ± 44 | - | - | - | - |
| Maspero et al., 2018 [17] | Pelvis | 59 Prostate 18 Rectum 14 cervix | In House cGAN | - | Body Prostate: 1 ± 6 Rectum: 2 ± 9 Cervix: 4 ± 10 | Body Prostate: 65 ± 10 Rectum: 56 ± 5 Cervix: 59 ± 6 | Target: ≤ 0.3% | 3%/ 3 mm: 97% 2%/ 2 mm: 91% | - | - |
| Wang et al., 2022 [30] | Brain | 24 | In House CUT vs cycleCUT | - | - | CUT Body: 78.05 ± 8.26 cycleCUT Body: 69 ± 5.68 | - | 3%/ 3 mm & 2%/ 2 mm: ≥ 97% | - | - |
| Fridström et al., 2024 [31] | Prostate | 10 | MRI Planner v.2.3 | - | - | Body: 71 Femoral Heads: 126 | VMAT CTV: -0.8 ± 0.8 OARs: -0.4 ± 0.2 IMPT CTV: 0.1 ± 0.1 OARs: -0.3 ± 0.4 | Global 3%/ 3mm VMAT: 99.8 IMPT: 97.8% Global 2%/2 mm VMAT: 99.4% IMPT: 95.8% Local 3%/3mm: VMAT:98.9% IMPT: 97.2% Local 2%/2 mm: VMAT: 97.7% IMPT: 94.9% | - | - |
| Dinkla et al., 2019 [14] | H&N | 34 | In House DCNN: U-Net | Body: 0.98 ± 0.01 Bone: 0.70 ± 0.07 Air: 0.79 ± 0.08 | Body: 9 ± 11 Soft tissue: -2 ± 5 Bone: 106 ± 51 Air: -57 ± 41 | Body: 75 ± 9 Soft tissue: 35 ± 3 Bone: 214 ± 26 Air: 103 ± 24 | < 1% | - | - | - |
| Siversson et al., 2015 [21] | Prostate | 10 | MRI Planner v.1.1 | - | - | Body: 36.5 ± 4.1 Soft tissue: 34.0 ± 4.1 Bone: 98.1 ± 10.3 | 0.0 ± 0.2 | 2%/1 mm: 100% for SPACE and 99.7% for FSE | - | - |
| Persson et al., 2017 [34] | Prostate | 170 | MRI Planner | - | - | - | < 0.3% | 2%/2 mm: 99.7% | - | - |
| Ma et al., 2021 [35] | H&N | 20 | In House U-net vs. GAN | - | U-net Body: -9.3 ± 16.9 GAN Body: -8.7 ± 17.3 | U-net Body: 102.6 ± 11.4 GAN Body: 104.2 ± 12.5 | Target: ≤ 0.4% OARs: ≤ 0.1% | 3%/ 2mm: 99.1% | - | - |
| O’Corner et al., 2022 [36] | Pelvis | 40 (20 Male & 20 Female) | In House Bulk Density vs. Tissue Class Density vs. Hybrid Atlas vs. cGAN | - | Bulk Density Body: 8 ± 13.7 Soft tissue: 7.8 ± 14.7 Bone: 5.7 ± 39.3 Tissue Class Density Body: -9.8 ± 7.3 Soft tissue: -8.6 ± 7.5 Bone: -25.8 ± 39.7 Hybrid Atlas Body: -2 ± 9 Soft tissue: 4.1 ± 8.5 Bone: -78.0 ± 35.3 cGAN Body: -2.5 ± 5.8 Soft tissue: -0.7 ± 6.3 Bone: -46.0 ± 19.6 | Bulk Density Body: 89.1 ± 7.7 Soft tissue: 76.1 ± 6.7 Bone: 244.1 ± 10 Tissue Class Density Body: 58.8 ± 10.4 Soft tissue: 44.6 ± 8.5 Bone: 228.2 ± 11.2 Hybrid Atlas Body: 57.4 ± 8.0 Soft tissue: 47.3 ± 7.8 Bone: 186.9 ± 17.9 cGAN Body: 34.7 ± 5.1 Soft tissue: 25.2 ± 3.4 Bone: 109.4 ± 12.3 | < 1.0% for all sCT methods | All sCT methods 3%/3 mm ≥ 99.8% All sCT methods 2%/2 mm ≥ 99.6% 1%/1 mm Bulk density: 93.6% Tissue Clase Density: 95.3% Hybrid Atlas: 94.8% cGAN: 97.3% | - | - |
| Gonzalez-Moya et al., 2021 [41] | Brain & Prostate | 16 Brain 31 Prostate | SIEMENS | - | - | - | < 0.5% | 1%/1 mm Brain: 73.2% Prostate: 84.7% | - | Translation: 3.5 mm Rotation: 0.2° |
| Farjam et al., 2019 [42] | Prostate | 23 | In House vs MRCAT | - | - | In House Body: 47 ± 5 Bone: 116 ± 12 Fat: 36 ± 6 Muscle: 47 ± 5 MRCAT Body: 65 ± 5 Bone: 172 ± 9 Fat: 43 ± 7 Muscle: 42 ± 4 |  | In House 3%/2 mm: 97.9% 2%/2 mm: 97.1% 1%/1 mm: 92.7% MRCAT 3%/2 mm: 97.9% 2%/2 mm: 97.1% 1%/1 mm: 90.6 | - | - |
| Maspero et al., 2018 [45] | Rectum | 20 | MRCAT | Body: 8 ± 6 | Body: 52 ± 3 | - | Mean: −0.3% | 2%/2 mm: 95.2% |  | Translation: <1 mm Rotation: <0.5° |
| Hsu et al., 2022 [44] | Prostate | 57 | cGAN | - | Body: -6.8 ± 3.6 Soft tissue: -1.2 ± 2.4 Bone: -128.0 ± 35.2 Air: 388.4 ± 131.3 | Body: 30.1 ± 4.2 Soft tissue: 19.6 ± 2.3 Bone: 158.5 ± 26.0 Air: 396.5 ± 126.1 | < 1% for PTV an OARs | 2%/2 mm: 99.9% 1%/1 mm: 99.0% | - | - |

Note. CUT: Contrastive Unpaired Translation Network

cycleCUT: CUT + cGAN