Supplementary Text S1.

The PAKAAR is controlled by a central microcontroller unit (OpenRB-150, Robotis, South Korea) that regulates actuation and data acquisition with a sampling interval of 20 ms. Anterior translation is generated by a linear actuator (L16-150, Actuonix Motion Devices, Canada) and monitored by a push-pull load cell (CBF30-100, CAS, South Korea). Rotational torque is generated by a rotational actuator (XM540-W270, Robotis Co., Ltd., South Korea) equipped with a 1:5 reduction gear module to increase maximal torque output. The applied torque is measured by a torque sensor (TQ-54T10-11HB8, CASKOREA, South Korea) with a repeatability of 0.1% RO. Both the load cell and torque sensor were calibrated using class M1 weights prior to experimentation (Supplementary Fig. S1).

Measurement precision is ensured by high-fidelity sensors. Anterior translation is detected by a laser displacement sensor (CD22-100A2, OPTEX FA Co., Ltd., Japan) with a repeat accuracy of 20 µm. Rotational angles of the tibia and femur are measured by wireless sensor units, each consisting of a gyroscopic sensor (GY-521, InvenSense, USA; repeatability 0.00763 deg/s) and a Bluetooth-enabled microcontroller (ESP-WROOM-32, Espressif, China).

Supplementary Text S2.

Intrarater reliability was assessed using intraclass correlation coefficient (ICC) (3,1), based on a two-way mixed-effects model (consistency, single measurement), appropriate for evaluating repeated measurements by a single fixed rater across subjects. The standard error of measurement (SEM) and minimal detectable change (MDC) were calculated using the following formulas: SEM=SD×1-ICC, and MDC=1.96×2×SEM, where SD represents the standard deviation. Interrater reliability was assessed using ICC (2,1), based on a two-way random-effects model (absolute agreement, single measurement), treating both raters and limbs as random samples from a larger population.

Supplementary Table S2. Intrarater and interrater reliability of PAKAAR and Lachmeter for anterior translation across ligamentous conditions

|  |  | ICC (3,1) [95% CI] |  | SEM |  | MDC |  | P valueb |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Ligament condition | Device | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Rater 1 | Rater 2 | ICC (2,1) [95% CI] | P valuec | r | P valued | P valuee |
| Intact | PAKAAR | 0.967 [0.94; 0.98] | 0.951 [0.91; 0.98] | 0.537 | 0.641 | 1.488 | 1.776 | .410 | .603 | 0.982 [0.83; 1.00] | .609 | 0.958 | <.001 | <.001 |
|  | Lachmeter | 0.880 [0.65; 0.98] | 0.889 [0.67; 0.98] | 0.673 | 0.716 | 1.865 | 1.984 |  |  | 0.960 [0.77; 0.99] |  | 0.861 | <.001 |  |
| sACL | PAKAAR | 0.952 [0.91; 0.98] | 0.960 [0.93; 0.98] | 0.85 | 0.775 | 2.357 | 2.148 | .309 | .263 | 0.964 [0.77; 0.99] | .336 | 0.95 | <.001 | <.001 |
|  | Lachmeter | 0.769 [0.43; 0.96] | 0.774 [0.44; 0.96] | 0.915 | 0.805 | 2.536 | 2.232 |  |  | 0.837 [0.31; 0.97] |  | 0.685 | <.001 |  |
| sACL + sALL | PAKAAR | 0.949 [0.91; 0.98] | 0.937 [0.89; 0.97] | 0.957 | 1.075 | 2.653 | 2.979 | .656 | .308 | 0.980 [0.48; 1.00] | .141 | 0.952 | <.001 | <.001 |
|  | Lachmeter | 0.898 [0.69; 0.98] | 0.708 [0.33; 0.94] | 0.855 | 0.838 | 2.369 | 2.323 |  |  | 0.797 [0.05; 0.97] |  | 0.773 | <.001 |  |
| Injured | PAKAAR | 0.950 [0.92; 0.97] | 0.947 [0.92; 0.97] | 0.91 | 0.949 | 2.523 | 2.629 | .225 | .090 | 0.986 [0.88; 1.00] | .004 | 0.951 | <.001 | <.001 |
|  | Lachmeter | 0.851 [0.69; 0.95] | 0.761 [0.54; 0.91] | 0.899 | 0.826 | 2.491 | 2.290 |  |  | 0.810 [0.15; 0.97] |  | 0.725 | <.001 |  |
| All | PAKAAR | 0.964 [0.95; 0.98] | 0.959 [0.94; 0.97] | 0.812 | 0.863 | 2.251 | 2.393 | .352 | .294 | 0.988 [0.81; 1.00] | .006 | 0.964 | <.001 | <.001 |
|  | Lachmeter | 0.931 [0.87; 0.97] | 0.915 [0.84; 0.96] | 0.838 | 0.812 | 2.322 | 2.250 |  |  | 0.913 [0.51; 0.99] |  | 0.888 | <.001 |  |

aICC, intraclass correlation coefficient; CI, confidence interval; MDC, minimal detectable change; r, Pearson’s correlation coefficient; sACL, sectioned anterior cruciate ligament; sALL, sectioned anterolateral ligament; PAKAAR, portable automated knee arthrometer for measuring anterior translation, central axial rotation, and medial pivotal rotation of tibia

bComparison of intrarater ICCs (3,1) between PAKAAR and Lachmeter using Fisher’s z-transformation and z-tests

cComparison of interrater ICCs (2,1) between PAKAAR and Lachmeter using Fisher’s z-transformation and z-tests

dInterrater correlation using Pearson’s correlation coefficients

eComparison of Pearson’s correlation coefficients between PAKAAR and Lachmeter using Fisher’s z-transformation and z-tests

fThe injured condition comprised both sACL and sACL + sALL.

Supplementary Table S3. Intrarater and interrater reliability of PAKAAR for central axial internal rotation across measurement site and ligamentous condition

|  |  | Intrarater ICC (3,1) [95% CI] |  | SEM |  | MDC |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Measurement site | Ligament condition | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Interrater ICC (2,1) [95% CI] | Inter -rater r | P valueb |
| Bone | Intact | 0.953 [0.89; 0.98] | 0.982 [0.96; 0.99] | 0.621 | 0.410 | 1.721 | 1.136 | 0.985 [0.83; 1.00] | 0.975 | <.001 |
|  | sACL | 0.940 [0.87; 0.97] | 0.963 [0.92; 0.98] | 0.670 | 0.597 | 1.858 | 1.654 | 0.967 [0.93; 0.99] | 0.946 | <.001 |
|  | sACL + sALL | 0.968 [0.93; 0.99] | 0.945 [0.88; 0.98] | 0.679 | 0.808 | 1.883 | 2.240 | 0.971 [0.93; 0.99] | 0.953 | <.001 |
|  | Injuredc | 0.964 [0.94; 0.98] | 0.956 [0.92; 0.98] | 0.680 | 0.744 | 1.885 | 2.061 | 0.974 [0.95; 0.99] | 0.954 | <.001 |
|  | All | 0.968 [0.95; 0.98] | 0.968 [0.95; 0.98] | 0.667 | 0.653 | 1.849 | 1.811 | 0.980 [0.97; 0.99] | 0.964 | <.001 |
| Skin | Intact | 0.933 [0.85; 0.97] | 0.965 [0.92; 0.98] | 2.318 | 1.831 | 6.425 | 5.076 | 0.961 [0.91; 0.98] | 0.938 | <.001 |
|  | sACL | 0.928 [0.84; 0.97] | 0.928 [0.84; 0.97] | 2.612 | 2.584 | 7.240 | 7.161 | 0.946 [0.88; 0.98] | 0.911 | <.001 |
|  | sACL + sALL | 0.959 [0.91; 0.98] | 0.945 [0.88; 0.98] | 2.177 | 2.794 | 6.034 | 7.746 | 0.973 [0.94; 0.99] | 0.954 | <.001 |
|  | Injuredc | 0.951 [0.91; 0.97] | 0.944 [0.90; 0.97] | 2.393 | 2.717 | 6.633 | 7.531 | 0.965 [0.94; 0.98] | 0.942 | <.001 |
|  | All | 0.950 [0.92; 0.97] | 0.952 [0.92; 0.97] | 2.367 | 2.479 | 6.561 | 6.872 | 0.966 [0.95; 0.98] | 0.945 | <.001 |

aICC, intraclass correlation coefficient; CI, confidence interval; SEM, standard error of measurement; MDC, minimal detectable change; r, Pearson’s correlation coefficient; sACL, sectioned anterior cruciate ligament; sALL, sectioned anterolateral ligament

bInterrater correlation using Pearson’s correlation coefficients

cThe injured condition comprised both sACL and sACL + sALL.

Supplementary Table S4. Regression and Correlation Between Skin-Based and Bone-Based central axial internal rotation with and without cadaver-wise Adjustment

|  |  | Regression and correlation without cadaver-wise adjustment |  |  |  | Cadaver-wise adjusted regression and correlation |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Ligamentous condition | R2 | Regression equation | r | P valueb | R2 | r | P valueb |
| Tibial CAIR | Intact | 0.841 | Bone = 0.67 × Skin + 0.51 | 0.917 | <.001 | 0.931 | 0.965 | <.001 |
|  | sACL | 0.775 | Bone = 0.62 × Skin + 1.17 | 0.880 | <.001 | 0.910 | 0.954 | <.001 |
|  | sACL + sALL | 0.840 | Bone = 0.64 × Skin + 0.26 | 0.917 | <.001 | 0.938 | 0.969 | <.001 |
|  | Injuredc | 0.829 | Bone = 0.63 × Skin + 0.80 | 0.911 | <.001 | 0.909 | 0.954 | <.001 |
|  | All | 0.838 | Bone = 0.63 × Skin + 0.88 | 0.915 | <.001 | 0.912 | 0.955 | <.001 |
| Femoral CAIR | Intact | 0.905 | Bone = 3.55 × Skin + 0.78 | 0.951 | <.001 | 0.924 | 0.961 | <.001 |
|  | sACL | 0.800 | Bone = 3.88 × Skin + 0.07 | 0.895 | <.001 | 0.938 | 0.968 | <.001 |
|  | sACL + sALL | 0.892 | Bone = 3.55 × Skin - 0.27 | 0.944 | <.001 | 0.971 | 0.985 | <.001 |
|  | Injuredc | 0.851 | Bone = 3.53 × Skin + 0.29 | 0.923 | <.001 | 0.940 | 0.969 | <.001 |
|  | All | 0.865 | Bone = 3.53 × Skin + 0.49 | 0.930 | <.001 | 0.914 | 0.956 | <.001 |
| Net CAIR | Intact | 0.572 | Bone = 0.24 × Skin + 2.10 | 0.756 | <.001 | 0.941 | 0.970 | <.001 |
|  | sACL | 0.689 | Bone = 0.25 × Skin + 3.04 | 0.830 | <.001 | 0.898 | 0.948 | <.001 |
|  | sACL + sALL | 0.451 | Bone = 0.22 × Skin + 5.03 | 0.672 | <.001 | 0.886 | 0.941 | <.001 |
|  | Injuredc | 0.594 | Bone = 0.25 × Skin + 3.59 | 0.771 | <.001 | 0.860 | 0.927 | <.001 |
|  | All | 0.612 | Bone = 0.26 × Skin + 2.60 | 0.783 | <.001 | 0.833 | 0.913 | <.001 |

ar, Pearson’s correlation coefficient; sACL, sectioned anterior cruciate ligament; sALL, sectioned anterolateral ligament; CAIR, central axial internal rotation

bCorrelation between skin-based and bone-based CAIR using Pearson’s correlation coefficients

cThe injured condition comprised both sACL and sACL + sALL.

Supplementary Table S5. Intrarater and interrater reliability of PAKAAR for medial pivotal rotation across measurement site and ligamentous condition

|  |  | Intrarater ICC (3,1) [95% CI] |  | SEM |  | MDC |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Measurement site | Ligament condition | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Rater 1 | Rater 2 | Interrater ICC (2,1) [95% CI] | Inter -rater r | P valueb |
| Bone | Intact | 0.935 [0.86; 0.97] | 0.969 [0.93; 0.99] | 0.630 | 0.477 | 1.747 | 1.323 | 0.977 [0.92; 0.99] | 0.963 | <.001 |
|  | sACL | 0.956 [0.90; 0.98] | 0.972 [0.94; 0.99] | 0.833 | 0.742 | 2.308 | 2.057 | 0.983 [0.96; 0.99] | 0.968 | <.001 |
|  | sACL + sALL | 0.968 [0.93; 0.99] | 0.964 [0.92; 0.98] | 0.694 | 0.850 | 1.924 | 2.356 | 0.955 [0.9; 0.98] | 0.946 | <.001 |
|  | Injuredc | 0.966 [0.94; 0.98] | 0.971 [0.95; 0.98] | 0.774 | 0.807 | 2.146 | 2.238 | 0.972 [0.95; 0.98] | 0.962 | <.001 |
|  | All | 0.976 [0.96; 0.98] | 0.977 [0.96; 0.99] | 0.736 | 0.763 | 2.039 | 2.116 | 0.983 [0.97; 0.99] | 0.974 | <.001 |
| Skin | Intact | 0.903 [0.79; 0.96] | 0.948 [0.88; 0.98] | 2.983 | 2.069 | 8.269 | 5.735 | 0.978 [0.89; 0.99] | 0.950 | <.001 |
|  | sACL | 0.969 [0.93; 0.99] | 0.963 [0.92; 0.98] | 1.851 | 1.922 | 5.130 | 5.328 | 0.970 [0.93; 0.99] | 0.953 | <.001 |
|  | sACL + sALL | 0.916 [0.82; 0.96] | 0.951 [0.89; 0.98] | 3.020 | 2.195 | 8.370 | 6.083 | 0.989 [0.97; 1.0] | 0.955 | <.001 |
|  | Injuredc | 0.942 [0.9; 0.97] | 0.956 [0.92; 0.98] | 2.533 | 2.102 | 7.022 | 5.828 | 0.979 [0.96; 0.99] | 0.955 | <.001 |
|  | All | 0.935 [0.9; 0.96] | 0.959 [0.93; 0.97] | 2.691 | 2.084 | 7.459 | 5.777 | 0.981 [0.97; 0.99] | 0.956 | <.001 |

aICC, intraclass correlation coefficient; CI, confidence interval; SEM, standard error of measurement; MDC, minimal detectable change; r, Pearson’s correlation coefficient; sACL, sectioned anterior cruciate ligament; sALL, sectioned anterolateral ligament

bInterrater correlation using Pearson’s correlation coefficients

cThe injured condition comprised both sACL and sACL + sALL.

Supplementary Table S6. Regression and correlation between skin- and bone-referenced medial pivotal rotation with and without cadaver-wise adjustment

|  |  | Regression and correlation without cadaver-wise adjustment |  |  |  | Cadaver-wise adjusted regression and correlation |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Ligamentous condition | R2 | Regression equation | r | P valueb | R2 | r | P valueb |
| Tibial MPR | Intact | 0.777 | Bone = 0.71 × Skin - 1.97 | 0.882 | <.001 | 0.918 | 0.958 | <.001 |
|  | sACL | 0.882 | Bone = 0.70 × Skin - 2.01 | 0.939 | <.001 | 0.963 | 0.982 | <.001 |
|  | sACL + sALL | 0.916 | Bone = 0.74 × Skin - 2.69 | 0.957 | <.001 | 0.946 | 0.973 | <.001 |
|  | Injuredc | 0.902 | Bone = 0.73 × Skin - 2.48 | 0.950 | <.001 | 0.944 | 0.972 | <.001 |
|  | All | 0.872 | Bone = 0.72 × Skin - 2.30 | 0.934 | <.001 | 0.921 | 0.959 | <.001 |
| Femoral MPR | Intact | 0.665 | Bone = 3.74 × Skin + 0.95 | 0.816 | <.001 | 0.886 | 0.941 | <.001 |
|  | sACL | 0.899 | Bone = 3.77 × Skin + 0.08 | 0.948 | <.001 | 0.961 | 0.980 | <.001 |
|  | sACL + sALL | 0.854 | Bone = 2.42 × Skin + 1.21 | 0.924 | <.001 | 0.947 | 0.973 | <.001 |
|  | Injuredc | 0.800 | Bone = 2.69 × Skin + 1.38 | 0.894 | <.001 | 0.816 | 0.903 | <.001 |
|  | All | 0.698 | Bone = 2.94 × Skin + 1.44 | 0.835 | <.001 | 0.751 | 0.867 | <.001 |
| Net MPR | Intact | 0.393 | Bone = 0.17 × Skin + 0.73 | 0.627 | <.001 | 0.937 | 0.968 | <.001 |
|  | sACL | 0.583 | Bone = 0.31 × Skin + 0.96 | 0.763 | <.001 | 0.961 | 0.980 | <.001 |
|  | sACL + sALL | 0.755 | Bone = 0.36 × Skin + 1.87 | 0.869 | <.001 | 0.911 | 0.954 | <.001 |
|  | Injuredc | 0.652 | Bone = 0.35 × Skin + 1.03 | 0.808 | <.001 | 0.809 | 0.900 | <.001 |
|  | All | 0.574 | Bone = 0.36 × Skin - 0.38 | 0.758 | <.001 | 0.715 | 0.846 | <.001 |

ar, Pearson’s correlation coefficient; sACL, sectioned anterior cruciate ligament; sALL, sectioned anterolateral ligament; MPR, medial pivotal rotation

bCorrelation between skin-based and bone-based CAIR using Pearson’s correlation coefficients

cThe injured condition comprised both sACL and sACL + sALL.

Supplementary Fig. S1 Logarithmic calibration curves for (A) the load cell and (B) the torque sensor, both calibrated using a set of class M1 weights.