Supplementary materials

(1)Proteomics Analysis:

### Sample Preparation:

### For each sample, 2 μg of total peptides were separated and analyzed.Protein extraction was performed using a lysis buffer with 4% SDS and 100 mM DTT, followed by sonication.

### Instrument:

Proteomics analysis was conducted using a nano-UPLC (EASY-nLC1200) coupled to a Q Exactive HFX Orbitrap instrument (Thermo Fisher Scientific) with a nano-electrospray ion source.

### LC Parameters:

Chromatographic separation was performed on a nano C18 column, with a gradient elution of water with 0.1% formic acid (A) and acetonitrile with 0.1% formic acid (B).

### Gradient Program:

The gradient elution program was as follows:

- 0–5 min: 5% B

- 5–35 min: 5–35% B

- 35–40 min: 35–80% B

### Mass Spectrometry Parameters:

The MS/MS analysis was performed using a Q Exactive HFX Orbitrap in data-dependent acquisition (DDA) mode and positive ionization mode.

- Resolution: 60,000 for MS1 and 15,000 for MS2.

- Scan range: m/z 300–2000.

- Collision energy: 30 eV for higher-energy collisional dissociation (HCD).

- Spray voltage: 3.2 kV for both positive and negative modes.

- Sheath gas flow rate: 10–40 arb.

- Capillary temperature: 320°C.

### Software and Data Processing:

Raw MS files were processed using Proteome Discoverer (PD) software (Version 2.4.0.305) with the built-in Sequest HT search engine. The MS spectra were searched against the UniProt FASTA database for Rattus norvegicus (uniprot-Rattus norvegicus-10116-2021-8.fasta).

- Fixed modifications: Carbamidomethyl [C], TMT Pro (K), TMT Pro (N-term).

- Variable modifications: Oxidation (M), Acetyl (Protein N-term).

- Peptide identification: Performed with an initial precursor mass deviation of up to 10 ppm and a fragment mass deviation of 0.02 Da.

### Quantification and Data Normalization:

- Unique peptides and Razor peptides were used for protein quantification, while the total peptide amount was used for normalization.

- A 1.5-fold change was set as the threshold for significance, determined by analyzing inner-quartile data from control experiments using ln-ln plots. The Pierson’s correlation coefficient (R) was 0.98, and 95–99% of the normalized intensities fell between the set fold changes.

- Both t-tests and fold change thresholds needed to pass to consider the protein significant.

### Multivariate Analyses:

- 2D Hierarchical Cluster Analysis (HCA) HeatMaps and Principal Component Analysis (PCA) plots were performed using Qlucore Omics Explorer (Qlucore, Lund Sweden).

- Gene ontology assignments and pathway analysis were executed using MetaCore (GeneGO Inc., St. Joseph, MI, USA). Interactions identified in MetaCore were manually validated through full-text articles.

(2)Metabolomics Detection

·  Instrument: Thermo Scientific Vanquish UHPLC-Q Exactive system and Thermo Q Precision mass spectrometer

·  Column: Hyperil Gold C18 column (2.1 mm × 100 mm, 1.9 μm)

·  Mobile Phases: Not specified directly, but based on common practices, it is likely to be water with 0.1% formic acid (A) and acetonitrile with 0.1% formic acid (B)

·  Gradient Elution: Not mentioned in the text, you may need to retrieve the specific gradient from experimental records (it usually involves increasing the percentage of B over time)

·  Flow Rate: 0.3 mL/min

·  Injection Volume: 5 μL

Ionization Mode: Positive and negative ionization modes

·  Mass Spectrometry Parameters:

- Capillary voltage: Spray voltage of 3.2 kV for both positive and negative modes

- Cone voltage: 30 V

- Source temperature: Capillary temperature of 320°C (likely equivalent to source temperature)

- Mass range: m/z 100–1200

- Sheath gas flow rate: Set between 10–40 arb

- Data Acquisition: HCD scan used for data-dependent acquisition (DDA) MS/MS experiments

- Dynamic Exclusion: Applied to remove redundant information in MS/MS spectra