Additional file

A novel micro-ECoG recording method for recording multisensory neural activity from the parietal to temporal cortices in mice

Susumu Setogawa, Ryota Kanda, Shuto Tada, Takuya Hikima, Yoshito Saitoh, Mikiko Ishikawa, Satoshi Nakada, Fumiko Seki, Keigo Hikishima, Hideyuki Matsumoto, Kenji Mizuseki, Osamu Fukayama, Makoto Osanai, Hiroto Sekiguchi*, and Noriaki Ohkawa*

*Correspondence: nohkawa@dokkyomed.ac.jp or sekiguchi@ee.tut.ac.jp

This file includes:

Supplementary Method

Fig. S1

Supplementary Method

Channel crosstalk

The crosstalk between electrode channels was investigated using an equivalent circuit model as shown in Fig. S1a. This circuit models two adjacent recording channels, where Vi1 and Vi2 are input voltages of each channel, Vo1 and Vo2 are output voltages of each channel, Ze is electrode-electrolyte impedance, Zc is coupling capacitance impedance, Zp is parasitic capacitance impedance to ground, and ZA is input impedance of amplifier, respectively. The cross talk is evaluated by the ratio of the signal on the input voltage Vi1 to the output voltage Vo2 of the adjacent channel when Vi2=0. In an ideal circuit, since Zp and ZA can be removed as infinite, the crosstalk can be easily evaluated as expressed in the following equation:

Vo2Vi1=Ze2Ze+Zc

Next, we consider the value of Zc from the design of the mECoG sheet. The coupling capacitance Cc can be expressed:

Cc=εptLs

where εp is the permittivity of dielectric (parylene C), t and L are the thickness (250 nm) and length (15 mm) of metal wiring, and s is spacing (20 mm) between the metal wiring layer, respectively. Based on the device parameters, the Cc is calculated to be 5.1 fF, which corresponds to the impedance Zc of 31 GW at 1 kHz. The average Ze in the mECoG sheet as shown in Fig. 2d was 0.5 MW. Therefore, the crosstalk Vo2/Vi1 is calculated to be 1.61×10-5 (­48 dB). There is almost no influence of channel crosstalk in the developed mECoG sheet. Additionally, the effect of spacing between electrode channels on the calculated crosstalk for Ze of 0.5 and 5 MW was calculated as shown in Fig. S1b. The crosstalk increases with decreasing spacing and impedance. Therefore, the mECoG sheet with low impedance electrode and wide spacing are more effective for ECoG measurement.

Fig. S1. Channel crosstalk. (a) Equivalent circuit model representation of two adjacent recording channels with electrode-electrolyte impedance Ze, coupling capacitance impedance Zc, parasitic capacitance impedance to ground Zp, and input impedance of amplifier ZA. (b) Crosstalk as a function of spacing between two electrode channels at different electrode impedance Ze of 0.5 and 5 MW with thickness t of 0.25 mm, length L of 15 mm of metal wiring layer, and spacing s of 20 mm.