Fig S1: Chromatographic column Luna С8 (100 mm х 4.6 mm 3 μm) and mobile phase - ACN-MetOH-0.5% TEA pH 3.06 (35-30-35)

Fig S2: Chromatographic column Luna С8 (100 mm х 4.6 mm 3 μm) and mobile phase - ACN-MetOH-0.7% TEA pH 2.8 (35-35-30)

Fig S3: Chromatographic column Luna С8 (100 mm х 4.6 mm 3 μm) and mobile phase - ACN-MetOH-0.7% TEA pH 3.06 (30-35-35)

Fig S4: Chromatographic column Zorbax SB-Phenyl (150 mm х 4.6 mm 3.5 μm)

ACN-MetOH-0.7% TEA pH 3.06 (30-35-35)

Fig S5: Chromatographic column Zorbax SB-C8 (150 mm х 4.6 mm 3.5 μm)

ACN-MetOH-0.7% TEA pH 3.06 (30-35-35)

Fig S6: Chromatographic column Luna С18 (100 mm х 4.6 mm 3 μm)

ACN-MetOH-0.7% TEA pH 3.06 (30-35-35)

Fig S7: Chromatographic column InertClone ODS (150 mm х 4.6 mm 3 μm)

ACN-MetOH-0.7% TEA pH 3.06 (30-35-35)

Fig S8: Pareto chart of the standardized effects and residual plots for NTP

Fig S9: Pareto chart of the standardized effects and residual plots for TF

Fig S10: Pareto chart of the standardized effects and residual plots for Rs

Fig S11: Pareto chart of the standardized effects and residual plots for Rt

Table S1: Preliminary studies in HPLC method development

| Chromatographic condition Parameters | ACN-MeOH-0.7 % TEA pH 3.06 (35:35:30) T 30° C, 1 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 30° C, 1 ml/min | ACN-MeOH-0.7 % TEA pH 3.06 (35:30:35) T 30° C, 1 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 28° C, 1 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 32° C, 1 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 30° C, 0.8 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 30° C, 1.2 ml/min | ACN-MeOH-0.7 %TEA pH 2.71 (30:35:35) T 32° C, 1.2 ml/min | ACN-MeOH-0.7 %TEA pH 2.91 (30:35:35) T 32° C, 1.2 ml/min | ACN-MeOH-0.7 %TEA pH 3.06 (30:35:35) T 32° C, 1.2 ml/min | Acceptability criteria |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Tailing F | AML- 1.37 NIF- 0 LER- 0 NIT- 1.27 NIM- 1.22 | AML- 1.38 NIF- 1.41 LER- 1.24 NIT- 1.20 NIM- 1.17 | AML- 1.24 NIF- 0 LER- 0 NIT- 0 NIM- 0 | AML- 1.35 NIF- 1.35 LER- 1.23 NIT- 1.22 NIM- 1.18 | AML- 1.39 NIF- 1.45 LER- 1.24 NIT- 1.22 NIM- 1.18 | AML- 1.37 NIF- 1.41 LER- 1.23 NIT- 1.17 NIM- 1.15 | AML- 1.35 NIF- 1.25 LER- 1.39 NIT- 1.16 NIM- 1.2 | AML- 1.41 NIF- 1.45 LER- 1.23 NIT- 1.21 NIM- 1.18 | AML- 1.35 NIF- 1.25 LER- 1.39 NIT- 1.16 NIM- 1.2 | AML- 1.43 NIF- 1.46 LER- 1.23 NIT- 1.22 NIM- 1.17 | 0.8-1.5 |
| Resolution | AML- NIF- 5.39 LER- 1.49 NIT- 6.73 NIM- 4.18 | AML- NIF- 6.46 LER- 5.48 NIT- 6.70 NIM- 6.37 | AML- NIF- 10.426 LER- 0 NIT- 0 NIM- 0 | AML- NIF- 6.43 LER- 5.90 NIT- 6.52 NIM- 6.45 | AML- NIF- 6.45 LER- 5.47 NIT- 6.54 NIM- 6.27 | AML- NIF- 6.82 LER- 6.03 NIT- 6.86 NIM- 6.67 | AML- NIF- 6.08 LER- 5.04 NIT- 6.25 NIM- 5.99 | AML- NIF- 6.04 LER- 5.20 NIT- 6.17 NIM- 5.96 | AML- NIF- 6.02 LER- 4.76 NIT- 6.50 NIM- 5.94 | AML- NIF- 5.96 LER- 5.63 NIT- 5.63 NIM- 5.95 | > 2.0 |
| K’ («died» volume 1.278) | AML- 0.87 NIF- 1.48 LER- 1.66 NIT- 2.51 NIM- 3.08 | AML- 1.15 NIF- 2.11 LER-2.89 NIT- 3.95 NIM- 5.06 | AML- 0.12 NIF- 1.97 LER- 2.24 NIT- 3.05 NIM- 3.18 | AML- 1.41 NIF- 2.16 LER- 3.03 NIT- 4.08 NIM- 5.25 | AML- 1.27 NIF- 2.08 LER- 2.84 NIT- 3.84 NIM- 4.92 | AML- 1.85 NIF- 2.86 LER- 3.87 NIT- 5.14 NIM- 6.53 | AML- 0.87 NIF- 1.55 LER- 2.18 NIT- 3.02 NIM- 3.91 | AML- 0.89 NIF- 1.56 LER- 2.20 NIT- 3.03 NIM- 3.92 | AML- 0.89 NIF- 1.54 LER- 2.13 NIT- 3.00 NIM- 3.89 | AML- 0.88 NIF- 1.54 LER- 2.13 NIT- 3.00 NIM- 3.89 | 1-10 |
| Retention time, min | AML- 2.40 NIF- 3.17 LER- 3.40 NIT- 4.48 NIM- 5.20 | AML- 2.93 NIF- 3.98 LER- 4.98 NIT- 6.32 NIM- 7.75 | AML- 1.43 NIF- 3.79 LER- 4.14 NIT- 5.18 NIM- 5.34 | AML- 2.99 NIF- 4.04 LER- 5.15 NIT- 6.49 NIM- 7.99 | AML- 2.90 NIF- 3.93 LER- 4.91 NIT- 6.19 NIM- 7.56 | AML- 3.64 NIF- 4.93 LER- 6.22 NIT- 7.85 NIM- 9.62 | AML- 2.40 NIF- 3.26 LER- 4.06 NIT- 5.14 NIM- 6.28 | AML- 2.42 NIF- 3.27 LER- 4.09 NIT- 5.15 NIM- 6.29 | AML- 2.41 NIF- 3.25 LER- 4.00 NIT- 5.12 NIM- 6.25 | AML- 2.40 NIF- 3.25 LER- 4.00 NIT- 5.12 NIM- 6.25 | - |
| NTP | AML- 4849 NIF- 7237 LER- 7374 NIT- 11674 NIM- 13069 | AML- 5936 NIF- 8394 LER- 10485 NIT- 14951 NIM- 16404 | AML- 553 NIF- 5209 LER- 0 NIT- 15457 NIM- 0 | AML- 6034 NIF- 8410 LER- 10618 NIT- 14906 NIM- 16241 | AML- 6073 NIF- 8586 LER- 10634 NIT- 15044 NIM- 16565 | AML- 6673 NIF- 9646 LER- 11714 NIT- 16328 NIM- 17898 | AML- 5247 NIF- 7580 LER- 9317 NIT- 13474 NIM- 14925 | AML- 5365 NIF- 7625 LER- 9436 NIT- 13399 NIM- 14838 | AML- 5280 NIF- 7541 LER- 9291 NIT- 13332 NIM- 14840 | AML- 5290 NIF- 7602 LER- 9612 NIT- 13369 NIM- 14889 | >2000 |

Table S2: The choice of injection volume

| Injection volume, µL | 3 | 4 | 5 | Acceptability criteria |
| --- | --- | --- | --- | --- |
| Tailing F | AML- 1.39 NIF- 1.46 LER- 1.22 NIT- 1.21 NIM- 1.16 | AML- 1.39 NIF- 1.38 LER- 1.24 NIT- 1.22 NIM- 1.17 | AML- 1.40 NIF- 1.43 LER- 1.25 NIT- 1.21 NIM- 1.17 | 0.8-1.5 |
| Resolution | AML- NIF- 5.95 LER- 5.44 NIT- 6.03 NIM- 6.02 | AML- NIF- 5.85 LER- 4.91 NIT- 6.29 NIM- 5.93 | AML- NIF- 5.73 LER- 4.54 NIT- 6.26 NIM- 5.79 | >2.0 |
| NTP | AML- 5595 NIF- 7836 LER- 9773 NIT- 13614 NIM- 15065 | AML- 5233 NIF- 7181 LER- 9174 NIT- 13401 NIM- 14910 | AML- 4958 NIF- 6712 LER- 8437 NIT- 12667 NIM- 14083 | >5000 |

Table S3: Risk assessment: Critical method variables vs HPLC method responses

| No | Input method variables (Xn) | Responses (Yn) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  | Consideration |
|  |  | TF | tR | Rs | NTP | A |  |
| 1 | TEA (%V/V) | H | H | H | H | M | Optimized |
| 2 | pH | H | H | H | H | M | Optimized |
| 3 | MeOH (%V/V) | H | H | H | H | M | Optimized |
| 4 | Buffer concentration | H | M | M | M | M | Controlled |
| 5 | Column temperature | M | M | M | M | M | Controlled |
| 6 | Column make | H | H | H | H | H | Default |
| 7 | Column type | H | H | H | H | H | Controlled |
| 8 | Buffer type | H | H | H | H | H | Controlled |
| 9 | Column length | M | M | M | M | M | Default |
| 10 | Column particle size | M | M | M | M | M | Default |
| 11 | Detector | N/A | N/A | N/A | N/A | N/A | Default |
| 12 | Injection volume | L | L | L | L | L | Default |
| 13 | Reagent purity | L | L | L | L | L | Controlled |
| 14 | Vendors (reagents) | L | L | L | L | L | Controlled |