Manuscript: Performance and usability evaluation of three LDH-based malaria rapid diagnostic tests in Kédougou, Senegal
Supplementary Table 2. Label comprehension questionnaire results for the Pf (pLDH) and Pf (pLDH/HRPII) tests.
	Question
	Correct responses
n (%)

	
	BIOCREDIT
Pf (pLDH) Test
	BIOCREDIT Pf (pLDH/HRPII) Test

	True or false: the test can be used to detect infection with Pf parasite causing malaria in humans.
	10 (100.0%)
	16 (100.0%)

	What does the test measure?
	10 (100.0%)
	13 (81.3%)

	Which plasmodium antigen can be detected using the test?
	9 (90.0%)
	15 (93.8%)

	The test can be used with which type(s) of samples?
	7 (70.0%)
	16 (100.0%)

	At what temperature should the test kit be stored?
	10 (100.0%)
	15 (93.8%)

	When should you apply the assay buffer to the test device?
	10 (100.0%)
	16 (100.0%)

	How much blood is required to run the test?
	7 (70.0%)
	12 (75.0%)

	How many drops of assay buffer should you add to the test device?
	5 (50.0%)
	10 (62.5%)

	How many test lines can you see on the device (including control line)?
	9 (90.0%)
	16 (100.0%)

	How long should you wait to interpret the test results?
	10 (100.0%)
	16 (100.0%)

	True or false: You can read the test result after 35 minutes.
	7 (70.0%)
	8 (50.0%)

	True or false: with the presence of any test line, no matter how faint, the result is considered positive
	8 (80.0%)
	15 (93.8%)

	Can you re-use the test device?
	10 (100.0%)
	16 (100.0%)

	Where should the blood sample be applied to the test device?
	10 (100.0%)
	16 (100.0%)

	Where should the assay buffer be applied to the test device?
	10 (100.0%)
	16 (100.0%)



	Key

	>85% 

	70%-85% 

	<70% 




