Supplementary Information
Additional file 2: Table S1. Comparative results between the MiLab™ automated microscopy platform and the thin blood smear microscopy (n=400). Table S2. Comparative results between the thin blood smear microscopy and the NM-PCR (n=400).

Additional file 2: Table S1. 
	
	 
	Thin blood smear microscopy

	
	 
	Positive Pf
	Positive Pv
	Positive Po
	Positive Pm
	Positive Pf+Po
	Negative
	Hemolysis

	miLabTM platform1
	Pf positive suspected
	144
	0
	3
	2
	0
	15
	1

	
	Pv positive suspected
	6
	1
	0
	1
	0
	1
	0

	
	Pf/Pv positive suspected 
	0
	0
	2
	0
	1
	0
	0

	
	Po positive
	0
	0
	2
	0
	0
	1
	0

	
	Pm positive
	0
	0
	0
	1
	0
	0
	0

	
	Plasmodium spp.
	0
	0
	0
	0
	0
	1
	0

	
	Negative
	8
	4
	1
	1
	0
	152
	2

	
	Error
	12
	0
	0
	0
	0
	13
	3

	
	Incomplete
	8
	1
	0
	0
	0
	9
	4

	
	Total samples
(n=400)
	178
	6
	8
	5
	1
	192
	10


1Categorised results after reclassification of ‘review needed’ results by an expert microscopist. Pf: P. falciparum, Pv: P. vivax, Po: P. ovale, Pm: P. malariae.



Additional file 2: Table S2. 
	
	
	NM-PCR

	
	
	Positive 
Pf
	Positive 
Pv
	Positive 
Po
	Positive 
Pm
	Mixed infection 
(Pf+Pm)
	Mixed infection
 (Pf+Po)
	Mixed infection 
(Pm+Po)
	Mixed infection 
(Pf+Pv)
	Negative

	Thin blood smear microscopy
	Positive Pf
	177
	0
	0
	0
	1
	0
	0
	0
	0

	
	Positive Pv
	0
	6
	0
	0
	0
	0
	0
	0
	0

	
	Positive Po
	0
	0
	7
	0
	0
	0
	1
	0
	0

	
	Positive Pm
	0
	0
	0
	3
	2
	0
	0
	0
	0

	
	Mixta Pf+Po
	0
	0
	0
	0
	0
	1
	0
	0
	0

	
	Negative
	98
	5
	6
	9
	1
	1
	0
	1
	71

	
	Hemolysis
	8
	0
	1
	1
	0
	0
	0
	0
	0

	
	Total samples
(n=400)
	283
	11
	14
	13
	4
	2
	1
	1
	71


Pf: P. falciparum, Pv: P. vivax, Po: P. ovale, Pm: P. malariae.

