Supplementary Information

Additional file 2: Table S1. Comparative results between the MiLab™ automated microscopy platform and the thin blood smear microscopy (n=400). Table S2. Comparative results between the thin blood smear microscopy and the NM-PCR (n=400).

Additional file 2: Table S1.

|  |  | Thin blood smear microscopy |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive Pf | Positive Pv | Positive Po | Positive Pm | Positive Pf+Po | Negative | Hemolysis |
| miLabTM platform1 | Pf positive suspected | 144 | 0 | 3 | 2 | 0 | 15 | 1 |
|  | Pv positive suspected | 6 | 1 | 0 | 1 | 0 | 1 | 0 |
|  | Pf/Pv positive suspected | 0 | 0 | 2 | 0 | 1 | 0 | 0 |
|  | Po positive | 0 | 0 | 2 | 0 | 0 | 1 | 0 |
|  | Pm positive | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
|  | Plasmodium spp. | 0 | 0 | 0 | 0 | 0 | 1 | 0 |
|  | Negative | 8 | 4 | 1 | 1 | 0 | 152 | 2 |
|  | Error | 12 | 0 | 0 | 0 | 0 | 13 | 3 |
|  | Incomplete | 8 | 1 | 0 | 0 | 0 | 9 | 4 |
|  | Total samples (n=400) | 178 | 6 | 8 | 5 | 1 | 192 | 10 |

1Categorised results after reclassification of ‘review needed’ results by an expert microscopist. Pf: P. falciparum, Pv: P. vivax, Po: P. ovale, Pm: P. malariae.

Additional file 2: Table S2.

|  |  | NM-PCR |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Positive Pf | Positive Pv | Positive Po | Positive Pm | Mixed infection (Pf+Pm) | Mixed infection (Pf+Po) | Mixed infection (Pm+Po) | Mixed infection (Pf+Pv) | Negative |
| Thin blood smear microscopy | Positive Pf | 177 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 |
|  | Positive Pv | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 |
|  | Positive Po | 0 | 0 | 7 | 0 | 0 | 0 | 1 | 0 | 0 |
|  | Positive Pm | 0 | 0 | 0 | 3 | 2 | 0 | 0 | 0 | 0 |
|  | Mixta Pf+Po | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 |
|  | Negative | 98 | 5 | 6 | 9 | 1 | 1 | 0 | 1 | 71 |
|  | Hemolysis | 8 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 |
|  | Total samples (n=400) | 283 | 11 | 14 | 13 | 4 | 2 | 1 | 1 | 71 |

Pf: P. falciparum, Pv: P. vivax, Po: P. ovale, Pm: P. malariae.