Appendix B: Minor Issues/Annoyances with Delica EMS 9D Robotic TCD System

Minor Issues

- Set-up Time – Intubated, sedated TBI patients with multi-modal monitoring, it takes ~30 minutes with one operator to set everything up safely, given the head-band must be disassembled and reassembled around the patient’s head, ensuring other invasive cranial monitoring, central access and endotracheal tube are not disturbed.

- C-collar – Given some manipulation of the patient’s head to secure the device in these complex patients, authors were uncomfortable placing in those with C-collar and uncertain stability status of the cervical spine.

- Decompressive Craniectomy – Given the probe holder is a head-band, secured circumferentially, this cannot be safely applied in those patients with a craniectomy.

- Headband Thickness – Head-band is ~1-1.5 inches thick, leaving only a small area to apply bifrontal NIRS pads in the presence of a triple bolt.

- Neoprene Pads - Though exchangeable, the pads take time to clean properly with disinfectant

- Headband Size – Despite being readily adjustable, we’ve found patients with small stature may have head circumferences that are too small for the manual straps and ratcheting system to tighten effectively.

C-collar = cervical collar, NIRS = near infrared spectroscopy.