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	Minor Issues

	1. Set-up Time – Intubated, sedated TBI patients with multi-modal monitoring, it takes ~30 minutes with one operator to set everything up safely, given the head-band must be disassembled and reassembled around the patient’s head, ensuring other invasive cranial monitoring, central access and endotracheal tube are not disturbed.

	2. C-collar – Given some manipulation of the patient’s head to secure the device in these complex patients, authors were uncomfortable placing in those with C-collar and uncertain stability status of the cervical spine.

	3. Decompressive Craniectomy – Given the probe holder is a head-band, secured circumferentially, this cannot be safely applied in those patients with a craniectomy.

	4. Headband Thickness – Head-band is ~1-1.5 inches thick, leaving only a small area to apply bifrontal NIRS pads in the presence of a triple bolt.

	5. Neoprene Pads -  Though exchangeable, the pads take time to clean properly with disinfectant 

	6. Headband Size – Despite being readily adjustable, we’ve found patients with small stature may have head circumferences that are too small for the manual straps and ratcheting system to tighten effectively. 


C-collar = cervical collar, NIRS = near infrared spectroscopy.

