Automated high-throughput quantification of plasma p-tau217 and APOE-ε4 for Alzheimer’s disease diagnosis and cognitive decline in a memory cohort

Jochen Titeca1, Marta Scarioni1,, Matthijs Oyaert2,$, Tim Van Langenhove1,$,#

- Cognitive Centre, Department of Neurology, Ghent University Hospital, Ghent, Belgium

- Department of Laboratory Medicine, Ghent University Hospital, Ghent, Belgium

$ Both authors contributed equally

Supplemental Table 1. Analytical performance (within-run, between-run, and within-laboratory imprecision, bias and total error) of the p-tau181 and p-tau217 (RUO) assays on the Roche cobas pro e801 determined by measuring Roche PC p-tau181 and Roche PC p-tau217 QC-solutions in a 5x5 protocol (CLSI E15-A3)

|  | Plasma p-tau181 |  | Plasma p-tau217 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Level 1 | Level 2 | Level 1 | Level 2 | Level 3 |  |
| Mean concentration (pg/mL) | 0.67 | 2.00 | 0.12 | 0.70 | 1.02 |  |
| Target (pg/mL) | 0.70 | 2.02 | 0.11 | 0.66 | 0.98 |  |
| Imprecision |  |  |  |  |  |  |
| SDwr (pg/mL) | 0.02 | 0.03 | 0.01 | 0.01 | 0.01 |  |
| CVwr (%) | 3.5 | 1.7 | 5.2 | 1.9 | 1.3 |  |
| SDbr (pg/mL) | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 |  |
| CVbr (%) | 1.6 | 0.5 | 6.8 | 3.1 | 0.5 |  |
| SDwl (pg/mL) | 0.03 | 0.04 | 0.01 | 0.03 | 0.01 |  |
| CVwl (%) | 3.8 | 1.7 | 8.6 | 3.6 | 1.4 |  |

Abbreviations: WR: within-run; BR: between-run, WL: within-lab, SD: standard deviation, CV: coefficient of variation, CVwr: within-run coefficient, Abs: absolute

Supplemental Table 2. Demographic, clinical and neurological characteristics of study cohort: other neurodegenerative disorders (ND) vs. psychiatric or psychological causes of cognitive complaints (PSY).

|  | ND (n=54) | PSY (n=30) | p-value | ES |
| --- | --- | --- | --- | --- |
| Age, median (IQR, Q1, Q3) | 71.5 (12, 65.0, 77.25) | 66.0 (11, 61.75, 72.25) | 0.012 | 0.27 |
| Male, n (%) | 63% | 66,7% | 0,734 | 0.04 |
| MoCA at baseline, median (IQR, Q1, Q3) | 19.5 (9, 15.25, 23.75) | 26 (4, 23, 27) | <0,001 | 0,47 |
| eGFR, median (IQR, Q1, Q3) | 87.0 (26, 67.5, 93.25) | 83.5 (22, 70.25, 92.0) | 0,978 | 0.003 |
| 2 or more lobar microbleeds on MRI | 9,4% | 3,4% | 0,32 | 0.11 |
| Fazekas grade 3, n (%) | 3,7% | 6,7% | 0,541 | 0.07 |
| Elecsys p-tau181, median (IQR, Q1, Q3) | 0.90 (0.338, 0.725, 1.063) | 0.763 (0.277, 0.664, 0.941) | 0,024 | 0.25 |
| Elecsys p-tau217, median (IQR, Q1, Q3) | 0.249 (0.123, 0.198, 0.322) | 0.237 (0.077, 0.199, 0.275) | 0.403 | 0,09 |
| Lumipulse p-tau217, median (IQR, Q1, Q3) | 0.097 (0.06, 0.079, 0.139) | 0.077 (0.048, 0.066, 0.114) | 0.017 | 0,09 |
| Lumipulse Aβ42, median (IQR, Q1, Q3) | 21.75 (5.8, 18.55, 24.3) | 20.2 (5.0, 17.8, 22.8) | 0.17 | 0.138 |
| Lumipulse Aβ40, median (IQR, Q1, Q3) | 278.7 (86.9, 239.4, 326.25) | 238.3 (46.5, 220.3, 266.8) | 0.012 | 0.28 |
| Lumipulse p-tau217 / Aβ42 | 0.005 (0.002, 0.004, 0.006) | 0.004 (0.001 (0.004, 0.005) | 0.219 | 0.14 |
| Lumipulse p-tau217 / (Aβ42/Aβ40) | 1.276 (1.060, 0.976, 2.031) | 0.917 (0.70, 0.814, 1.511) | 0.010 | 0.283 |

IQR: interquartile range, ES: estimated effect size, MOCA: Montreal Cognitive Assessment

Supplemental Table 3. Diagnostic performance using single and two cut-off approach in AD vs. ND

|  | Single cut-off approach |  |  |  | Two cut-off approach |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Parameter | AUC (95%CI) | Optimal cut-off value (pg/mL) | Sensitivity (95% CI) | Specificity (95% CI) | Lower cut-off (pg/mL) | Upper cut-off (pg/mL) | % inter-mediates |
| Elecsys p-tau181 | 0.881(0.820-0.942) | 1.155 | 85.4%(78.6-92.2%) | 85.2%(76.4-95.6%) | 0.936 | 1.915 | 59.2% |
| Elecsys p-tau217 | 0.932(0.890-0.974) | 0,405 | 89.3%(83.3-95.3%) | 88.9%(80.2-97.6%) | 0.275 | 0.555 | 31.8% |
| Lumipulse p-tau217 | 0.938(0.896-0.979) | 0.2465 | 85.4%(78.6-92.2%) | 94.3%(87.4-100%) | 0.1305 | 0.2665 | 26.1% |
| Lumipulse p-tau217/Aβ42 | 0.947(0.907-0.987) | 0.0105 | 90.3%(85.4-95.1%) | 96.2%(91.5-100%) | 0.0065 | 0.0135 | 3.8% |

Supplemental Table 4. Diagnostic performance using single and two cut-off approach in AD vs. PSY

|  | Single cut-off approach |  |  |  | Two cut-off approach |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Parameter | AUC (95%CI) | Optimal cut-off value (pg/mL) | Sensitivity (95% CI) | Specificity (95% CI) | Lower cut-off (pg/mL) | Upper cut-off (pg/mL) | % inter-mediates |
| Elecsys p-tau181 | 0.944(0.905-0.984) | 1.130 | 86.4%(77.7-92.4%) | 90.0%(73.1-92.4%) | 0.936 | 1.385 | 15.0% |
| Elecsys p-tau217 | 0.953(0.915-0.991) | 0.3715 | 92.2%(85.6-96%) | 96.6%(82.4-100%) | 0.2695 | 0.3715 | 6.0% |
| Lumipulse p-tau217 | 0.969(0.942-0.997) | 0.180 | 90.3%(83.6-95.3%) | 96.7%(83.4-99.9%) | 0.1305 | 0.1800 | 5.2% |
| Lumipulse p-tau217/Aβ42 | 0.970 (0.940-1.00) | 0.0085 | 93.2%(88.9-96.4%) | 96.3(86.0-100%) | 0.0065 | 0.0085 | <1% |

Supplemental Table 5. Diagnostic performance using single cut-off approach defining borderline cases as AD

| Parameter | AUC (95%CI) | Optimal cut-off value (pg/mL) | Sensitivity (95% CI) | Specificity (95% CI) |
| --- | --- | --- | --- | --- |
| Elecsys p-Tau181 | 0.867 (0.816-0.919) | 1.1550 | 78.2% (66.8-89.6%) | 86.9% (78.4-95.4%) |
| Elecsys p-Tau217 | 0.891 (0.843-0.938) | 0.405 | 82.4% (71.9-92.9%) | 91.6% (84.6-98.6%) |
| Lumipulse p-Tau217 | 0.911 (0.868-0.953) | 0.465 | 76.5% (64.7-88.3%) | 95.2% (89.8-100%) |
| Lumipulse p-tau217/Aβ42 | 0.931 (0.895-0.968) | 0.0109 | 83.2% (72.8-93.6%) | 92.4% (85.7-99.1%) |

Supplemental Figure 1. Linearity of plasma p-tau181 and plasma p-tau217 using a patient specimen diluted according to the manufacturer’s instructions. The initial p-tau181 concentration was 2.16 pg/mL (panel A) and the initial p-tau217 concentration was 1.09 pg/mL (panel B). The dotted line represents the linear fitting line.