Supplementary Materials

Table S1. Comparison of CNR and QH,j for different energy windows and sphere diameters at the acquisition time of 20 s/frame

| Energy window | Acquisition time | CNR1 | CNR2 | CNR3 | CNR4 | CNR5 | CNR6 | QH,1(%) | QH,2(%) | QH,3(%) | QH,4(%) | QH,5(%) | QH,6(%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 94keV±40% | 20 | 190.26 | 66.15 | 66.11 | 43.53 | 58.06 | 41.97 | 46.96 | 16.33 | 16.32 | 10.75 | 14.33 | 10.36 |
| 80keV±30% | 20 | 70.29 | 38.52 | 26.36 | 23.31 | 26.17 | 27.10 | 37.42 | 20.51 | 14.03 | 12.41 | 13.93 | 14.43 |
| 100keV±30% | 20 | 42.42 | 25.14 | 11.25 | 13.30 | 9.52 | 10.05 | 33.54 | 19.87 | 8.90 | 10.51 | 7.52 | 7.95 |
| 90keV±20% | 20 | 48.75 | 23.76 | 19.55 | 19.64 | 15.79 | 10.57 | 33.76 | 16.46 | 13.54 | 13.60 | 10.93 | 7.32 |
| 110keV±20% | 20 | 48.51 | 28.22 | 16.27 | 13.11 | 11.77 | 16.99 | 45.09 | 26.22 | 15.12 | 12.18 | 10.94 | 15.79 |
| 90keV±10% | 20 | 31.57 | 20.52 | 10.94 | 7.54 | 17.12 | 10.82 | 37.20 | 24.18 | 12.89 | 8.88 | 20.17 | 12.75 |
| 100keV±10% | 20 | 31.19 | 19.42 | 11.08 | 9.96 | 11.50 | 8.89 | 40.12 | 24.97 | 14.26 | 12.82 | 14.79 | 11.43 |

Table S2. Comparison of CNR and QH,j for different energy windows and sphere diameters at the acquisition time of 40 s/frame

| Energy window | Acquisition time | CNR1 | CNR2 | CNR3 | CNR4 | CNR5 | CNR6 | QH,1(%) | QH,2(%) | QH,3(%) | QH,4(%) | QH,5(%) | QH,6(%) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 94keV±40% | 40 | 106.01 | 52.39 | 40.96 | 33.26 | 25.83 | 19.37 | 38.64 | 19.10 | 14.93 | 12.12 | 9.41 | 7.06 |
| 80keV±30% | 40 | 61.45 | 30.24 | 20.48 | 14.96 | 17.17 | 13.33 | 37.19 | 18.30 | 12.40 | 9.05 | 10.39 | 8.07 |
| 100keV±30% | 40 | 120.37 | 48.49 | 50.78 | 28.36 | 33.07 | 23.59 | 40.63 | 16.37 | 17.14 | 9.57 | 11.16 | 7.96 |
| 90keV±20% | 40 | 41.49 | 20.96 | 14.61 | 14.97 | 13.76 | 10.24 | 40.67 | 20.55 | 14.32 | 14.68 | 13.49 | 10.04 |
| 110keV±20% | 40 | 54.48 | 25.35 | 21.55 | 12.72 | 13.74 | 9.67 | 44.94 | 20.91 | 17.78 | 10.49 | 11.33 | 7.98 |
| 90keV±10% | 40 | 43.37 | 25.80 | 19.98 | 3.94 | 3.44 | 6.76 | 38.15 | 22.70 | 17.58 | 3.47 | 3.02 | 5.95 |
| 100keV±10% | 40 | 44.69 | 21.75 | 21.54 | 11.51 | 14.77 | 15.36 | 43.49 | 21.16 | 20.96 | 11.20 | 14.37 | 14.94 |

Table S3. Comparison of CNR3, QH,3 and COV at different energy windows and acquisition times

| Energy window | Acquisition time | CNR3 | QH,3(%) | COV(%) | ∆Clung(%) |
| --- | --- | --- | --- | --- | --- |
| 94keV±40% | 20 | 66.11063 | 16.31887 | 2.379058 | 93.95904 |
| 80keV±30% | 20 | 26.36073 | 14.03329 | 5.130838 | 100.6209 |
| 100keV±30% | 20 | 11.25377 | 8.897227 | 7.619788 | 88.16705 |
| 90keV±20% | 20 | 19.55076 | 13.54126 | 6.675467 | 92.00633 |
| 110keV±20% | 20 | 16.26805 | 15.11875 | 8.957086 | 100.5195 |
| 90keV±10% | 20 | 10.93637 | 12.88871 | 11.35853 | 90.82931 |
| 100keV±10% | 20 | 11.08371 | 14.25714 | 12.39748 | 91.98356 |
| 94keV±40% | 40 | 40.95738 | 14.92954 | 3.513181 | 93.04315 |
| 80keV±30% | 40 | 20.4818 | 12.3966 | 5.833387 | 92.10435 |
| 100keV±30% | 40 | 50.77594 | 17.13838 | 3.253104 | 84.61483 |
| 90keV±20% | 40 | 14.6104 | 14.32387 | 9.44897 | 90.71846 |
| 110keV±20% | 40 | 21.55038 | 17.77597 | 7.949952 | 85.42155 |
| 90keV±10% | 40 | 19.98114 | 17.57802 | 8.478828 | 101.2701 |
| 100keV±10% | 40 | 21.53899 | 20.95839 | 9.378188 | 98.16972 |

| Table S4. CNR (40s-20s) Test statisticsc |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 94keV±40% | 80keV±30% | 100keV±30% | 90keV±20% | 110keV±20% | 100keV±10% | 90keV±10% |
| Z | -2.201a | -2.201a | -0.105b | -2.201a | -2.201b | -2.201b | -0.314b |
| Asymptotic significance (bilateral) | 0.028 | 0.028 | 0.917 | 0.028 | 0.028 | 0.028 | 0.753 |
| a. Based on positive rank |  |  |  |  |  |  |  |
| b. Based on negative rank |  |  |  |  |  |  |  |
| c. Wilcoxon signed rank test |  |  |  |  |  |  |  |

| Table S5. QH（40s-20s）Test statisticsc |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 94keV±40% | 80keV±30% | 100keV±30% | 90keV±20% | 110keV±20% | 100keV±10% | 90keV±10% |
| Z | -1.363a | -2.201b | -0.943a | -2.201a | -1.153b | -0.524b | -1.363a |
| Asymptotic significance (bilateral) | 0.173 | 0.028 | 0.345 | 0.028 | 0.249 | 0.600 | 0.173 |
| a. Based on positive rank |  |  |  |  |  |  |  |
| b. Based on negative rank |  |  |  |  |  |  |  |
| c. Wilcoxon signed rank test |  |  |  |  |  |  |  |

Table S6. Summary of optimal parameters of 90Y BECT imaging

| 90Y BECT imaging settings | Optimal parameters |
| --- | --- |
| Energy window center and width | 94 keV±40% (56.4–131.6 keV) |
| Acquisition time | 20 seconds / frame |
| Collimator | High-energy general-purpose collimator (HEGPC) |