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Table S1. Intralaboratory reproducibility of HPV Selfy using two different Real Time PCR machines. HPV Selfy assay was performed twice on 397 samples using two different Real Time PCR instruments: Quant Studio 5, indicated by the manufacturer’s protocol, and Light Cycler. Overall concordance observed was 99.5% (kappa value of 0.97, almost perfect agreement).
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HPV Selfy - Light Cycler HPV selfy - Quant Studio 5
HR-HPV negative HR-HPV positive Total



HR-HPV negative Negative 363 0 363
HR-HPV positive Positive 2 32 34



Total 365 32 397
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