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Numerical Analysis:  Photons depth analysis (current protocol: 'Custom_0')
D100 67.3%
D200 38.9%
D50 86.5%
Dmax 100.0 %
Qi 0.58
R100 14.9 mm
R50 154.3 mm
Data Information:
ScanType RadiationType Energy Inline Crossline SSD
Depth Dose Photons 6MV 100.0 mm 100.0 mm 1000.0 mm
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Numerical Analysis: Photons profile analysis (current protocol: ‘Custom_0")
Flatness | Symmetry | Penumbra | FieldWidth | Deviation | Center |
105.2 % 100.4 % 7.7 mm - 7.6 mm 1142 mm 0.0 mm

T 9% 100.8 % 1L5mm-116mm 4034 mm 0.0 mm

Data Information:

ScanType | RadiationType | Energy | Inline | Crossline | SSD |
Inline Photons 100.0 mm 100.0 mm 1000.0 mm

Crossline Photons 350.0 mm 350.0 mm 1000.0 mm
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Numerical Analysis:

100 4

il

D100 634%
D200 345%
DS0 852%
Dmax 1000 %
Ds 542 %
Qi 054
R100 15.3 mm
RSO 139.0 mm
Ruser! 61.1 mm
Ruser3 | 2929 mm

Data Information:

Depth [an]

Photons depth analysis (current protocol: 'VenusX')

ScanType RadiationType

Energy

Inline

Crossline

SSD

Depth Dose  Photons

6MV(FFF)

100.0 mm

100.0 mm

800.0 mm
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Numerical Analysis:

0ff axis [mm]

Photons profile analysis (current protocol: 'VenusX')

Flatness Symmetry Penumbra FieldWidth Deviation Center

1142 % 101.5 % 8.1 mm - 8.6 mm 99.0 mm 124% -0.5 mm

1715 % 1013 % 14.9 mm - 13.9 mm 392.3 mm 416 % 0.1 mm
Data Information:

ScanType RadiationType | Energy Inline Crossline SSD

Crossline Photons 6MV(FFF) 100.0 mm 100.0 mm 800.0 mm

Crossline Photons 6MV(FFF) 400.0 mm 400.0 mm 800.0 mm





