Supplementary materials

Optimization and validation of bifenazate and bifenazate-diazene quantification in agricultural products using a reduction method

Table S1. The instrumental condition of LC-MS/MS for quantitative anlaysis of bifenazate and bifenazate diazene

| Instrument | Agilent Technologies 6420 Triple Quad LC/MS (Agilent, Santa Clara, CA, USA) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Column | Poroshell 120 SB–C18 (2.1 × 100 mm, 2.7 µm) |  |  |  |  |  |  |
| Flow | 0.2 mL min–1 |  |  |  |  |  |  |
| Mobile phase | A : 0.1% Formic acid with 5mM ammonium formate in water B : 0.1% Formic acid with 5mM ammonium formate in methanol |  |  |  |  |  |  |
|  | Time (min) |  | %A |  |  | %B |  |
|  | 5 |  | 20 |  |  | 80 |  |
| Injection volume | 2 μL |  |  |  |  |  |  |
| Ionization source | Electrospray ionization (ESI) |  |  |  |  |  |  |
| Polarity | Positive |  |  |  |  |  |  |
| Source temperature | 150℃ |  |  |  |  |  |  |
| Capillary voltage | 4.0 kV |  |  |  |  |  |  |
| Gas temperature | 325℃ (N2, 360 L hr–1) |  |  |  |  |  |  |
| MRM | Analyte | Retention time (min) |  | Precursor ion (m/z) | Product ion (m/z) |  | Collision energy (V) |
|  | Bifenazate | 2.1 |  | 301.2 | 198.0 |  | 10 |
|  |  |  |  |  | 170.1 |  | 20 |
|  | Bifenazate diazene | 3.1 |  | 299.2 | 213.1 |  | 7 |
|  |  |  |  |  | 197.2 |  | 20 |

Table S2. Statistical parameters obtained from the analysis of variance (ANOVA) for the models for bifenazate conversion rate and recovery rate from reaction time and temperature

| Parameter |  | Conversion rate (%) | Recvoery rate (%) |
| --- | --- | --- | --- |
| Intercept, X0 |  | −46.9*** | −7.54*** |
| Linear | X1a) | 3.06*** | 3.77 |
|  | X2b) | 22.9*** | 11.4*** |
| Quadratic | X12 | −0.0166*** | −0.0343*** |
|  | X12 | −1.15*** | −0.451*** |
| Cross Product | X1 X2 | −0.199*** | −0.124*** |
| R2 |  | 0.9883 | 0.9816 |
| F-value (model) |  | 190*** | 85.9*** |
| F-value (lack of fit) |  | 0.077 | 0.070 |

a) X1 = Temperature (℃), b) X2 = reaction time (h); Level of significance * p < 0.05, ** p < 0.01, *** p < 0.001.

Table S3. Precision of bifenazate in agricultural sampels

| Agricultural product | Concentration (mg kg-1) | Recovery (CV) |  |
| --- | --- | --- | --- |
|  |  | Inter-day (n=5) | Intra-day (3day, n=5/day) |
| Pepper | 0.01 | 80.6 (5.1) | 100.5 (2.0) |
|  | 0.02 | 82.7 (3.2) | 97.7 (1.7) |
|  | 0.1 | 81.7 (2.4) | 85.8 (3.2) |
| Mandarin | 0.01 | 95.9 (1.1) | 90.3 (12) |
|  | 0.02 | 92.9 (8.1) | 101.1 (4.5) |
|  | 0.1 | 91.5 (2.0) | 92.9 (8.1) |
| Brown rice | 0.01 | 90.7 (5.0) | 103.3 (5.0) |
|  | 0.02 | 85.8 (6.1) | 105.3 (4.3) |
|  | 0.1 | 85.5 (3.6) | 82.7 (6.0) |

Table S4. System suitability of LC-MS/MS for bifenazate in agricultural sampels

| Agricultural product | Concentration (mg kg-1) | Instrumental repeatability (n=5) |  |
| --- | --- | --- | --- |
|  |  | Peak area (mean ± standard deviation) | CV (%) |
| Pepper | 0.01 | 354.7 ± 30.8 | 8.7 |
|  | 0.1 | 3445.2 ± 92.9 | 2.7 |
| Mandarin | 0.01 | 353.4 ± 23.8 | 6.8 |
|  | 0.1 | 3401.4 ± 119.2 | 3.5 |
| Brown rice | 0.01 | 345.0 ± 19.6 | 5.7 |
|  | 0.1 | 3229.8 ± 85.5 | 2.6 |