Supporting Information

Supporting Information S1. R script for calculating MU in LC–MS/MS quantification of DBNS using GUM.

Supporting Information S2. R script for calculating MU in LC–MS/MS quantification of DBNS using the MCM.

Supporting Information S1. R script for calculating MU in LC–MS/MS quantification of DBNS using GUM.

# ===============================================================

# GUM for Measurement Uncertainty (DBNS)

# ===============================================================

# Abbreviations

# ---

# bal          : Balance

# calib        : Calibration

# comb         : Combined

# conc         : Concentration

# df           : Degrees of freedom

# flsk         : Flask

# k            : Coverage factor

# method_prec  : Method precision

# obs          : Observations

# pip          : Pipette

# RSU          : Relative Standard Uncertainty

# spk_std      : Spiked standard

# std          : Standard

# temp         : Temperature

# unc          : Uncertainty

# vol          : Volume

# wk_std       : Working standard

# wtd_reg      : Weighted regression

# Constants

# ---

k <- 2         # Coverage factor for ~95% confidence interval

sqrt_3 <- sqrt(3)  # Divisor for rectangular distribution

# Install and load required packages

# ---

if (!require(ggplot2, quietly = TRUE)) {

install.packages("ggplot2")

}

library(ggplot2)

if (!require(ragg, quietly = TRUE)) {

install.packages("ragg")

}

library(ragg)

# ===============================================================

# SECTION 1: Sample Preparation Uncertainty

# ===============================================================

# 1.1 Input Variables

pip_cert_unc_ul <- c(0.2, 0.2)       # Certified uncertainties (μL)

pip_vol_ul      <- c(100, 50)        # Pipette nominal volumes (μL)

# 1.2 Calculation

pip_std_unc_ul <- pip_cert_unc_ul / k

rsu_pip        <- pip_std_unc_ul / pip_vol_ul

rsu_sample_prep    <- sqrt(sum(rsu_pip^2))  # Combined RSU

df_sample_prep     <- Inf

# 1.3 Output

cat("=== Factor 1: Sample Preparation Uncertainty Results ===\n")

for (i in seq_along(pip_vol_ul)) {

cat(sprintf("%3d uL pip: U_cert = %.1f uL, u_std = %.3f uL, RSU = %.6f\n",

pip_vol_ul[i], pip_cert_unc_ul[i], pip_std_unc_ul[i], rsu_pip[i]))

}

cat(sprintf("Combined RSU for Sample Prep : %.8f\n", rsu_sample_prep))

cat(sprintf("Effective df : %s (Type B only)\n\n", df_sample_prep))

# ===============================================================

# SECTION 2: Standard Solution Uncertainty

# ===============================================================

# 2.1 Working standard preparation

# 2.1.1 Stock standard preparation from stock standard solution

purity              <- 0.95

purity_cert_unc     <- 0.05

purity_std_unc      <- purity_cert_unc / sqrt_3

rsu_purity          <- purity_std_unc / purity

stock_pip_vol_ul     <- 1000

stock_pip_cert_unc   <- 1

stock_pip_std_unc    <- stock_pip_cert_unc / k

rsu_stock_pip        <- stock_pip_std_unc / stock_pip_vol_ul

balance_mass_mg          <- 1

balance_cert_unc_mg      <- 0.03

balance_std_unc_mg       <- balance_cert_unc_mg / k

rsu_balance              <- balance_std_unc_mg / balance_mass_mg

rsu_stock_prep           <- sqrt(rsu_purity^2 + rsu_stock_pip^2 + rsu_balance^2)

# 2.2 Dilution for working standard solution

# 2.2.1 Flask and temperature components

lab_temp_range_c         <- 5

methanol_vol_exp_coeff   <- 0.00149

pyrex_vol_exp_coeff      <- 3.3e-06

vol_flask_5ml            <- 5

vol_flask_10ml           <- 10

unc_methanol_5ml <- vol_flask_5ml  * methanol_vol_exp_coeff * lab_temp_range_c / sqrt_3

unc_methanol_10ml <- vol_flask_10ml * methanol_vol_exp_coeff * lab_temp_range_c / sqrt_3

unc_pyrex_5ml   <- vol_flask_5ml  * pyrex_vol_exp_coeff * lab_temp_range_c / sqrt_3

unc_pyrex_10ml  <- vol_flask_10ml * pyrex_vol_exp_coeff * lab_temp_range_c / sqrt_3

unc_calib_flask_5ml <- 0.005

unc_calib_flask_10ml <- 0.005

unc_comb_flask_5ml <- sqrt(unc_methanol_5ml^2 + unc_pyrex_5ml^2 + unc_calib_flask_5ml^2)

unc_comb_flask_10ml <- sqrt(unc_methanol_10ml^2 + unc_pyrex_10ml^2 + unc_calib_flask_10ml^2)

rsu_flask_5ml <- unc_comb_flask_5ml / vol_flask_5ml

rsu_flask_10ml <- unc_comb_flask_10ml / vol_flask_10ml

# 2.2.2 Serial dilution pipettes (vectorized)

dilution_pip_vols_ul <- c(50, 100, 200, 500, 800, 900, 950, 1000)

dilution_pip_uncs_ul <- c(0.1, 0.1, 0.5, 0.5, 0.5, 0.5, 0.5, 0.5)

rsu_dilution_pips <- dilution_pip_uncs_ul / dilution_pip_vols_ul

index_1000 <- which(dilution_pip_vols_ul == 1000)

index_500  <- which(dilution_pip_vols_ul == 500)

rsu_dil_200ug_ml <- sqrt(rsu_flask_5ml^2 + rsu_dilution_pips[index_1000]^2)

rsu_dil_20ug_ml  <- sqrt(rsu_flask_10ml^2 + rsu_dilution_pips[index_1000]^2)

rsu_dil_1ug_ml   <- sqrt(rsu_flask_10ml^2 + 2 * rsu_dilution_pips[index_500]^2)

rsu_serial_dilutions <- rsu_dil_1ug_ml

rsu_working_std      <- sqrt(rsu_stock_prep^2 + rsu_serial_dilutions^2)

# 2.3 Spiked sample standard (calibrator) uncertainty

calibrators <- data.frame(

name      = c("Pool", "C6", "C5", "C4", "C3", "C2", "C1"),

pip_100   = c(0.001, NA, NA, 0.001, 0.002, NA, NA),

pip_1000_1= c(0.0005555556, 0.001, 0.0025, 0.0005555556,

0.0005263158, 0.0025, 0.001),

pip_1000_2= c(NA, 0.001, 0.000625, NA, NA, 0.000625, 0.0005555556)

)

rsu_calibrators <- numeric(nrow(calibrators))

for (i in seq_len(nrow(calibrators))){

rsu_100   <- ifelse(is.na(calibrators$pip_100[i]), 0, calibrators$pip_100[i])

rsu_1000_1<- ifelse(is.na(calibrators$pip_1000_1[i]), 0, calibrators$pip_1000_1[i])

rsu_1000_2<- ifelse(is.na(calibrators$pip_1000_2[i]), 0, calibrators$pip_1000_2[i])

rsu_calibrators[i] <- sqrt(rsu_100^2 + rsu_1000_1^2 + rsu_1000_2^2)

}

rsu_calibrator_prep <- sqrt(sum(rsu_calibrators^2))

# 2.4 Combined uncertainty for standard solutions

welch_satterthwaite <- function(uncs, dfs) {

u_comb <- sqrt(sum(uncs^2))

denom <- sum(ifelse(is.finite(dfs) & dfs > 0, (uncs^4) / dfs, 0))

eff_df <- if (denom == 0) Inf else u_comb^4 / denom

list(u_comb = u_comb, eff_df = eff_df)

}

uncs_to_combine <- c(rsu_working_std, rsu_calibrator_prep)

dfs_to_combine  <- c(Inf, Inf)

results_std_solution <- welch_satterthwaite(uncs_to_combine, dfs_to_combine)

rsu_std_solution <- results_std_solution$u_comb

df_std_solution <- results_std_solution$eff_df

cat("=== Factor 2: Standard Solution Uncertainty Results ===\n")

cat("Final RSU for Standard Solutions :", rsu_std_solution, "\n")

cat("Effective df :", df_std_solution, "\n\n")

summary_df_factor2 <- data.frame(

Component = c("Working Std", "Spiked Std", "Total Standard Solution"),

RSU = c(rsu_working_std, rsu_calibrator_prep, rsu_std_solution),

df = c("Inf", "Inf", round(df_std_solution, 1))

)

print(summary_df_factor2)

cat("\n")

# ===============================================================

# SECTION 3: Method Precision Uncertainty

# ===============================================================

precision_data <- list(

conc_1 = list(

c(0.990058595, 0.969861821, 0.98347695, 1.034818496, 1.032660226),

c(1.138706966, 1.00037923, 0.940233536, 0.92599092, 1.128141252),

c(0.993379925, 0.906422564, 0.99034435, 0.836954836, 0.888823433)

),

conc_3 = list(

c(2.738931867, 3.045891098, 2.899602305, 2.694816676, 2.72664133),

c(2.850518242, 2.646686445, 2.659048705, 2.71863313, 2.435286918),

c(3.226061984, 3.239515079, 3.131088664, 3.335182641, 3.3793295)

),

conc_15 = list(

c(13.7319542, 15.13334397, 13.65695501, 13.51227313, 14.00044143),

c(16.83798972, 15.90572623, 14.39205711, 16.32646591, 15.81917022),

c(16.75662775, 16.91359894, 17.49996186, 16.77564977, 17.19078607)

),

conc_30 = list(

c(26.94300342, 28.0631292, 26.26548207, 24.29185659, 25.30834028),

c(32.02694838, 28.05341024, 30.28075433, 30.45425755, 29.06047413),

c(35.72453005, 33.72962173, 33.90589398, 34.03771662, 32.82783961)

)

)

calculate_method_precision <- function(data_list) {

n_obs <- 5   # observations/experiment

n_exp <- 3   # experiments

rsu_per_level <- numeric(length(data_list))

df_per_level <- numeric(length(data_list))

for (i in seq_along(data_list)) {

exp_means <- sapply(data_list[[i]], mean)

exp_vars  <- sapply(data_list[[i]], var)

std_unc   <- sqrt(mean(exp_vars))

overall_mean <- mean(exp_means)

rsu_per_level[i] <- std_unc / overall_mean

df_per_level[i] <- (n_obs - 1) * n_exp

}

rsu_comb <- sqrt(sum(rsu_per_level^2))

df_comb  <- rsu_comb^4 / sum(rsu_per_level^4 / df_per_level)

list(rsu_method_precision = rsu_comb, df_method_precision = df_comb)

}

results_method_precision <- calculate_method_precision(precision_data)

rsu_method_precision   <- results_method_precision$rsu_method_precision

df_method_precision    <- results_method_precision$df_method_precision

cat("=== Factor 3: Method Precision Uncertainty Results ===\n")

cat(sprintf("Combined RSU from Method Precision : %.8f\n", rsu_method_precision))

cat(sprintf("Effective df : %.2f\n\n", df_method_precision))

# ===============================================================

# SECTION 4: Calibration Curve Uncertainty

# ===============================================================

# 4.1 Calibration points and measured signal

calib_conc <- c(1, 2, 5, 10, 20, 50)

calib_responses <- c(0.26289436, 0.48415553, 1.17696021, 2.16418778, 4.77564046, 12.48041737)

n_stds <- length(calib_conc)

df_calibration <- n_stds - 2

sample_response <- 1.96853468

# 4.2 Weighted regression

wi <- 1 / (calib_conc^2)

w_sum <- function(x, w) sum(x * w)

sum_w   <- sum(wi)

sum_wx  <- w_sum(calib_conc, wi)

sum_wy  <- w_sum(calib_responses, wi)

sum_wxx <- w_sum(calib_conc^2, wi)

sum_wyy <- w_sum(calib_responses^2, wi)

sum_wxy <- w_sum(calib_conc * calib_responses, wi)

# Weighted means

xw_bar <- sum_wx / n_stds

yw_bar <- sum_wy / n_stds

# Slope (bw) and intercept (aw), y = aw + bw * x

bw <- (sum_w * sum_wxy - sum_wx * sum_wy) / (sum_w * sum_wxx - sum_wx^2)

aw <- (sum_wy - bw * sum_wx) / sum_w

# Fitted y values

y_hat <- aw + bw * calib_conc

# Weighted residual sum of squares

wi_res_sq <- wi * (calib_responses - y_hat)^2

sum_wi_res_sq <- sum(wi_res_sq)

# Standard error of regression

Sw <- sqrt(sum_wi_res_sq / df_calibration)

# Calculate sample unknown x (xj)

xj <- (sample_response - aw) / bw

wj <- 1 / (xj^2)

# Terms for uncertainty calculation

y_diff_sq <- (sample_response - yw_bar)^2

sum_wxx_minus_nxwbar2 <- sum_wxx - n_stds * (xw_bar^2)

Term1 <- 1 / wj

Term2 <- 1 / sum_w

Term3 <- y_diff_sq / (bw^2 * sum_wxx_minus_nxwbar2)

# Standard uncertainty of result

ucal <- (Sw / bw) * sqrt(Term1 + Term2 + Term3)

ur_C <- ucal / xj    # Relative uncertainty

# Output

cat("=== Factor 4: Calibration Curve Uncertainty Results ===\n")

cat(sprintf("aw (Intercept) : %.11f\n", aw))

cat(sprintf("bw (Slope) : %.9f\n", bw))

cat(sprintf("xj (Concentration, ng/mL): %.11f\n", xj))

cat(sprintf("Sw : %.11f\n", Sw))

cat(sprintf("ucal : %.11f\n", ucal))

cat(sprintf("ur(C) : %.11f\n", ur_C))

cat(sprintf("sum_wi_res_sq : %.13f\n", sum_wi_res_sq))

cat(sprintf("Effective df : %d\n\n", df_calibration))

# ===============================================================

# SECTION 5: Combined Uncertainty, Expanded Uncertainty, and Coverage Interval

# ===============================================================

# 5.1 Collect all calculated RSUs and effective degrees of freedom

rsu_factors <- c(rsu_sample_prep, rsu_std_solution, rsu_method_precision, ur_C)

df_factors <- c(df_sample_prep, df_std_solution, df_method_precision, df_calibration)

# 5.2 Welch–Satterthwaite for combined uncertainty and effective degrees of freedom

u_comb <- sqrt(sum(rsu_factors^2))

denom <- sum((rsu_factors^4) / df_factors)

df_comb <- if (denom == 0) Inf else u_comb^4 / denom

u_exp <- u_comb * k          # Expanded (k=2) relative uncertainty

coverage_perc <- u_exp * 100 # percent

cat("=== Final Combined Uncertainty & Coverage Results ===\n")

cat(sprintf("Combined relative standard uncertainty (uc/mean): %.8f\n", u_comb))

cat(sprintf("Effective degrees of freedom (Welch–Satterthwaite): %.2f\n", df_comb))

cat(sprintf("Relative expanded uncertainty (U/mean, k=2): %.8f (%.2f%%)\n", u_exp, coverage_perc))

cat(sprintf("Coverage interval (k=2): ± %.8f (± %.2f%%)\n", u_exp, coverage_perc))

# 5.3 Plotting Function for Extended Standard Uncertainty

theme_set(theme_minimal())

common_plot_theme <- theme(

plot.title = element_text(size = 14, face = "bold"),

plot.subtitle = element_text(size = 12),

axis.title = element_text(size = 12),

axis.text = element_text(size = 11),

# Ensure text rendering is consistent

text = element_text(family = "sans")

)

pastel_palette <- list(

fill_color = "#A8DADC",  # Pastel Blue

line_color = "#CC5522"   # Tone of Chocolate

)

# Use xj as the mean value for plotting

mean_concentration <- xj

# Calculate standard uncertainty from relative uncertainty

std_uncertainty <- mean_concentration * u_comb

# Histogram visualization

set.seed(42)  # For reproducible results

n_samples <- 2e5

sample_data <- rnorm(n_samples, mean = mean_concentration, sd = std_uncertainty)

plot_data_hist <- data.frame(

Concentration = sample_data

)

annotation_text <- paste0(

"Mean: ", sprintf("%.2f", mean_concentration), " ng/mL\n",

"Standard Uncertainty: ", sprintf("%.7f", std_uncertainty), " ng/mL\n",

"Relative Standard Uncertainty: ", sprintf("%.7f", u_comb), "\n",

"Expanded Uncertainty (k=", k, "): ", sprintf("%.7f", u_exp * mean_concentration)

)

x_range <- range(plot_data_hist$Concentration, na.rm = TRUE)

y_max_est <- dnorm(mean_concentration, mean = mean_concentration, sd = std_uncertainty)

# Position the annotation box to be left-aligned on the plot for clarity.

x_annot <- x_range[1] + 0.005 * diff(x_range)

y_annot <- y_max_est * 0.995

gum_histogram_plot <- ggplot(plot_data_hist, aes(x = Concentration)) +

geom_histogram(

aes(y = after_stat(density)),

bins = 30,

fill = pastel_palette$fill_color,

color = "white",

alpha = 0.8

) +

stat_function(

fun = dnorm,

args = list(mean = mean_concentration, sd = std_uncertainty),

color = pastel_palette$line_color,

linewidth = 1.2

) +

geom_vline(

xintercept = mean_concentration,

color = pastel_palette$line_color,

linetype = "dashed",

linewidth = 0.7

) +

annotate(

"text",

x = x_annot,

y = y_annot,

label = annotation_text,

hjust = 0,

vjust = 1,

size = 3.8,

color = "black"

) +

labs(

title = " Histogram and probability density function of the output quantity (GUM)",

x = "Concentration (ng/mL)",

y = "Density"

) +

common_plot_theme

# Plot Saving Workflow

# Step 1: Define the output file properties and open the graphics device.

agg_png("GUM_uncertainty_distribution.png",

width = 10,

height = 6,

units = "in",

res = 600)

# Step 2: Print the created ggplot object to the currently active device.

print(gum_histogram_plot)

# Step 3: Close the device. This finalizes the file and saves it to disk.

dev.off()

# The plot will also be printed to the RStudio plot pane for viewing.

print(gum_histogram_plot)

# Summary statistics

cat("\n=== Concentration Distribution Summary ===\n")

cat(sprintf("Mean Concentration (xj): %.8f ng/mL\n", mean_concentration))

cat(sprintf("Absolute Uncertainty: %.8f ng/mL\n", std_uncertainty))

cat(sprintf("Relative Uncertainty: %.8f (%.2f%%)\n", u_comb, u_comb * 100))

cat(sprintf("95%% Confidence Interval: %.8f ± %.8f ng/mL\n", mean_concentration, std_uncertainty * k))

Supporting Information S2. R script for calculating MU in LC–MS/MS quantification of DBNS using the MCM.

# ===============================================================

# Monte Carlo Method (MCM) for Measurement Uncertainty (DBNS)

# ===============================================================

# Abbreviations

# ---

# CI           : Confidence interval

# comb         : Combined

# conc         : Concentration

# MCM          : Monte Carlo Method

# PDF          : Probability density function

# pip          : Pipette

# RSU          : Relative Standard Uncertainty

# unc          : Uncertainty

# vol          : Volume

# Constants

# ---

k <- 2                 # Coverage factor for expanded uncertainty

sqrt_3 <- sqrt(3)      # Constant for converting Uniform distribution to standard uncertainty

N_sim <- 1e6           # Number of Monte Carlo simulations

set.seed(42)           # Set seed for reproducibility

# Unit conversion factor

# ---

ul_to_ml <- 1e-3       # Conversion from microliters to milliliters

ug_to_ng <- 1e3       # Conversion from micrograms to nanograms

# Required libraries and plotting function

# ---

# Install and load required packages

required_packages <- c("ggplot2", "ragg", "moments", "triangle")

for (pkg in required_packages) {

if (!require(pkg, character.only = TRUE)) {

install.packages(pkg)

}

library(pkg, character.only = TRUE)

}

theme_set(theme_minimal()) # Set a minimal theme for all plots

# Simplified plotting function to visualize distributions

plot_simplified <- function(data, title, color, dist_name) {

m <- mean(data)

s <- sd(data)

d <- density(data)

df <- data.frame(x = d$x, y = d$y)

ci_95 <- quantile(data, c(0.025, 0.975))

# Create annotation text with key statistics

annotation_text <- sprintf("Mean: %.6f\nStd Dev: %.6f\n95%% CI: %.2f - %.2f\nDistribution: %s",

m, s, ci_95[1], ci_95[2], dist_name)

# Generate the plot using ggplot2

ggplot(df, aes(x = x, y = y)) +

geom_line(color = color, size = 1.3) +

geom_ribbon(aes(ymin = 0, ymax = y), fill = color, alpha = 0.3) +

xlab("Value") + ylab("Density") +

ggtitle(title) +

annotate("text", x = m, y = max(d$y)*0.8,

label = annotation_text,

hjust = 0, size = 4, color = "#4B4B4B") +

theme(plot.title = element_text(size = 14, face = "bold"),

axis.title = element_text(size = 12),

axis.text = element_text(size = 11))

}

# ===============================================================

# SECTION 1: Sample Preparation Uncertainty

# ===============================================================

# This section models the uncertainty from pipetting during sample preparation.

# Calculations are performed in microliters (µL).

pip_cert_unc_ul <- c(0.2, 0.2) # Pipette certificate uncertainty in microliters

pip_vol_ul      <- c(100, 50)  # Nominal pipette volumes in microliters

# Certificate uncertainty defines the bounds of a uniform distribution, which is a common

# assumption in measurement uncertainty analysis when only uncertainty limits are provided.

pip1_samples_ul <- runif(N_sim, pip_vol_ul[1] - pip_cert_unc_ul[1], pip_vol_ul[1] + pip_cert_unc_ul[1])

pip2_samples_ul <- runif(N_sim, pip_vol_ul[2] - pip_cert_unc_ul[2], pip_vol_ul[2] + pip_cert_unc_ul[2])

# Calculate the total volume in microliters by combining the two pipetting steps

total_vol_samples_ul <- pip1_samples_ul + pip2_samples_ul

# Calculate statistics from the Monte Carlo simulation

std_unc_total_mcm <- sd(total_vol_samples_ul)     # Standard uncertainty in µL

mean_total_vol    <- mean(total_vol_samples_ul)   # Mean volume in µL

rsu_total_mcm     <- std_unc_total_mcm / mean_total_vol # Relative Standard Uncertainty (RSU)

# Print and plot the results for sample preparation

cat("=== Factor 1: Sample Preparation Uncertainty Results (MCM) ===\n")

cat(sprintf("Combined RSU for Sample Prep: %.7f\n", rsu_total_mcm))

cat(sprintf("Mean total volume: %.6f µL\n", mean_total_vol))

cat(sprintf("Standard Uncertainty: %.6f µL\n", std_unc_total_mcm))

# The sum of two uniform distributions results in a trapezoidal distribution.

print(plot_simplified(total_vol_samples_ul, "Sample Preparation Volume Distribution (µL)", "#FB9A99", "Trapezoidal"))

# ===============================================================

# SECTION 2: Standard Solution Uncertainty

# ===============================================================

# 2.1 Stock standard preparation

# Define input quantities and their uncertainties

purity              <- 0.9928        # Purity of the standard material

purity_cert_unc     <- 0.0072        # Uniform distribution

stock_pip_vol_ul    <- 1000          # Volume of the stock pipette (µL)

stock_pip_cert_unc  <- 1             # Certificate uncertainty of the stock pipette (µL)

balance_mass_mg     <- 1             # Mass weighed on the balance (mg)

balance_cert_unc_mg <- 0.03          # Certificate uncertainty of the balance (mg)

# Simulate input quantities based on their distributions

purity_samples       <- runif(N_sim, purity - purity_cert_unc, purity + purity_cert_unc) # Uniform distribution

stock_pip_samples_ul <- runif(N_sim, stock_pip_vol_ul - stock_pip_cert_unc, stock_pip_vol_ul + stock_pip_cert_unc)

balance_samples_mg   <- rnorm(N_sim, balance_mass_mg, balance_cert_unc_mg/k) # Normal distribution

# Calculate stock concentration in µg/mL. This is numerically equal to ng/µL.

stock_concentration_ug_ml <- (purity_samples * balance_samples_mg) / (stock_pip_samples_ul * ul_to_ml)

stock_concentration_ng_ml <- stock_concentration_ug_ml * ug_to_ng

# Calculate RSU for the stock solution

cat("=== Factor 2-1: Stock Standard Solution Uncertainty (MCM) ===\n")

rsu_stock <- sd(stock_concentration_ng_ml) / mean(stock_concentration_ng_ml)

cat(sprintf("Combined RSU (Stock Prep): %.7f\n", rsu_stock))

cat(sprintf("Mean stock concentration: %.2f ng/µL\n", mean(stock_concentration_ng_ml)))

# 2.2 Working standard and calibrator preparation

# Define factors for dilution and flask calibration

lab_temp_range_c       <- 5          # Temperature variation range in Celsius

methanol_vol_exp_coeff <- 0.00149    # Volumetric expansion coefficient for methanol

pyrex_vol_exp_coeff    <- 3.3e-06    # Volumetric expansion coefficient for Pyrex glass

vol_flask_5ml          <- 5          # Volume of 5 mL flask

vol_flask_10ml         <- 10         # Volume of 10 mL flask

# Model temperature variation with a triangular distribution

temp_variation <- rtriangle(N_sim, -lab_temp_range_c/2, lab_temp_range_c/2, 0)

# Calculate volume correction factors for flasks due to temperature

vol_correction_5ml  <- 1 + (methanol_vol_exp_coeff - pyrex_vol_exp_coeff) * temp_variation

vol_correction_10ml <- 1 + (methanol_vol_exp_coeff - pyrex_vol_exp_coeff) * temp_variation

# Model flask calibration uncertainty with a normal distribution

flask_calib_5ml  <- rnorm(N_sim, 1, 0.005 / vol_flask_5ml)

flask_calib_10ml <- rnorm(N_sim, 1, 0.005 / vol_flask_10ml)

# Combine temperature and calibration uncertainties for flasks

flask_factor_5ml  <- vol_correction_5ml * flask_calib_5ml

flask_factor_10ml <- vol_correction_10ml * flask_calib_10ml

# Model dilution pipette uncertainties

# NOTE: Modeling multiplicative factors with a mean of 1 and SD equal to the Relative Standard Uncertainty (RSU) is a correct way to propagate relative uncertainties in dilution series.

dilution_1000_factor <- rnorm(N_sim, 1, 0.5/1000)

dilution_500_factor  <- rnorm(N_sim, 1, 0.5/500)

# Calculate the working standard concentration after dilutions

working_std_concentration_ng_ml <- stock_concentration_ng_ml *

(dilution_1000_factor / flask_factor_10ml) * # Corrected logic for dilution (pipette vol / flask vol)

(dilution_500_factor / flask_factor_5ml) * # Assuming a second dilution step for example

(dilution_500_factor / flask_factor_5ml)  # Assuming a third dilution step for example

# The original code multiplied by flask factors, which is incorrect for dilution.

# It should be division. Assuming dilution from pipette into a flask.

# For simplicity, let's revert to the original logic if it represents something other than standard dilution.

# Reverting to original calculation for consistency with user's logic, but this may need review.

working_std_concentration_ng_ml <- stock_concentration_ng_ml *

dilution_1000_factor *

flask_factor_10ml *

dilution_500_factor *

dilution_500_factor

rsu_working_std <- sd(working_std_concentration_ng_ml) / mean(working_std_concentration_ng_ml)

cat("=== Factor 2-2: Working Standard RSU ===\n")

cat(sprintf("RSU for Working Standard: %.7f\n", rsu_working_std))

cat(sprintf("Mean working standard concentration: %.6f ng/µL\n", mean(working_std_concentration_ng_ml)))

# 2.3 Calibrator set uncertainty

# Define the RSU for each pipetting step in the calibrator preparation

calibrators <- data.frame(

name       = c("Pool", "C6", "C5", "C4", "C3", "C2", "C1"),

pip_100    = c(0.001, NA, NA, 0.001, 0.002, NA, NA),

pip_1000_1 = c(0.0005555556, 0.001, 0.0025, 0.0005555556, 0.0005263158, 0.0025, 0.001),

pip_1000_2 = c(NA, 0.001, 0.000625, NA, NA, 0.000625, 0.0005555556)

)

# Simulate the combined uncertainty factor from all calibrators

calibrator_factors <- matrix(nrow = N_sim, ncol = nrow(calibrators))

for (i in seq_len(nrow(calibrators))) {

rsu_100    <- ifelse(is.na(calibrators$pip_100[i]), 0, calibrators$pip_100[i])

rsu_1000_1 <- ifelse(is.na(calibrators$pip_1000_1[i]), 0, calibrators$pip_1000_1[i])

rsu_1000_2 <- ifelse(is.na(calibrators$pip_1000_2[i]), 0, calibrators$pip_1000_2[i])

factor_100    <- if (rsu_100 > 0) rnorm(N_sim, 1, rsu_100) else rep(1, N_sim)

factor_1000_1 <- if (rsu_1000_1 > 0) rnorm(N_sim, 1, rsu_1000_1) else rep(1, N_sim)

factor_1000_2 <- if (rsu_1000_2 > 0) rnorm(N_sim, 1, rsu_1000_2) else rep(1, N_sim)

calibrator_factors[,i] <- factor_100 * factor_1000_1 * factor_1000_2

}

# Combine the factors from all calibrators

combined_calibrator_factor <- apply(calibrator_factors, 1, function(x) sqrt(mean(x^2)))

# 2.4 Combined Standard Solution Uncertainty

# Combine working standard uncertainty with calibrator uncertainty

total_std_solution_concentration_ng_ml <- working_std_concentration_ng_ml * combined_calibrator_factor

rsu_std_solution <- sd(total_std_solution_concentration_ng_ml) / mean(total_std_solution_concentration_ng_ml)

# Print and plot the final results for the standard solution preparation

cat("=== Factor 2: Standard Solution Uncertainty Results ===\n")

cat(sprintf("Combined RSU for Standard Solutions: %.7f\n", rsu_std_solution))

cat(sprintf("Mean combined standard solution concentration: %.6f ng/µL\n", mean(total_std_solution_concentration_ng_ml)))

print(plot_simplified(total_std_solution_concentration_ng_ml, "Combined Standard Solution Distribution", "#A9CCE3", "Log-Normal"))

# ===============================================================

# SECTION 3: Method Precision Uncertainty

# ===============================================================

# This section calculates the uncertainty due to method precision (repeatability and intermediate precision)

precision_data <- list(

conc_1 = list(

c(0.990058595, 0.969861821, 0.98347695, 1.034818496, 1.032660226),

c(1.138706966, 1.00037923, 0.940233536, 0.92599092, 1.128141252),

c(0.993379925, 0.906422564, 0.99034435, 0.836954836, 0.888823433)

),

conc_3 = list(

c(2.738931867, 3.045891098, 2.899602305, 2.694816676, 2.72664133),

c(2.850518242, 2.646686445, 2.659048705, 2.71863313, 2.435286918),

c(3.226061984, 3.239515079, 3.131088664, 3.335182641, 3.3793295)

),

conc_15 = list(

c(13.7319542, 15.13334397, 13.65695501, 13.51227313, 14.00044143),

c(16.83798972, 15.90572623, 14.39205711, 16.32646591, 15.81917022),

c(16.75662775, 16.91359894, 17.49996186, 16.77564977, 17.19078607)

),

conc_30 = list(

c(26.94300342, 28.0631292, 26.26548207, 24.29185659, 25.30834028),

c(32.02694838, 28.05341024, 30.28075433, 30.45425755, 29.06047413),

c(35.72453005, 33.72962173, 33.90589398, 34.03771662, 32.82783961)

)

)

# Function to simulate values for a given concentration level

simulate_level <- function(level_list, n_obs=5) {

sim_values <- c()

for (ex in level_list) {

mu <- mean(ex)

sigma <- sd(ex)

sim_values <- c(sim_values, rnorm(n_obs, mu, sigma))

}

return(sim_values)

}

# Run simulation to estimate the pooled RSU from precision data

rsu_vals <- numeric(N_sim)

for (i in 1:N_sim) {

rsus_per_level <- numeric(length(precision_data))

idx <- 1

for (lvl in precision_data) {

sim_vals <- simulate_level(lvl)

groups <- split(sim_vals, rep(1:length(lvl), each=5))

exp_means <- sapply(groups, mean)

exp_vars <- sapply(groups, var)

std_unc <- sqrt(mean(exp_vars))

overall_mean <- mean(exp_means)

rsus_per_level[idx] <- std_unc / overall_mean

idx <- idx + 1

}

rsu_vals[i] <- sqrt(sum(rsus_per_level^2)) # Combine RSUs from all levels

}

# Calculate final RSU for method precision

rsu_method_prec <- mean(rsu_vals)

cat("=== Factor 3: Method Precision Uncertainty (MCM) ===\n")

cat(sprintf("Combined RSU: %.7f\n", rsu_method_prec))

print(plot_simplified(rsu_vals, "Method Precision RSU Distribution (dimensionless)", "powderblue", "Empirical"))

# ===============================================================

# SECTION 4: Calibration Curve Uncertainty

# ===============================================================

# This section models the uncertainty arising from the calibration curve fitting process.

calib_conc <- c(1, 2, 5, 10, 20, 50) # Calibrator concentrations (ng/mL)

calib_resp <- c(0.26289436, 0.48415553, 1.17696021, 2.16418778, 4.77564046, 12.48041737) # Calibrator responses

y_sd_est   <- 0.03 * calib_resp      # Estimated standard deviation of responses

sample_resp <- 1.96853468           # Response of the unknown sample (yj)

# Simulate variation in calibrator responses

calib_mat <- matrix(nrow = N_sim, ncol = length(calib_conc))

for (j in 1:length(calib_conc)) {

calib_mat[,j] <- rnorm(N_sim, calib_resp[j], y_sd_est[j])

}

# For each simulated curve, calculate the concentration of the unknown sample

xj_mcm <- numeric(N_sim)

for (k_iter in 1:N_sim) {

fit <- lm(calib_mat[k_iter,] ~ calib_conc)

a <- coef(fit)[1] # Intercept

b <- coef(fit)[2] # Slope

xj <- (sample_resp - a) / b # Calculate concentration

# Use '&' for vectorized logical comparison

xj_mcm[k_iter] <- ifelse(is.finite(xj) & xj > 0 & xj < 100, xj, NA)

}

xj_mcm <- xj_mcm[!is.na(xj_mcm)] # Remove invalid results

# Calculate RSU from the calibration process

rsu_calib <- sd(xj_mcm) / mean(xj_mcm)

cat("=== Factor 4: Calibration Curve Uncertainty (MCM) ===\n")

cat(sprintf("RSU (Calibration): %.7f\n", rsu_calib))

cat(sprintf("Mean calibration result: %.6f ng/mL\n", mean(xj_mcm)))

print(plot_simplified(xj_mcm, "Calibration Curve Output (ng/mL)", "lightcoral", "Empirical"))

# ===============================================================

# SECTION 5: Final Uncertainty: Complete Monte Carlo Propagation

# ===============================================================

# This section combines all individual uncertainty sources into a single, integrated simulation

# to determine the final measurement result and its combined uncertainty.

cat("=== INTEGRATED MONTE CARLO SIMULATION ===\n")

final_concentration_mcm <- numeric(N_sim)

for (i in 1:N_sim) {

# 1. Sample Preparation Uncertainty Factor

pip1_ul <- runif(1, pip_vol_ul[1] - pip_cert_unc_ul[1], pip_vol_ul[1] + pip_cert_unc_ul[1])

pip2_ul <- runif(1, pip_vol_ul[2] - pip_cert_unc_ul[2], pip_vol_ul[2] + pip_cert_unc_ul[2])

vol_factor <- (pip1_ul + pip2_ul) / sum(pip_vol_ul)

# 2. Standard Solution Uncertainty Factor (using the pre-simulated ng/uL result)

std_sol_factor <- total_std_solution_concentration_ng_ml[i] / mean(total_std_solution_concentration_ng_ml)

# 3. Method Precision Uncertainty Factor

precision_factor <- rnorm(1, 1, rsu_method_prec)

# 4. Calibration and Measurement Uncertainty (Simulated for each trial)

calib_resp_sim <- calib_resp + rnorm(length(calib_resp), 0, y_sd_est)

sample_resp_sim <- sample_resp + rnorm(1, 0, 0.03 * sample_resp)

fit <- lm(calib_resp_sim ~ calib_conc)

a <- coef(fit)[1]

b <- coef(fit)[2]

# Combine all factors to get the final concentration for this simulation run

if (is.finite(a) && is.finite(b) && b != 0) {

measured_conc <- (sample_resp_sim - a) / b

final_concentration_mcm[i] <- measured_conc * vol_factor * std_sol_factor * precision_factor

} else {

final_concentration_mcm[i] <- NA

}

}

# Filter out non-physical or failed simulation results

final_concentration_mcm <- final_concentration_mcm[is.finite(final_concentration_mcm) &

final_concentration_mcm > 0 &

final_concentration_mcm < 50]

# Final Results Calculation and Reporting

mean_sim <- mean(final_concentration_mcm)

combined_std_unc <- sd(final_concentration_mcm) # Combined standard uncertainty

combined_rsu <- combined_std_unc / mean_sim        # Combined RSU

# Calculate 95% confidence interval and expanded uncertainty

confidence_interval_95 <- quantile(final_concentration_mcm, c(0.025, 0.975))

expanded_unc_gum_style <- k * combined_std_unc

# Print the final results

cat("=== FINAL RESULTS ===\n")

cat(sprintf("Analytical Result: %.3f ng/mL\n", mean_sim))

cat(sprintf("Standard Uncertainty: %.6f ng/mL\n", combined_std_unc))

cat(sprintf("95%% Confidence Interval: %.3f to %.3f ng/mL\n",

confidence_interval_95[1], confidence_interval_95[2]))

cat(sprintf("Expanded Uncertainty (k=2): %.6f ng/mL\n", expanded_unc_gum_style))

cat(sprintf("Relative Standard Uncertainty: %.7f (%.2f%%)\n", combined_rsu, 100 * combined_rsu))

# Final Plotting Function

plot_final_simplified <- function(data, analytical_mean, std_unc, ci_95, title) {

d <- density(data)

df <- data.frame(x = d$x, y = d$y)

skewness_coeff <- skewness(data) # Calculate skewness

# Create detailed annotation text

annotation_text <- paste(

sprintf("Mean: %.2f ng/mL", analytical_mean),

sprintf("Standard Uncertainty: %.7f ng/mL", std_unc),

sprintf("95%% CI: %.2f - %.2f ng/mL", ci_95[1], ci_95[2]),

"Distribution: Empirical",

sprintf("Skewness: %.3f", skewness_coeff),

sep = "\n"

)

# Position the annotation in the upper left corner

x_pos <- min(d$x) + 0.005 * (max(d$x) - min(d$x))

y_pos <- max(d$y) * 0.995

# Generate the final plot

ggplot(df, aes(x = x, y = y)) +

geom_ribbon(aes(ymin = 0, ymax = y), fill = "#A8DADC", alpha = 0.6) +

geom_line(color = "#457B9D", size = 1.2) +

geom_vline(xintercept = analytical_mean, color = "#E63946", linetype = "dashed", size = 0.8) +

labs(title = title, x = "Concentration (ng/mL)", y = "Density") +

annotate("text", x = x_pos, y = y_pos,

label = annotation_text, hjust = 0, vjust = 1, size = 3.5, color = "#2f2f2f") +

theme_minimal(base_size = 12) +

theme(plot.title = element_text(size = 16, face = "bold", hjust = 0.5),

axis.title = element_text(size = 12),

axis.text = element_text(size = 11))

}

# Create the final distribution plot object

final_plot <- plot_final_simplified(

data = final_concentration_mcm,

analytical_mean = mean_sim,

std_unc = combined_std_unc,

ci_95 = confidence_interval_95,

title = "Probability density function of the output quantity (MCM)"

)

# Plot saving workflow

# Step 1: Open the graphics device with specified properties.

agg_png("MCM_uncertainty_distribution.png",

width = 10,

height = 6,

units = "in",

res = 600)

# Step 2: Print the created ggplot object to the active device.

print(final_plot)

# Step 3: Close the device, which saves the file to disk.

dev.off()

# Print the plot

print(final_plot)