Supplemental Table 1: Mean and standard deviation of SNR values (3 replicates) obtained for all spheres of the IEC phantoms visualized by both cameras (≥ 8-mm in diameter). Values are given for analog- and digital-PET images obtained with and without TOF (TOF and noTOF respectively) and with 1 to 10 OSEM iterations, as well as with NEMA-defined reconstruction parameters.

| SNR |  |  | Mean |  |  |  |  |  |  |  |  | SD |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | Hot sphere diameters (mm) |  |  |  |  |  |  | Cold sphere diameters (mm) |  | Hot sphere diameters (mm) |  |  |  |  |  |  | Cold sphere diameters (mm) |  |
|  |  | Number of Iterations | 8 | 10 | 13 | 17 | 22 | 28 | 37 | 28 | 37 | 8 | 10 | 13 | 17 | 22 | 28 | 37 | 28 | 37 |
| noTOF | Digital | 1 | 4.3 | 5.9 | 8.9 | 16.6 | 22.5 | 27.4 | 35.6 | 8.1 | 7.9 | 0.5 | 1.8 | 1.0 | 2.2 | 3.4 | 0.9 | 3.0 | 1.3 | 1.0 |
|  |  | 2 | 4.4 | 7.4 | 12.8 | 19.6 | 25.6 | 29.1 | 34.3 | 7.6 | 7.9 | 0.5 | 0.8 | 1.1 | 1.6 | 1.5 | 1.1 | 2.3 | 0.7 | 0.5 |
|  |  | 3 | 4.3 | 8.0 | 13.5 | 18.7 | 22.9 | 25.5 | 28.1 | 6.4 | 6.9 | 0.5 | 0.8 | 0.9 | 1.2 | 0.9 | 0.7 | 1.6 | 0.4 | 0.3 |
|  |  | 4 | 4.2 | 8.2 | 13.3 | 17.3 | 20.4 | 22.5 | 24.0 | 5.7 | 6.1 | 0.4 | 0.7 | 0.8 | 1.0 | 0.7 | 0.5 | 1.2 | 0.3 | 0.2 |
|  |  | 5 | 4.0 | 8.2 | 12.8 | 16.0 | 18.3 | 20.1 | 21.2 | 5.1 | 5.6 | 0.4 | 0.6 | 0.7 | 0.8 | 0.6 | 0.3 | 0.9 | 0.3 | 0.2 |
|  |  | 6 | 3.9 | 8.1 | 12.1 | 14.8 | 16.7 | 18.3 | 19.2 | 4.7 | 5.2 | 0.4 | 0.5 | 0.6 | 0.7 | 0.5 | 0.2 | 0.7 | 0.2 | 0.2 |
|  |  | 7 | 3.7 | 7.9 | 11.5 | 13.8 | 15.5 | 17.0 | 17.7 | 4.4 | 4.8 | 0.4 | 0.4 | 0.6 | 0.6 | 0.5 | 0.1 | 0.6 | 0.2 | 0.2 |
|  |  | 8 | 3.6 | 7.7 | 10.9 | 12.9 | 14.5 | 15.8 | 16.4 | 4.2 | 4.6 | 0.4 | 0.4 | 0.5 | 0.6 | 0.4 | 0.0 | 0.5 | 0.2 | 0.1 |
|  |  | 9 | 3.5 | 7.5 | 10.4 | 12.2 | 13.6 | 14.9 | 15.4 | 4.0 | 4.3 | 0.4 | 0.3 | 0.5 | 0.5 | 0.4 | 0.0 | 0.4 | 0.2 | 0.1 |
|  |  | 10 | 3.4 | 7.3 | 9.9 | 11.6 | 12.9 | 14.1 | 14.6 | 3.8 | 4.1 | 0.4 | 0.3 | 0.5 | 0.5 | 0.3 | 0.1 | 0.4 | 0.2 | 0.1 |
|  | Analog | 1 | 4.1 | 3.3 | 5.9 | 12.5 | 16.9 | 18.6 | 30.5 | 7.5 | 8.1 | 0.3 | 2.6 | 4.1 | 4.1 | 5.9 | 0.8 | 2.1 | 0.5 | 0.1 |
|  |  | 2 | 4.2 | 5.7 | 11.1 | 17.9 | 23.8 | 27.0 | 33.2 | 8.2 | 9.0 | 0.3 | 2.4 | 2.1 | 1.8 | 1.9 | 1.9 | 1.6 | 0.4 | 0.7 |
|  |  | 3 | 4.2 | 6.8 | 12.3 | 17.9 | 22.3 | 25.4 | 28.2 | 7.2 | 8.0 | 0.3 | 2.1 | 1.0 | 1.3 | 1.8 | 1.8 | 2.5 | 0.3 | 0.5 |
|  |  | 4 | 4.1 | 6.9 | 12.2 | 16.8 | 19.9 | 22.7 | 25.1 | 6.4 | 7.1 | 0.3 | 1.9 | 0.5 | 1.3 | 1.7 | 1.6 | 2.2 | 0.3 | 0.5 |
|  |  | 5 | 4.0 | 6.8 | 11.6 | 15.3 | 17.6 | 20.4 | 22.5 | 5.5 | 6.2 | 0.3 | 1.7 | 0.5 | 1.5 | 2.1 | 1.4 | 2.0 | 0.6 | 0.8 |
|  |  | 6 | 3.9 | 6.8 | 11.2 | 14.5 | 16.3 | 18.6 | 20.5 | 5.3 | 6.0 | 0.2 | 1.6 | 0.4 | 1.2 | 1.4 | 1.3 | 1.8 | 0.3 | 0.4 |
|  |  | 7 | 3.8 | 6.6 | 10.6 | 13.5 | 15.1 | 17.1 | 18.9 | 5.0 | 5.6 | 0.2 | 1.6 | 0.4 | 1.1 | 1.3 | 1.2 | 1.7 | 0.2 | 0.4 |
|  |  | 8 | 3.8 | 6.4 | 10.0 | 12.7 | 14.1 | 15.9 | 17.6 | 4.7 | 5.3 | 0.2 | 1.5 | 0.4 | 1.0 | 1.2 | 1.1 | 1.6 | 0.2 | 0.4 |
|  |  | 9 | 3.7 | 6.2 | 9.6 | 12.0 | 13.3 | 15.0 | 16.6 | 4.5 | 5.0 | 0.2 | 1.4 | 0.4 | 1.0 | 1.1 | 1.0 | 1.5 | 0.2 | 0.4 |
|  |  | 10 | 3.7 | 6.2 | 9.2 | 11.7 | 12.8 | 14.2 | 15.7 | 4.4 | 4.9 | 0.2 | 1.4 | 0.4 | 0.9 | 0.8 | 1.0 | 1.4 | 0.0 | 0.1 |
| TOF | Digital | 1 | 7.1 | 14.5 | 24.0 | 30.6 | 36.5 | 40.1 | 43.5 | 10.6 | 12.0 | 0.8 | 1.5 | 1.0 | 1.2 | 0.5 | 0.0 | 0.9 | 0.1 | 0.2 |
|  |  | 2 | 7.2 | 14.1 | 20.3 | 24.0 | 26.9 | 29.0 | 30.7 | 8.1 | 8.8 | 0.8 | 1.4 | 0.8 | 1.1 | 0.8 | 0.3 | 0.4 | 0.2 | 0.2 |
|  |  | 3 | 6.9 | 12.9 | 17.3 | 19.9 | 21.9 | 23.6 | 24.9 | 6.8 | 7.3 | 0.7 | 1.2 | 0.7 | 0.9 | 0.7 | 0.3 | 0.2 | 0.2 | 0.2 |
|  |  | 4 | 6.6 | 11.8 | 15.3 | 17.4 | 19.1 | 20.5 | 21.6 | 6.1 | 6.5 | 0.7 | 1.1 | 0.6 | 0.8 | 0.7 | 0.3 | 0.1 | 0.2 | 0.2 |
|  |  | 5 | 6.2 | 11.0 | 14.0 | 15.7 | 17.2 | 18.5 | 19.5 | 5.6 | 5.9 | 0.7 | 1.0 | 0.6 | 0.8 | 0.6 | 0.3 | 0.1 | 0.2 | 0.2 |
|  |  | 6 | 5.9 | 10.3 | 13.0 | 14.5 | 15.9 | 17.1 | 18.0 | 5.2 | 5.5 | 0.6 | 0.9 | 0.6 | 0.7 | 0.6 | 0.3 | 0.0 | 0.2 | 0.2 |
|  |  | 7 | 5.7 | 9.7 | 12.2 | 13.7 | 15.0 | 16.1 | 16.9 | 5.0 | 5.2 | 0.6 | 0.8 | 0.5 | 0.7 | 0.5 | 0.3 | 0.0 | 0.1 | 0.2 |
|  |  | 8 | 5.5 | 9.3 | 11.6 | 13.0 | 14.2 | 15.3 | 16.1 | 4.7 | 4.9 | 0.6 | 0.8 | 0.5 | 0.7 | 0.5 | 0.3 | 0.0 | 0.1 | 0.2 |
|  |  | 9 | 5.3 | 8.9 | 11.1 | 12.5 | 13.6 | 14.7 | 15.4 | 4.6 | 4.7 | 0.6 | 0.7 | 0.5 | 0.6 | 0.5 | 0.3 | 0.0 | 0.1 | 0.2 |
|  |  | 10 | 5.1 | 8.6 | 10.8 | 12.0 | 13.1 | 14.1 | 14.8 | 4.4 | 4.6 | 0.5 | 0.7 | 0.5 | 0.6 | 0.5 | 0.3 | 0.0 | 0.1 | 0.2 |
|  |  | NEMA |  | 10.8 | 12.8 | 15.4 | 16.3 | 18.0 | 18.9 | 5.5 | 5.9 |  | 1.3 | 0.8 | 0.7 | 0.8 | 0.3 | 0.2 | 0.2 | 0.2 |
|  | Analog | 1 | 4.3 | 8.8 | 16.7 | 23.9 | 31.2 | 37.5 | 44.4 | 11.0 | 12.3 | 0.3 | 1.6 | 1.6 | 1.8 | 2.0 | 1.5 | 2.6 | 0.0 | 0.3 |
|  |  | 2 | 4.8 | 10.1 | 16.9 | 22.1 | 25.7 | 29.1 | 31.9 | 7.8 | 8.6 | 0.3 | 1.6 | 1.5 | 1.3 | 1.4 | 1.2 | 1.8 | 0.9 | 0.8 |
|  |  | 3 | 4.9 | 9.7 | 15.2 | 19.0 | 21.3 | 23.6 | 25.5 | 6.8 | 7.4 | 0.3 | 1.6 | 1.4 | 1.1 | 1.1 | 0.9 | 1.5 | 0.5 | 0.4 |
|  |  | 4 | 5.0 | 9.2 | 13.7 | 16.8 | 18.6 | 20.5 | 22.1 | 6.2 | 6.7 | 0.3 | 1.5 | 1.2 | 0.9 | 1.0 | 0.8 | 1.3 | 0.3 | 0.2 |
|  |  | 5 | 4.9 | 8.6 | 12.5 | 15.1 | 16.6 | 18.0 | 19.1 | 5.7 | 6.1 | 0.3 | 1.4 | 1.1 | 0.8 | 1.0 | 0.7 | 1.1 | 0.2 | 0.0 |
|  |  | 6 | 4.9 | 8.3 | 11.7 | 14.0 | 15.4 | 16.8 | 18.2 | 5.3 | 5.7 | 0.3 | 1.3 | 1.0 | 0.8 | 0.9 | 0.7 | 1.0 | 0.1 | 0.0 |
|  |  | 7 | 4.8 | 7.9 | 11.0 | 13.1 | 14.3 | 15.6 | 16.9 | 5.0 | 5.4 | 0.3 | 1.2 | 1.0 | 0.7 | 0.8 | 0.6 | 1.0 | 0.1 | 0.0 |
|  |  | 8 | 4.8 | 7.6 | 10.4 | 12.4 | 13.5 | 14.7 | 15.9 | 4.8 | 5.1 | 0.3 | 1.2 | 0.9 | 0.7 | 0.8 | 0.6 | 0.9 | 0.1 | 0.1 |
|  |  | 9 | 4.7 | 7.3 | 9.9 | 11.8 | 12.9 | 14.0 | 15.1 | 4.6 | 4.9 | 0.3 | 1.1 | 0.9 | 0.6 | 0.8 | 0.6 | 0.9 | 0.0 | 0.1 |
|  |  | 10 | 4.7 | 7.0 | 9.6 | 11.3 | 12.4 | 13.3 | 14.4 | 4.5 | 4.8 | 0.3 | 1.1 | 0.8 | 0.6 | 0.7 | 0.5 | 0.8 | 0.0 | 0.1 |
|  |  | NEMA |  | 9.6 | 13.2 | 15.7 | 17 | 18.4 | 19.8 | 6.1 | 6.5 |  | 0.7 | 1.3 | 0.9 | 0.6 | 0.3 | 0.3 | 0.1 | 0.3 |