Table S1. FWHM and FWTM measured at the center of the FOV and with a 5cm radial and tangential offset. Radial, tangential and axial resolution measurements are expressed in mm and presented for the different reconstruction settings.
	
	Centre
	X=+5, y=+5

	
	Radial
	Tangential
	Axial
	Radial
	Tangential
	Axial

	
	FWHM
	FWTM
	FWHM
	FWTM
	FWHM
	FWTM
	FWHM
	FWTM
	FWHM
	FWTM
	FWHM
	FWTM

	[bookmark: _GoBack]FBP
	3,86
	7,03
	3,90
	7,08
	3,91
	7,13
	4,20
	7,69
	4,23
	7,77
	4,22
	7,78

	3D-OSEM i5s5
	3,25
	5,92
	3,36
	6,13
	3,37
	6,37
	3,65
	6,60
	3,67
	6,64
	4,58
	6,67

	3D-OSEM i10s5
	3,20
	5,80
	3,30
	5,99
	3,30
	6,11
	3,63
	6,58
	3,64
	6,60
	4,11
	6,67

	3D-OSEM i15s5
	3,16
	5,75
	3,22
	5,89
	3,21
	5,87
	3,60
	6,57
	3,62
	6,64
	3,96
	6,65

	3D-OSEM i20s5
	3,13
	5,68
	3,17
	5,78
	3,17
	5,77
	3,59
	6,56
	3,62
	6,60
	3,67
	6,64

	3D-OSEM i25s5
	3,10
	5,66
	3,15
	5,73
	3,15
	5,75
	3,59
	6,55
	3,61
	6,58
	3,82
	6,62

	3D-OSEM i30s5
	3,06
	5,62
	3,13
	5,70
	3,14
	5,74
	3,58
	6,53
	3,60
	6,57
	3,78
	6,61

	PSF i5s5
	2,29
	3,92
	2,19
	3,70
	2,37
	3,86
	2,16
	3,89
	2,24
	4,03
	2,24
	4,16

	PSF i10s5
	1,91
	3,71
	1,86
	3,41
	1,89
	3,55
	1,86
	3,39
	1,92
	3,50
	1,96
	3,60

	PSF i15s5
	1,85
	3,53
	1,83
	3,32
	1,86
	3,50
	1,79
	3,24
	1,82
	3,32
	1,84
	3,36

	PSF i20s5
	1,80
	3,35
	1,79
	3,28
	1,85
	3,48
	1,78
	3,22
	1,79
	3,27
	1,81
	3,50

	PSF i25s5
	1,79
	3,34
	1,78
	3,25
	1,84
	3,45
	1,77
	3,20
	1,77
	3,22
	1,79
	3,27

	PSF i30s5
	1,80
	3,35
	1,77
	3,24
	1,83
	3,45
	1,78
	3,21
	1,78
	3,25
	1,79
	3,27



Table S2. Image quality parameters calculated for the 18F-filled NEMA NU 4-2008 phantom positioned in the center of the FOV (C-SC-). Results are presented for the different reconstruction settings tested.
	Recons
	SORwater
	%SD water
	SORair
	% SD air
	Uniformity %SD
	RC 5mm
	%SD 5mm
	RC 4mm
	%SD 4mm
	RC 3mm
	%SD 3mm
	RC 2mm
	%SD 2mm
	RC 1mm
	%SD 1mm

	FBP
	0,18
	11,49
	0,11
	20,45
	1,74
	0,64
	1,85
	0,49
	2,25
	0,29
	2,82
	0,16
	4,19
	0,04
	13,17

	3D-OSEM i3s5
	0,24
	5,98
	0,19
	7,44
	1,39
	0,77
	2,28
	0,64
	2,06
	0,39
	2,65
	0,20
	3,07
	0,03
	8,91

	3D-OSEM i5s5
	0,20
	6,56
	0,15
	8,73
	1,75
	0,76
	2,77
	0,63
	2,42
	0,41
	2,69
	0,22
	3,11
	0,04
	7,31

	3D-OSEM i10s5
	0,17
	7,29
	0,13
	9,76
	2,21
	0,74
	3,53
	0,62
	2,92
	0,41
	3,15
	0,22
	3,56
	0,06
	6,74

	3D-OSEM i15s5
	0,17
	7,91
	0,10
	12,85
	2,49
	0,74
	3,94
	0,63
	3,17
	0,41
	3,59
	0,22
	3,85
	0,06
	6,84

	3D-OSEM i20s5
	0,17
	8,51
	0,10
	14,29
	2,69
	0,74
	4,21
	0,63
	3,33
	0,41
	3,90
	0,22
	4,06
	0,06
	6,94

	3D-OSEM i25s5
	0,17
	9,08
	0,10
	15,47
	2,85
	0,74
	4,38
	0,63
	3,46
	0,41
	4,15
	0,22
	4,22
	0,07
	7,06

	3D-OSEM i30s5
	0,17
	9,60
	0,10
	16,46
	2,98
	0,74
	4,50
	0,63
	3,58
	0,41
	4,34
	0,22
	4,33
	0,07
	7,12

	PSF i3s5
	0,23
	7,34
	0,19
	8,45
	4,46
	1,12
	7,53
	0,78
	6,98
	0,38
	8,75
	0,15
	10,74
	0,01
	8,76

	PSF i5s5
	0,17
	9,84
	0,13
	10,42
	4,55
	1,37
	7,35
	1,10
	7,36
	0,53
	11,36
	0,25
	11,64
	0,02
	8,44

	PSF i10s5
	0,12
	13,00
	0,08
	13,09
	5,15
	1,60
	7,63
	1,43
	7,11
	0,76
	9,84
	0,40
	13,59
	0,04
	9,03

	PSF i15s5
	0,10
	13,78
	0,06
	14,58
	5,57
	1,66
	8,07
	1,54
	7,07
	0,85
	9,28
	0,46
	13,16
	0,06
	11,66

	PSF i20s5
	0,10
	13,71
	0,05
	15,95
	5,55
	1,67
	7,55
	1,58
	6,42
	0,90
	8,56
	0,50
	13,23
	0,08
	13,43

	PSF i25s5
	0,10
	13,39
	0,05
	16,35
	5,73
	1,66
	7,32
	1,59
	6,21
	0,93
	7,97
	0,51
	13,30
	0,09
	14,20

	PSF i30s5
	0,10
	12,92
	0,04
	16,59
	5,85
	1,65
	7,31
	1,60
	6,46
	0,96
	7,87
	0,53
	12,78
	0,10
	14,58

	PSF i20s5 Gauss 2mm
	0.10
	14.97
	0.06
	17.56
	3.47
	1.10
	3.85
	0.92
	3.12
	0.56
	4.86
	0.28
	8.59
	0.06
	10.85



Table S3. Image quality parameters calculated for the 18F-filled NEMA NU 4-2008 phantom in off-center position with 3 scattering sources placed in the FOV (OC-SC+). Results are presented for the different reconstruction settings tested.
	Recons
	SORwater
	%SD water
	SORair
	% SD air
	Uniformity %SD
	RC 5mm
	%SD 5mm
	RC 4mm
	%SD 4mm
	RC 3mm
	%SD 3mm
	RC 2mm
	%SD 2mm
	RC 1mm
	%SD 1mm

	FBP
	0,17
	18,57
	0,08
	44,06
	3,79
	0,56
	4,30
	0,45
	5,63
	0,30
	7,73
	0,14
	9,32
	0,03
	28,43

	3D-OSEM i3s5
	0,24
	11,19
	0,20
	12,61
	2,21
	0,67
	2,36
	0,58
	3,24
	0,35
	2,36
	0,18
	5,18
	0,03
	9,02

	3D-OSEM i5s5
	0,20
	11,89
	0,15
	13,71
	2,58
	0,67
	3,29
	0,59
	3,72
	0,37
	2,83
	0,22
	5,38
	0,04
	8,76

	3D-OSEM i10s5
	0,17
	13,14
	0,12
	15,53
	3,15
	0,66
	3,75
	0,57
	4,32
	0,38
	3,85
	0,22
	6,04
	0,06
	10,25

	3D-OSEM i15s5
	0,17
	14,08
	0,11
	16,88
	3,53
	0,66
	4,05
	0,57
	4,70
	0,37
	4,61
	0,22
	6,58
	0,06
	11,67

	3D-OSEM i20s5
	0,16
	14,83
	0,10
	18,03
	3,81
	0,67
	4,30
	0,57
	4,98
	0,37
	5,15
	0,22
	6,97
	0,06
	12,68

	3D-OSEM i25s5
	0,10
	25,12
	0,10
	19,09
	4,02
	0,67
	4,49
	0,57
	5,20
	0,37
	5,54
	0,21
	7,26
	0,06
	13,48

	3D-OSEM i30s5
	0,10
	25,99
	0,01
	200,57
	4,19
	0,67
	4,64
	0,57
	5,38
	0,37
	5,83
	0,21
	7,48
	0,06
	14,15

	PSF i3s5
	0,25
	6,62
	0,21
	7,24
	3,80
	0,97
	6,43
	0,67
	5,93
	0,34
	6,67
	0,13
	9,85
	0,01
	9,51

	PSF i5s5
	0,19
	8,61
	0,15
	9,55
	3,76
	1,20
	6,52
	0,93
	6,02
	0,48
	8,20
	0,22
	6,67
	0,02
	9,22

	PSF i10s5
	0,13
	11,78
	0,09
	12,84
	4,61
	1,35
	7,31
	1,29
	6,11
	0,67
	7,70
	0,38
	9,59
	0,04
	10,61

	PSF i15s5
	0,10
	14,02
	0,07
	15,51
	5,25
	1,41
	9,76
	1,44
	7,36
	0,76
	8,23
	0,42
	9,92
	0,06
	10,30

	PSF i20s5
	0,10
	14,62
	0,06
	17,54
	4,52
	1,43
	9,86
	1,53
	7,23
	0,87
	9,30
	0,51
	11,30
	0,08
	10,19

	PSF i25s5
	0,09
	15,00
	0,05
	18,65
	4,73
	1,43
	9,93
	1,54
	7,43
	0,92
	10,09
	0,56
	13,55
	0,08
	10,28

	PSF i30s5
	0,09
	15,18
	0,05
	19,43
	4,91
	1,42
	9,85
	1,58
	7,65
	0,95
	10,41
	0,59
	14,99
	0,10
	11,57



Table S4. Image quality parameters calculated for the 68Ga-filled NEMA NU 4-2008 phantom positioned in the center of the FOV (C-SC-). Results are presented for the different reconstruction settings tested.
	Recons
	SORwater
	%SD water
	SORair
	% SD air
	Uniformity %SD
	RC 5mm
	%SD 5mm
	RC 4mm
	%SD 4mm
	RC 3mm
	%SD 3mm
	RC 2mm
	%SD 2mm
	RC 1mm
	%SD 1mm

	FBP
	0,19
	18,56
	0,14
	18,52
	3,43
	0,44
	3,97
	0,33
	3,78
	0,20
	5,33
	0,10
	6,84
	-
	-

	3D-OSEM i3s5
	0,24
	12,24
	0,21
	11,10
	2,98
	0,49
	3,49
	0,38
	3,03
	0,23
	3,94
	0,10
	5,43
	-
	-

	3D-OSEM i5s5
	0,21
	13,23
	0,17
	11,80
	3,14
	0,48
	3,85
	0,38
	3,25
	0,24
	4,67
	0,12
	6,45
	-
	-

	3D-OSEM i10s5
	0,18
	14,76
	0,14
	13,35
	3,45
	0,48
	4,33
	0,38
	3,59
	0,24
	5,81
	0,12
	8,02
	-
	-

	3D-OSEM i15s5
	0,17
	16,69
	0,12
	14,00
	3,67
	0,47
	4,63
	0,38
	3,83
	0,24
	6,59
	0,12
	8,86
	-
	-

	3D-OSEM i20s5
	0,17
	17,33
	0,12
	14,77
	3,84
	0,47
	4,85
	0,38
	4,04
	0,24
	7,14
	0,12
	9,60
	-
	-

	3D-OSEM i25s5
	0,17
	17,80
	0,12
	15,46
	3,98
	0,47
	5,02
	0,38
	4,21
	0,24
	7,55
	0,12
	10,25
	-
	-

	3D-OSEM i30s5
	0,16
	18,18
	0,11
	16,08
	4,10
	0,47
	5,16
	0,38
	4,37
	0,24
	7,84
	0,12
	10,82
	-
	-

	PSF i3s5
	0,24
	10,04
	0,21
	8,08
	1,51
	0,62
	2,36
	0,39
	3,17
	0,23
	5,38
	0,09
	6,53
	-
	-

	PSF i5s5
	0,19
	11,22
	0,16
	9,65
	1,53
	0,70
	2,60
	0,48
	2,36
	0,30
	2,90
	0,13
	5,31
	-
	-

	PSF i10s5
	0,14
	13,40
	0,11
	13,03
	2,28
	0,75
	3,77
	0,58
	2,99
	0,38
	2,46
	0,18
	4,44
	-
	-

	PSF i15s5
	0,12
	14,76
	0,09
	15,34
	2,63
	0,75
	4,05
	0,59
	3,56
	0,41
	3,24
	0,20
	5,12
	-
	-

	PSF i20s5
	0,12
	15,67
	0,08
	17,01
	2,81
	0,74
	4,14
	0,59
	3,73
	0,42
	3,92
	0,22
	6,07
	-
	-

	PSF i25s5
	0,11
	16,30
	0,07
	18,28
	2,94
	0,73
	4,16
	0,58
	3,78
	0,43
	4,28
	0,22
	6,91
	-
	-

	PSF i30s5
	0,11
	15,97
	0,07
	20,05
	3,04
	0,73
	4,16
	0,58
	3,83
	0,44
	4,50
	0,23
	7,66
	-
	-



Table S5. Image quality parameters calculated for the 68Ga-filled NEMA NU 4-2008 phantom in off-center position with 3 scattering sources placed in the FOV (OC-SC+). Results are presented for the different reconstruction settings tested.
	Recons
	SORwater
	%SD water
	SORair
	% SD air
	Uniformity %SD
	RC 5mm
	%SD 5mm
	RC 4mm
	%SD 4mm
	RC 3mm
	%SD 3mm
	RC 2mm
	%SD 2mm
	RC 1mm
	%SD 1mm

	FBP
	0,20
	19,15
	0,14
	27,01
	5,90
	0,31
	3,78
	0,25
	6,15
	0,16
	9,82
	0,08
	20,56
	-
	-

	3D-OSEM i3s5
	0,24
	11,16
	0,21
	9,98
	3,50
	0,45
	4,11
	0,33
	3,76
	0,21
	4,35
	0,09
	8,07
	-
	-

	3D-OSEM i5s5
	0,22
	12,58
	0,18
	10,65
	3,93
	0,48
	4,71
	0,36
	4,26
	0,24
	4,93
	0,11
	7,93
	-
	-

	3D-OSEM i10s5
	0,19
	15,22
	0,15
	12,18
	4,33
	0,48
	5,39
	0,36
	4,77
	0,24
	5,46
	0,12
	8,80
	-
	-

	3D-OSEM i15s5
	0,18
	17,08
	0,14
	13,45
	4,61
	0,48
	5,77
	0,36
	5,14
	0,24
	5,59
	0,12
	9,32
	-
	-

	3D-OSEM i20s5
	0,18
	18,49
	0,13
	14,46
	4,84
	0,48
	6,03
	0,36
	5,43
	0,24
	5,65
	0,12
	9,66
	-
	-

	3D-OSEM i25s5
	0,18
	19,61
	0,13
	15,32
	5,02
	0,48
	6,22
	0,36
	5,65
	0,24
	5,70
	0,12
	9,88
	-
	-

	3D-OSEM i30s5
	0,18
	20,53
	0,13
	16,05
	5,16
	0,47
	6,37
	0,36
	5,83
	0,24
	5,76
	0,12
	10,03
	-
	-

	PSF i3s5
	0,26
	7,85
	0,23
	7,39
	1,79
	0,62
	3,76
	0,38
	3,54
	0,21
	5,37
	0,07
	5,58
	-
	-

	PSF i5s5
	0,20
	9,52
	0,17
	9,25
	1,77
	0,72
	3,21
	0,47
	3,04
	0,27
	4,50
	0,11
	5,80
	-
	-

	PSF i10s5
	0,15
	12,24
	0,12
	12,35
	2,48
	0,77
	3,38
	0,55
	3,45
	0,36
	4,40
	0,16
	5,33
	-
	-

	PSF i15s5
	0,13
	13,84
	0,10
	14,32
	2,87
	0,77
	3,83
	0,56
	3,76
	0,40
	4,21
	0,19
	5,24
	-
	-

	PSF i20s5
	0,12
	14,95
	0,08
	15,70
	2,62
	0,76
	3,88
	0,56
	3,67
	0,41
	3,92
	0,20
	4,77
	-
	-

	PSF i25s5
	0,12
	15,79
	0,08
	16,75
	2,76
	0,75
	4,23
	0,56
	3,90
	0,42
	4,10
	0,21
	4,46
	-
	-

	PSF i30s5
	0,11
	16,47
	0,07
	17,57
	2,90
	0,75
	4,59
	0,56
	4,15
	0,43
	4,39
	0,22
	4,20
	-
	-



Table S6. RC max and RC A50 calculated for the 18F Micro Hollow Sphere phantom with a contrast ratio of 1/8 for C-SC- and OC-SC+ acquisitions.
	 
	C-SC-
	OC-SC+

	 
	RC max
	RC A50
	RC max
	RC A50

	  Φ spheres (mm)
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95

	3D-OSEM i3s5
	0,94
	0,77
	0,55
	0,37
	0,70
	0,58
	0,42
	0,30
	0,91
	0,75
	0,55
	0,32
	0,69
	0,54
	0,41
	0,26

	3D-OSEM i5s5
	0,91
	0,80
	0,61
	0,43
	0,69
	0,60
	0,46
	0,34
	0,90
	0,77
	0,60
	0,35
	0,67
	0,56
	0,44
	0,29

	3D-OSEM i10s5
	0,88
	0,78
	0,63
	0,47
	0,68
	0,59
	0,47
	0,37
	0,88
	0,76
	0,61
	0,37
	0,66
	0,55
	0,44
	0,30

	3D-OSEM i15s5
	0,87
	0,77
	0,63
	0,48
	0,68
	0,59
	0,47
	0,37
	0,87
	0,76
	0,60
	0,38
	0,66
	0,55
	0,44
	0,30

	3D-OSEM i20s5
	0,87
	0,77
	0,63
	0,48
	0,68
	0,58
	0,47
	0,37
	0,87
	0,76
	0,60
	0,38
	0,66
	0,55
	0,44
	0,30

	3D-OSEM i25s5
	0,88
	0,77
	0,63
	0,48
	0,68
	0,58
	0,47
	0,37
	0,87
	0,76
	0,60
	0,37
	0,66
	0,55
	0,44
	0,30

	3D-OSEM i30s5
	0,88
	0,77
	0,63
	0,48
	0,68
	0,58
	0,47
	0,37
	0,87
	0,76
	0,60
	0,37
	0,66
	0,55
	0,44
	0,30

	PSF i3s5
	1,43
	0,89
	0,49
	0,29
	0,99
	0,65
	0,38
	0,23
	1,30
	0,85
	0,48
	0,26
	0,91
	0,60
	0,36
	0,22

	PSF i5s5
	1,72
	1,16
	0,66
	0,36
	1,17
	0,84
	0,49
	0,29
	1,55
	1,14
	0,64
	0,32
	1,08
	0,80
	0,47
	0,26

	PSF i10s5
	1,90
	1,52
	0,94
	0,52
	1,29
	1,08
	0,69
	0,41
	1,72
	1,53
	0,93
	0,41
	1,20
	1,09
	0,67
	0,36

	PSF i15s5
	1,86
	1,68
	1,13
	0,63
	1,28
	1,19
	0,82
	0,50
	1,72
	1,73
	1,13
	0,52
	1,21
	1,21
	0,80
	0,42

	PSF i20s5
	1,79
	1,77
	1,25
	0,72
	1,25
	1,26
	0,90
	0,56
	1,68
	1,84
	1,28
	0,58
	1,19
	1,29
	0,90
	0,47

	PSF i25s5
	1,71
	1,82
	1,35
	0,80
	1,20
	1,30
	0,96
	0,61
	1,63
	1,90
	1,38
	0,63
	1,16
	1,34
	1,00
	0,51

	PSF i30s5
	1,63
	1,85
	1,42
	0,86
	1,17
	1,33
	1,04
	0,66
	1,58
	1,95
	1,47
	0,68
	1,13
	1,37
	1,07
	0,54

	PSF i20s5 Gauss 2mm
	1,43
	1,27
	0,84
	0,51
	1,02
	0,92
	0,62
	0,40
	1,40
	1,22
	0,80
	0,44
	0,98
	0,88
	0,59
	0,36



Table S7. RC max and RC A50 calculated for the 18F Micro Hollow Sphere phantom with a contrast ratio of 1/4 for C-SC- and OC-SC+ acquisitions.
	 
	C-SC-
	OC-SC+

	 
	RC Amax
	RC A50
	RC Amax
	RC A50

	  Φ spheres (mm)
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95

	3D-OSEM i3s5
	0,91
	0,72
	0,50
	0,37
	0,70
	0,57
	0,38
	0,33
	0,89
	0,71
	0,47
	0,32
	0,68
	0,55
	0,39
	0,29

	3D-OSEM i5s5
	0,91
	0,77
	0,56
	0,42
	0,70
	0,61
	0,46
	0,36
	0,90
	0,76
	0,52
	0,34
	0,69
	0,57
	0,43
	0,31

	3D-OSEM i10s5
	0,92
	0,77
	0,61
	0,47
	0,72
	0,61
	0,49
	0,40
	0,88
	0,76
	0,55
	0,37
	0,68
	0,58
	0,44
	0,33

	3D-OSEM i15s5
	0,92
	0,76
	0,62
	0,49
	0,72
	0,60
	0,50
	0,41
	0,88
	0,75
	0,55
	0,38
	0,68
	0,58
	0,45
	0,34

	3D-OSEM i20s5
	0,92
	0,75
	0,62
	0,50
	0,72
	0,60
	0,50
	0,41
	0,88
	0,74
	0,55
	0,38
	0,68
	0,58
	0,45
	0,34

	3D-OSEM i25s5
	0,93
	0,75
	0,62
	0,50
	0,71
	0,60
	0,50
	0,42
	0,89
	0,74
	0,55
	0,38
	0,68
	0,58
	0,45
	0,34

	3D-OSEM i30s5
	0,92
	0,75
	0,62
	0,50
	0,72
	0,60
	0,50
	0,42
	0,89
	0,74
	0,55
	0,38
	0,68
	0,58
	0,45
	0,34

	PSF i3s5
	1,08
	0,70
	0,42
	0,30
	0,83
	0,55
	0,36
	0,28
	1,07
	0,67
	0,41
	0,28
	0,78
	0,52
	0,35
	0,26

	PSF i5s5
	1,32
	0,90
	0,51
	0,34
	1,01
	0,69
	0,42
	0,31
	1,31
	0,86
	0,48
	0,30
	0,93
	0,65
	0,40
	0,28

	PSF i10s5
	1,57
	1,22
	0,67
	0,42
	1,18
	0,91
	0,53
	0,36
	1,58
	1,18
	0,63
	0,35
	1,11
	0,87
	0,50
	0,32

	PSF i15s5
	1,64
	1,40
	0,78
	0,49
	1,24
	1,02
	0,59
	0,41
	1,67
	1,02
	0,59
	0,41
	1,16
	0,99
	0,57
	0,34

	PSF i20s5
	1,65
	1,52
	0,86
	0,54
	1,24
	1,11
	0,67
	0,45
	1,69
	1,51
	0,82
	0,42
	1,18
	1,08
	0,63
	0,36

	PSF i25s5
	1,65
	1,60
	0,92
	0,58
	1,25
	1,16
	0,71
	0,48
	1,68
	1,60
	0,89
	0,45
	1,18
	1,16
	0,68
	0,38

	PSF i30s5
	1,64
	1,65
	0,96
	0,62
	1,23
	1,20
	0,74
	0,51
	1,66
	1,68
	0,94
	0,47
	1,17
	1,22
	0,71
	0,40

	PSF i20s5 Gauss 2mm
	1,34
	1,09
	0,65
	0,42
	1,01
	0,83
	0,53
	0,38
	1,32
	1,06
	0,62
	0,37
	0,96
	0,80
	0,50
	0,33



Table S8. RC max and RC A50 calculated for the 68Ga Micro Hollow Sphere phantom with a contrast ratio of 1/8 for C-SC- and OC-SC+ acquisitions
	 
	C-SC-
	OC-SC+

	 
	RC max
	RC A50
	RC max
	RC A50

	  Φ spheres (mm)
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95
	7.86
	6.23
	4.95
	3.95

	3D-OSEM i3s5
	0,63
	0,41
	0,29
	0,19
	0,46
	0,31
	0,22
	0,16
	0,59
	0,40
	0,27
	0,17
	0,43
	0,30
	0,21
	0,15

	3D-OSEM i5s5
	0,63
	0,42
	0,31
	0,21
	0,45
	0,31
	0,24
	0,17
	0,59
	0,41
	0,28
	0,18
	0,43
	0,31
	0,22
	0,16

	3D-OSEM i10s5
	0,62
	0,42
	0,32
	0,23
	0,46
	0,32
	0,25
	0,18
	0,59
	0,40
	0,29
	0,19
	0,43
	0,31
	0,23
	0,16

	3D-OSEM i15s5
	0,62
	0,41
	0,32
	0,23
	0,46
	0,32
	0,25
	0,19
	0,59
	0,40
	0,29
	0,20
	0,43
	0,31
	0,23
	0,17

	3D-OSEM i20s5
	0,62
	0,41
	0,32
	0,24
	0,46
	0,32
	0,25
	0,19
	0,59
	0,39
	0,30
	0,20
	0,43
	0,31
	0,23
	0,17

	3D-OSEM i25s5
	0,62
	0,41
	0,32
	0,24
	0,46
	0,32
	0,25
	0,19
	0,59
	0,39
	0,30
	0,20
	0,43
	0,31
	0,23
	0,17

	3D-OSEM i30s5
	0,62
	0,41
	0,32
	0,24
	0,46
	0,32
	0,25
	0,19
	0,59
	0,39
	0,30
	0,20
	0,43
	0,31
	0,23
	0,17

	PSF i3s5
	0,82
	0,48
	0,28
	0,17
	0,58
	0,35
	0,22
	0,15
	0,75
	0,45
	0,27
	0,16
	0,54
	0,33
	0,21
	0,14

	PSF i5s5
	0,92
	0,57
	0,34
	0,20
	0,64
	0,41
	0,26
	0,16
	0,84
	0,55
	0,31
	0,18
	0,60
	0,39
	0,24
	0,15

	PSF i10s5
	0,96
	0,68
	0,43
	0,24
	0,67
	0,47
	0,31
	0,19
	0,89
	0,65
	0,38
	0,21
	0,63
	0,45
	0,28
	0,18

	PSF i15s5
	0,94
	0,71
	0,47
	0,26
	0,67
	0,50
	0,35
	0,21
	0,87
	0,69
	0,42
	0,23
	0,63
	0,48
	0,31
	0,19

	PSF i20s5
	0,92
	0,73
	0,50
	0,28
	0,66
	0,51
	0,37
	0,22
	0,86
	0,70
	0,44
	0,25
	0,62
	0,49
	0,32
	0,20

	PSF i25s5
	0,90
	0,74
	0,53
	0,29
	0,65
	0,51
	0,38
	0,23
	0,85
	0,71
	0,45
	0,26
	0,61
	0,50
	0,33
	0,21

	PSF i30s5
	0,88
	0,74
	0,54
	0,30
	0,64
	0,52
	0,39
	0,24
	0,84
	0,72
	0,46
	0,27
	0,60
	0,50
	0,34
	0,21




