# Supplementary materials

Title: Bone marrow dosimetry in low volume mHSPC patients receiving Lu-177-PSMA therapy using SPECT/CT

Authors: Dagmar Grob1, Bastiaan M. Privé1,2, Constantijn H.J. Muselaers4, Niven Mehra5, James Nagarajah1, Mark W. Konijnenberg1,3, Steffie M.B. Peters1

1: Department of Medical Imaging, Radboud university medical center, Nijmegen, The Netherlands

2: Department of Radiation Oncology, Erasmus Medical Center, Rotterdam, The Netherlands

3: Department of Radiology and Nuclear Medicine, Erasmus Medical Center, Rotterdam, The Netherlands

4: Department of Urology, Radboud university medical center, Nijmegen, The Netherlands

5: Department of Medical Oncology, Radboud university medical center, Nijmegen, The Netherlands

Corresponding author:

Steffie M.B. Peters (ORCID: 0000-0002-0752-7134)

Radboud university medical center

Department of Medical Imaging

P.O. Box 9101

6500 HB Nijmegen, The Netherlands

T: +31 (6) 11621752

E: steffie.peters@radboudumc.nl

Online Resource 1: Study flowchart

Online Resource 2: Acquisition and reconstruction parameters of the imaging protocols

SPECT/CT and planar imaging was performed at 1, 24, 48, 72 and 168 hours after administration of [177Lu]Lu-PSMA on either a Symbia T16 or Symbia Intevo Bold system (Siemens Healthineers, Erlangen, Germany). SPECT/CT scans were acquired at three body regions: the pelvis, abdomen, and head-neck region (64 projections per detector, time per projection of 14 s, a 20% photon energy window at 208 keV, and dual-energy window for Compton scattering. SPECT data were reconstructed using OSEM reconstruction (Flash 3D with collimator detector response) using 4 iterations and 8 subsets, matrix size of 128 (resulting in a cuboid-shape voxel size of 4.8 mm3) and a smoothing Gaussian filter of 8.4 mm. The estimated reconstructed SPECT spatial resolution was 15 mm (FWHM). A low dose CT was performed (average DLP of 130 mGy·cm ) and the data were reconstructed using B31s kernel and 3.0 mm slice thickness resulting in a voxel size of 1.0 x 1.0 x 3.0 mm3. Of note, the SPECT image reconstruction approach takes into account corrections for scatter, CT-based attenuation, and dead-time.

Online Resource 3: Uncertainty analysis flowchart

- Error in SPECT camera specific calibration factor for 177Lu: 5% [Peters, 2020]

CF = 10.6 ± 0.5 cps/MBq low count statistics at late time points was accounted for by Poisson distribution.

- Data collection:

  - Quantitative SPECT at 5 time-points, within 45 minute scan-time

  - Activity concentration in blood samples at 9 time points; error: 5%

- Drawing of VOIs in SPECT data over organs and tumor lesions to determine counts:

  - Fixed size spherical VOI over central region of vertebrae; error: 10%

  - CT based VOI over organs; error 5%

- Determination of lesion and organ volume:

  - Lesion volume based on PET/CT and diagnostic CT [Jentzen, 2015];

Voxelisation and resolution error volume with lesion diameter d and voxel size a: 	δVV2=9 δdd2=1.5ad2 [Gear, 2018]

  - Organ volumes set at fixed ICRP-89 male phantom values; error 10%

  - Bone marrow volume relative to red marrow content in vertebrae and total volume, systematic uncertainty in patients due to age, health status and prior therapy; e.g. 16% difference between S-values by Olinda (with 1170 g0 and IDAC-dose (1394 g), hence a 20% uncertainty in the S-value.

- Fit to Time-Activity Curve

  - Mono-exponential fit SPECT TAC when R2 > 0.7, determine covariance matrix

  - Actual SPECT when R2 < 0.7 with exponential decay from last time-point

  - Three-exponential curve fit of blood data, determine covariance matrix

- Integration over time of the TAC, to determine the time-integrated activity Ã

  - Mono-exponential integration, error in TIA or Ã:

δAA2=δA0A02+δλeffλeff2-2 CovA0,λeffA0λeff

  - Trapezoid integration method; error:

δAA2=iδA(ti)A(ti)2

  - Three-exponential integration of blood concentration TAC, error in TIA or Ã:

δAA2=δA0A02+δPP2+δ%Fast%Fast2+δλFastλFast+λp2+δ%Slow%Slow2+δλSlowλSlow+λp+δλMediumλMedium+λp-covariances

- Absorbed dose calculation with MIRD equation D=A×S; error:

δDD2=δAA2+δSS2

Online Resource 4: Blood measurements for blood and bone marrow toxicity. Each coloured line represents a patient. Dotted vertical lines indicate the first and second therapeutic cycle with 177Lu-PSMA-617. (a): hemoglobin, (b): white blood cells and (c): thrombocytes