Supplemental Materials of Harmonisation of standard uptake values on long axial field of view PET/CT: A exploratory study

Yue Lia,☆, Zhu Zhanga,☆, Hanyu Baia, Xinmiao Lua,*, Hua Panga,*

a: Department of Nuclear Medicine, The First Affiliated Hospital of Chongqing Medical University, Chongqing 400016, China

☆:Yue Li and Zhu Zhang made equal contributions to this work and were considered as the cofirst authors.

*Corresponding author: Xinmiao Lu, Department of Nuclear Medicine, The First Affiliated Hospital of Chongqing Medical University, Chongqing 400042, P. R. China. Email: luxinmiao@sjtu.edu.cn. ORCID:0009-0007-1968-1976

*Corresponding author: Hua Pang, Department of Nuclear Medicine, The First Affiliated Hospital of Chongqing Medical University, Chongqing 400042, P. R. China. Email: ph1973@126.com. ORCID:0000-0001-6601-3208

Table S1. The range of patient body positions.

| No. | Gender | Head (cm) | Chest(cm) | Abdomen(cm) | Pelvis(cm) | Leg(cm) |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Male | 24.9 | 19.78 | 26.61 | 10.15 | 88.53 |
| 2 | Male | 25.98 | 16.62 | 24.63 | 12.92 | 20.29 |
| 3 | Male | 26.32 | 23.08 | 25.11 | 12.42 | 83.05 |
| 4 | Male | 26.73 | 18.47 | 18.23 | 13.66 | 66.58 |
| 5 | Male | 26.97 | 21.65 | 22.12 | 12.61 | 42.64 |
| 6 | Male | 26.8 | 19.84 | 23.86 | 12.98 | 24.1 |
| 7 | Male | 27.18 | 23.31 | 25.33 | 9.68 | 17.34 |
| 8 | Male | 27.53 | 23.23 | 22.39 | 12.31 | 21.53 |
| 9 | Male | 28.23 | 17.49 | 26.46 | 12.61 | 42.27 |
| 10 | Male | 27.81 | 18.6 | 24.79 | 12.99 | 16.6 |
| 11 | Male | 28 | 22.71 | 22.28 | 13.06 | 18.3 |
| 12 | Male | 26.94 | 25.93 | 18.76 | 10.01 | 84.95 |
| 13 | Male | 28.07 | 23.27 | 22.88 | 13.96 | 46.58 |
| 14 | Male | 27.48 | 19.14 | 22.53 | 12.23 | 66.23 |
| 15 | Male | 28.19 | 20.72 | 29.01 | 14.36 | 53.97 |
| 16 | Male | 29.35 | 22.73 | 21.53 | 12.86 | 41.73 |
| 17 | Male | 26.89 | 25.22 | 20.58 | 13.3 | 60.33 |
| 18 | Male | 29.11 | 24.18 | 21.41 | 11.78 | 39.92 |
| 19 | Male | 28.32 | 24.31 | 22.68 | 13.1 | 27.02 |
| 20 | Male | 26.78 | 24.88 | 22.1 | 12.31 | 21.53 |
| 21 | Male | 27.56 | 16.97 | 28.35 | 12.59 | 87.18 |
| 22 | Male | 28.64 | 22.23 | 25.4 | 13.35 | 59.54 |
| 23 | Male | 28.11 | 19.86 | 27.08 | 12.12 | 41.65 |
| 24 | Male | 25.99 | 16.05 | 23.49 | 13.41 | 79.73 |
| 25 | Female | 24.59 | 18.35 | 22.54 | 13.27 | 9.71 |
| 26 | Female | 25.12 | 16.86 | 24.4 | 13.1 | 35.84 |
| 27 | Female | 25.1 | 18.29 | 23.72 | 11.17 | 37.19 |
| 28 | Female | 25.29 | 20.83 | 22.62 | 12.9 | 19.05 |
| 29 | Female | 26.79 | 17.64 | 24.94 | 14.29 | 18.24 |
| 30 | Female | 24.93 | 15.8 | 26.58 | 12.21 | 74.88 |
| 31 | Female | 24.33 | 19.54 | 23.6 | 12.59 | 82.19 |
| 32 | Female | 26.08 | 19.23 | 25.94 | 11.77 | 61.35 |
| 33 | Female | 26.54 | 16.29 | 27.12 | 11.92 | 40.99 |
| 34 | Female | 25.02 | 17.51 | 25.32 | 12.36 | 50.98 |
| 35 | Female | 24.87 | 17.79 | 24.41 | 13.56 | 57.6 |
| 36 | Female | 26.4 | 21.29 | 23.53 | 13.41 | 35.84 |
| 37 | Female | 26.23 | 15.98 | 25.2 | 13.01 | 39.99 |
| 38 | Female | 25.99 | 17.29 | 23.47 | 13.8 | 18.83 |
| 39 | Female | 27.04 | 17.11 | 23.73 | 13.71 | 13.52 |
| 40 | Female | 26.62 | 18.43 | 23.93 | 13.58 | 15.95 |
| 41 | Female | 24.81 | 16.8 | 23.28 | 12.86 | 17.82 |
| 42 | Female | 25.13 | 16.17 | 23.97 | 12.51 | 19.74 |
| 43 | Female | 25.68 | 20.56 | 22.17 | 13.43 | 56.29 |
| 44 | Female | 25.11 | 19.41 | 22.26 | 13.89 | 17.6 |
| 45 | Female | 27.19 | 18.67 | 27.55 | 12.39 | 85.42 |

Table S2. Optimal Gaussian FWHM for SUVmax harmonisation at different positions and reconstruction parameters.

|  | Reconstruction parameters | Optimal Filter FWHM (mm) of RCmax under each position |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | Head Position (0-30 cm) | Chest position (30-60 cm) | Abdomen position (60-90 cm) |  | Pelvis position (90-120 cm) |  | Leg position (120-148 cm) |  |
| EARL 1 | 2i10s G2.0 | 7.6 | 7.5 |  | 7.5 |  | 7.5 |  | 7.5 |
|  | 2i20s | 7.9 | 7.9 |  | 8.0 |  | 7.9 |  | 7.9 |
|  | 3i10s G2.0 | 7.8 | 7.7 |  | 7.7 |  | 7.7 |  | 7.7 |
|  | 3i10s | 7.9 | 7.9 |  | 7.9 |  | 7.9 |  | 7.8 |
|  | 3i20s G2.0 | 7.8 | 7.8 |  | 7.9 |  | 7.8 |  | 7.7 |
|  | 3i20s | 8.0 | 8.0 |  | 8.0 |  | 8.0 |  | 7.8 |
| EARL 2 | 2i10s G2.0 | 5.1 | 4.7 |  | 4.6 |  | 4.5 |  | 4.3 |
|  | 2i20s | 5.5 | 5.6 |  | 5.0 |  | 5.4 |  | 4.5 |
|  | 3i10s G2.0 | 5.4 | 5.3 |  | 4.6 |  | 5.1 |  | 4.6 |
|  | 3i10s | 5.6 | 5.4 |  | 4.9 |  | 5.0 |  | 4.7 |
|  | 3i20s G2.0 | 5.5 | 5.4 |  | 4.7 |  | 5.4 |  | 4.0 |
|  | 3i20s | 5.7 | 5.6 |  | 5.1 |  | 5.6 |  | 4.2 |

Table S3. RCmax for each hot sphere after SUVmax harmonisation at different positions and reconstruction parameters with position-specific optimal Gaussian filter FWHM.

| Position | Reconstruction parameters | Sphere diameter (mm) | RCmax | RCmax-unity (EARL 1) | RCmax-unity (EARL 2) |
| --- | --- | --- | --- | --- | --- |
| Head position (0-30cm) | 2i10s G2.0 | 10 | 1.34 | 0.46 | 0.72 |
|  |  | 13 | 1.25 | 0.68 | 0.95 |
|  |  | 17 | 1.22 | 0.93 | 1.08 |
|  |  | 22 | 1.23 | 0.99 | 1.05 |
|  |  | 28 | 1.17 | 0.98 | 1.02 |
|  |  | 37 | 1.15 | 1.00 | 1.04 |
|  | 2i20s | 10 | 1.47 | 0.46 | 0.73 |
|  |  | 13 | 1.20 | 0.68 | 0.93 |
|  |  | 17 | 1.31 | 0.90 | 1.04 |
|  |  | 22 | 1.44 | 0.97 | 1.02 |
|  |  | 28 | 1.21 | 0.97 | 1.00 |
|  |  | 37 | 1.28 | 1.00 | 1.04 |
|  | 3i10s G2.0 | 10 | 1.41 | 0.45 | 0.71 |
|  |  | 13 | 1.19 | 0.67 | 0.92 |
|  |  | 17 | 1.23 | 0.90 | 1.04 |
|  |  | 22 | 1.29 | 0.97 | 1.02 |
|  |  | 28 | 1.19 | 0.96 | 0.99 |
|  |  | 37 | 1.20 | 1.00 | 1.04 |
|  | 3i10s | 10 | 1.50 | 0.46 | 0.71 |
|  |  | 13 | 1.22 | 0.68 | 0.92 |
|  |  | 17 | 1.27 | 0.90 | 1.04 |
|  |  | 22 | 1.37 | 0.97 | 1.02 |
|  |  | 28 | 1.23 | 0.96 | 0.99 |
|  |  | 37 | 1.24 | 1.00 | 1.04 |
|  | 3i20s G2.0 | 10 | 1.34 | 0.46 | 0.71 |
|  |  | 13 | 1.16 | 0.67 | 0.90 |
|  |  | 17 | 1.26 | 0.89 | 1.01 |
|  |  | 22 | 1.42 | 0.96 | 1.01 |
|  |  | 28 | 1.20 | 0.97 | 1.00 |
|  |  | 37 | 1.27 | 1.00 | 1.03 |
|  | 3i20s | 10 | 1.41 | 0.45 | 0.71 |
|  |  | 13 | 1.20 | 0.66 | 0.90 |
|  |  | 17 | 1.33 | 0.88 | 1.01 |
|  |  | 22 | 1.52 | 0.96 | 1.01 |
|  |  | 28 | 1.25 | 0.97 | 1.00 |
|  |  | 37 | 1.33 | 1.00 | 1.03 |
| Chest position (30-60cm) | 2i10s G2.0 | 10 | 1.31 | 0.46 | 0.75 |
|  |  | 13 | 1.31 | 0.69 | 1.00 |
|  |  | 17 | 1.20 | 0.93 | 1.11 |
|  |  | 22 | 1.18 | 1.00 | 1.05 |
|  |  | 28 | 1.15 | 1.00 | 1.04 |
|  |  | 37 | 1.18 | 1.01 | 1.05 |
|  | 2i20s | 10 | 1.53 | 0.46 | 0.71 |
|  |  | 13 | 1.22 | 0.68 | 0.93 |
|  |  | 17 | 1.25 | 0.89 | 1.01 |
|  |  | 22 | 1.20 | 0.96 | 0.99 |
|  |  | 28 | 1.25 | 0.98 | 1.02 |
|  |  | 37 | 1.24 | 1.01 | 1.03 |
|  | 3i10s G2.0 | 10 | 1.40 | 0.46 | 0.71 |
|  |  | 13 | 1.25 | 0.68 | 0.94 |
|  |  | 17 | 1.21 | 0.90 | 1.04 |
|  |  | 22 | 1.18 | 0.97 | 1.01 |
|  |  | 28 | 1.20 | 0.99 | 1.02 |
|  |  | 37 | 1.22 | 1.00 | 1.03 |
|  | 3i10s | 10 | 1.51 | 0.45 | 0.72 |
|  |  | 13 | 1.27 | 0.67 | 0.95 |
|  |  | 17 | 1.24 | 0.90 | 1.04 |
|  |  | 22 | 1.21 | 0.97 | 1.01 |
|  |  | 28 | 1.23 | 0.98 | 1.03 |
|  |  | 37 | 1.24 | 1.00 | 1.04 |
|  | 3i20s G2.0 | 10 | 1.39 | 0.46 | 0.71 |
|  |  | 13 | 1.17 | 0.68 | 0.92 |
|  |  | 17 | 1.22 | 0.88 | 0.98 |
|  |  | 22 | 1.18 | 0.95 | 0.99 |
|  |  | 28 | 1.25 | 0.98 | 1.01 |
|  |  | 37 | 1.24 | 1.01 | 1.04 |
|  | 3i20s | 10 | 1.49 | 0.45 | 0.71 |
|  |  | 13 | 1.21 | 0.66 | 0.92 |
|  |  | 17 | 1.25 | 0.87 | 0.98 |
|  |  | 22 | 1.22 | 0.95 | 0.99 |
|  |  | 28 | 1.28 | 0.98 | 1.01 |
|  |  | 37 | 1.27 | 1.01 | 1.04 |
| Abdomen position (60-90cm) | 2i10s G2.0 | 10 | 1.22 | 0.46 | 0.75 |
|  |  | 13 | 1.34 | 0.70 | 1.03 |
|  |  | 17 | 1.22 | 0.94 | 1.12 |
|  |  | 22 | 1.14 | 1.03 | 1.07 |
|  |  | 28 | 1.15 | 1.00 | 1.05 |
|  |  | 37 | 1.13 | 1.02 | 1.06 |
|  | 2i20s | 10 | 1.41 | 0.46 | 0.80 |
|  |  | 13 | 1.37 | 0.68 | 1.02 |
|  |  | 17 | 1.21 | 0.90 | 1.04 |
|  |  | 22 | 1.16 | 0.97 | 1.02 |
|  |  | 28 | 1.23 | 0.98 | 1.06 |
|  |  | 37 | 1.23 | 1.00 | 1.06 |
|  | 3i10s G2.0 | 10 | 1.30 | 0.46 | 0.80 |
|  |  | 13 | 1.29 | 0.69 | 1.04 |
|  |  | 17 | 1.19 | 0.92 | 1.08 |
|  |  | 22 | 1.13 | 1.00 | 1.04 |
|  |  | 28 | 1.19 | 0.98 | 1.06 |
|  |  | 37 | 1.14 | 1.01 | 1.06 |
|  | 3i10s | 10 | 1.38 | 0.45 | 0.78 |
|  |  | 13 | 1.34 | 0.68 | 1.03 |
|  |  | 17 | 1.22 | 0.91 | 1.07 |
|  |  | 22 | 1.14 | 0.99 | 1.04 |
|  |  | 28 | 1.21 | 0.98 | 1.06 |
|  |  | 37 | 1.18 | 1.01 | 1.06 |
|  | 3i20s G2.0 | 10 | 1.29 | 0.46 | 0.82 |
|  |  | 13 | 1.31 | 0.68 | 1.01 |
|  |  | 17 | 1.16 | 0.88 | 1.02 |
|  |  | 22 | 1.17 | 0.96 | 1.02 |
|  |  | 28 | 1.22 | 0.97 | 1.06 |
|  |  | 37 | 1.22 | 1.00 | 1.06 |
|  | 3i20s | 10 | 1.35 | 0.45 | 0.79 |
|  |  | 13 | 1.36 | 0.67 | 0.99 |
|  |  | 17 | 1.20 | 0.88 | 1.01 |
|  |  | 22 | 1.20 | 0.96 | 1.02 |
|  |  | 28 | 1.25 | 0.97 | 1.05 |
|  |  | 37 | 1.28 | 1.00 | 1.05 |
| Pelvic position (90-120cm) | 2i10s G2.0 | 10 | 1.25 | 0.46 | 0.77 |
|  |  | 13 | 1.31 | 0.69 | 1.02 |
|  |  | 17 | 1.19 | 0.94 | 1.12 |
|  |  | 22 | 1.19 | 0.99 | 1.06 |
|  |  | 28 | 1.13 | 0.99 | 1.05 |
|  |  | 37 | 1.18 | 1.02 | 1.08 |
|  | 2i20s | 10 | 1.41 | 0.46 | 0.73 |
|  |  | 13 | 1.29 | 0.68 | 0.94 |
|  |  | 17 | 1.20 | 0.91 | 1.03 |
|  |  | 22 | 1.23 | 0.95 | 0.99 |
|  |  | 28 | 1.16 | 0.97 | 1.01 |
|  |  | 37 | 1.24 | 1.00 | 1.06 |
|  | 3i10s G2.0 | 10 | 1.33 | 0.46 | 0.73 |
|  |  | 13 | 1.26 | 0.68 | 0.95 |
|  |  | 17 | 1.18 | 0.91 | 1.05 |
|  |  | 22 | 1.20 | 0.97 | 1.01 |
|  |  | 28 | 1.15 | 0.98 | 1.02 |
|  |  | 37 | 1.19 | 1.01 | 1.06 |
|  | 3i10s | 10 | 1.42 | 0.45 | 0.77 |
|  |  | 13 | 1.30 | 0.67 | 0.99 |
|  |  | 17 | 1.21 | 0.91 | 1.06 |
|  |  | 22 | 1.24 | 0.96 | 1.02 |
|  |  | 28 | 1.18 | 0.98 | 1.03 |
|  |  | 37 | 1.21 | 1.01 | 1.07 |
|  | 3i20s G2.0 | 10 | 1.30 | 0.46 | 0.71 |
|  |  | 13 | 1.26 | 0.67 | 0.91 |
|  |  | 17 | 1.17 | 0.89 | 1.00 |
|  |  | 22 | 1.21 | 0.94 | 0.97 |
|  |  | 28 | 1.16 | 0.97 | 1.00 |
|  |  | 37 | 1.24 | 1.00 | 1.05 |
|  | 3i20s | 10 | 1.37 | 0.45 | 0.71 |
|  |  | 13 | 1.34 | 0.66 | 0.91 |
|  |  | 17 | 1.22 | 0.88 | 1.00 |
|  |  | 22 | 1.26 | 0.94 | 0.97 |
|  |  | 28 | 1.20 | 0.97 | 1.00 |
|  |  | 37 | 1.28 | 1.00 | 1.05 |
| Leg position (120-149cm) | 2i10s G2.0 | 10 | 1.18 | 0.46 | 0.79 |
|  |  | 13 | 1.27 | 0.69 | 1.03 |
|  |  | 17 | 1.19 | 0.90 | 1.06 |
|  |  | 22 | 1.19 | 0.98 | 1.05 |
|  |  | 28 | 1.23 | 0.98 | 1.04 |
|  |  | 37 | 1.18 | 1.01 | 1.08 |
|  | 2i20s | 10 | 1.35 | 0.46 | 0.87 |
|  |  | 13 | 1.29 | 0.68 | 1.04 |
|  |  | 17 | 1.19 | 0.87 | 1.00 |
|  |  | 22 | 1.28 | 0.98 | 1.04 |
|  |  | 28 | 1.34 | 0.95 | 1.04 |
|  |  | 37 | 1.28 | 1.00 | 1.09 |
|  | 3i10s G2.0 | 10 | 1.23 | 0.45 | 0.78 |
|  |  | 13 | 1.23 | 0.68 | 1.00 |
|  |  | 17 | 1.17 | 0.88 | 1.02 |
|  |  | 22 | 1.19 | 0.96 | 1.02 |
|  |  | 28 | 1.25 | 0.97 | 1.02 |
|  |  | 37 | 1.22 | 1.00 | 1.07 |
|  | 3i10s | 10 | 1.29 | 0.46 | 0.80 |
|  |  | 13 | 1.29 | 0.68 | 1.01 |
|  |  | 17 | 1.19 | 0.88 | 1.02 |
|  |  | 22 | 1.24 | 0.96 | 1.02 |
|  |  | 28 | 1.29 | 0.97 | 1.03 |
|  |  | 37 | 1.24 | 1.00 | 1.07 |
|  | 3i20s G2.0 | 10 | 1.25 | 0.47 | 0.91 |
|  |  | 13 | 1.22 | 0.69 | 1.03 |
|  |  | 17 | 1.15 | 0.87 | 1.00 |
|  |  | 22 | 1.28 | 0.97 | 1.05 |
|  |  | 28 | 1.38 | 0.95 | 1.08 |
|  |  | 37 | 1.27 | 1.00 | 1.11 |
|  | 3i20s | 10 | 1.30 | 0.47 | 0.92 |
|  |  | 13 | 1.28 | 0.68 | 1.04 |
|  |  | 17 | 1.19 | 0.86 | 1.01 |
|  |  | 22 | 1.34 | 0.97 | 1.05 |
|  |  | 28 | 1.47 | 0.95 | 1.08 |
|  |  | 37 | 1.35 | 1.00 | 1.11 |

Table S4. Calculated RMSE at each scan position before and after harmonisation of SUVmax with position-specific optimal Gaussian filter

| Position | Reconstruction parameters | RMSE before EARL 1 harmonization | RMSE after EARL 1 harmonization | RMSE before EARL 2 harmonization | RMSE after EARL 2 harmonization |
| --- | --- | --- | --- | --- | --- |
| Head position + F-head | 2i10s G2.0 | 0.522 | 0.071 | 0.274 | 0.100 |
|  | 2i20s | 0.604 | 0.055 | 0.346 | 0.124 |
|  | 3i10s G2.0 | 0.545 | 0.054 | 0.299 | 0.127 |
|  | 3i10s | 0.597 | 0.055 | 0.346 | 0.124 |
|  | 3i20s G2.0 | 0.547 | 0.051 | 0.288 | 0.139 |
|  | 3i20s | 0.609 | 0.049 | 0.342 | 0.139 |
|  |  | 0.518 | 0.072 | 0.271 | 0.087 |
| Chest position + F-chest | 2i10s G2.0 | 0.588 | 0.054 | 0.350 | 0.130 |
|  | 2i20s | 0.539 | 0.057 | 0.299 | 0.118 |
|  | 3i10s G2.0 | 0.587 | 0.054 | 0.345 | 0.117 |
|  | 3i10s | 0.529 | 0.050 | 0.291 | 0.143 |
|  | 3i20s G2.0 | 0.580 | 0.047 | 0.337 | 0.140 |
|  | 3i20s | 0.493 | 0.081 | 0.246 | 0.077 |
|  |  | 0.567 | 0.052 | 0.322 | 0.109 |
| Abdomen position + F-abdoment | 2i10s G2.0 | 0.504 | 0.065 | 0.266 | 0.097 |
|  | 2i20s | 0.547 | 0.062 | 0.305 | 0.096 |
|  | 3i10s G2.0 | 0.513 | 0.048 | 0.267 | 0.117 |
|  | 3i10s | 0.556 | 0.050 | 0.302 | 0.117 |
|  | 3i20s G2.0 | 0.489 | 0.065 | 0.251 | 0.079 |
|  | 3i20s | 0.544 | 0.045 | 0.309 | 0.126 |
|  |  | 0.505 | 0.051 | 0.273 | 0.109 |
| Pelvic position + F-pelivis | 2i10s G2.0 | 0.549 | 0.047 | 0.313 | 0.105 |
|  | 2i20s | 0.498 | 0.041 | 0.262 | 0.143 |
|  | 3i10s G2.0 | 0.551 | 0.038 | 0.305 | 0.140 |
|  | 3i10s | 0.446 | 0.041 | 0.220 | 0.095 |
|  | 3i20s G2.0 | 0.525 | 0.040 | 0.301 | 0.124 |
|  | 3i20s | 0.455 | 0.033 | 0.235 | 0.116 |
|  |  | 0.498 | 0.034 | 0.273 | 0.115 |
| Leg position + F-leg | 2i10s G2.0 | 0.484 | 0.043 | 0.263 | 0.128 |
|  | 2i20s | 0.539 | 0.044 | 0.312 | 0.128 |
|  | 3i10s G2.0 | 0.522 | 0.071 | 0.274 | 0.100 |
|  | 3i10s | 0.604 | 0.055 | 0.346 | 0.124 |
|  | 3i20s G2.0 | 0.545 | 0.054 | 0.299 | 0.127 |
|  | 3i20s | 0.597 | 0.055 | 0.346 | 0.124 |

Table S5. Example of calculated RMSE at each scan position after SUVmax harmonisation with optimal Gaussian post-filter determined from another scan position

| Harmonization standards | Filter Position (optimal filter) | Scan Position (optimal filter) | Reconstruction parameters | Sphere diameter (mm) | RCmax | RMSE after harmonization with ummatched optimal filter | RMSE after harmonization with matched optimal filter |
| --- | --- | --- | --- | --- | --- | --- | --- |
| EARL 1 | F-leg | abdomen | 2i20s | 10mm | 0.46 | 0.058 | 0.052 |
|  | (7.9mm) | (8.0mm) |  | 13mm | 0.71 |  |  |
|  |  |  |  | 17mm | 0.91 |  |  |
|  |  |  |  | 22mm | 0.89 |  |  |
|  |  |  |  | 28mm | 0.92 |  |  |
|  |  |  |  | 37mm | 0.94 |  |  |
|  | F-leg | abdomen | 3i10s | 10mm | 0.50 | 0.067 | 0.062 |
|  | (7.8mm) | (7.9mm) |  | 13mm | 0.75 |  |  |
|  |  |  |  | 17mm | 0.96 |  |  |
|  |  |  |  | 22mm | 0.97 |  |  |
|  |  |  |  | 28mm | 1.02 |  |  |
|  |  |  |  | 37mm | 1.03 |  |  |
|  | F-leg | abdomen | 3i20s | 10mm | 0.50 | 0.061 | 0.050 |
|  | (7.8mm) | (8.0mm) |  | 13mm | 0.75 |  |  |
|  |  |  |  | 17mm | 0.95 |  |  |
|  |  |  |  | 22mm | 0.97 |  |  |
|  |  |  |  | 28mm | 1.02 |  |  |
|  |  |  |  | 37mm | 1.03 |  |  |
| EARL 2 | F-abdomen | leg | 2i20s | 10mm | 0.52 | 0.125 | 0.124 |
|  | (5.0mm) | (4.5mm) |  | 13mm | 0.79 |  |  |
|  |  |  |  | 17mm | 0.92 |  |  |
|  |  |  |  | 22mm | 0.92 |  |  |
|  |  |  |  | 28mm | 0.97 |  |  |
|  |  |  |  | 37mm | 1.00 |  |  |
|  | F-abdomen | leg | 3i10s | 10mm | 0.37 | 0.117 | 0.115 |
|  | (4.9mm) | (4.7mm) |  | 13mm | 0.54 |  |  |
|  |  |  |  | 17mm | 0.67 |  |  |
|  |  |  |  | 22mm | 0.69 |  |  |
|  |  |  |  | 28mm | 0.72 |  |  |
|  |  |  |  | 37mm | 0.75 |  |  |
|  | F-abdomen | leg | 3i20s | 10mm | 0.49 | 0.132 | 0.128 |
|  | (5.1mm) | (4.2mm) |  | 13mm | 0.73 |  |  |
|  |  |  |  | 17mm | 0.92 |  |  |
|  |  |  |  | 22mm | 0.90 |  |  |
|  |  |  |  | 28mm | 0.97 |  |  |
|  |  |  |  | 37mm | 0.99 |  |  |