Supplementary Table S1:  Reproduced from Perrey, S., Quaresima, V. & Ferrari, M. "Muscle Oximetry in Sports Science: An Updated Systematic Review." Sports Med 54, 975–996 (2024).

Commercial muscle NIRS oximeters utilized in the sports science articles included in Table S1

| Device | Manufacturer | Technique | Light source | No. of wl (SD, mm) | Time- resolution (Hz) | No. of channels | Measurable parameters | Website |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Oxymon MkIII | Artinis Medical Systems, Elst, The Netherlands | Multi-distance CW | Laser diodes | 2 (50) | 250 | 1 | ΔtHb/ΔO2Hb/ΔHHb | http://www.artinis.com |
| PortaLitea,b,c | Artinis Medical Systems, Elst, The Netherlands | Multi-distance CW | LEDs | 2 (40) | 50 | 1 | TSI, ΔtHb/ΔO2Hb/ΔHHb | http://www.artinis.com |
| PortaMona,b,d,e | Artinis Medical Systems, Elst, The Netherlands | Multi-distance CW | LEDs | 2 (40) | 10 | 1 | TSI, ΔtHb/ΔO2HbΔHHb | http://www.artinis.com |
| Hb11, Hb14f | Astem, Kyoto, Japan | Multi-distance CW | LEDs | 2 (3) | 1 | 1 | StO2/ΔtHb/ΔO2Hb/ΔHHb | http://www.astem-jp.com |
| Humon Hexa,b,g,h | Dynometric, Inc, Cambridge, USA | Multi-distance CW | LEDs | 2 (2.4) | 4 | 1 | SmO2 | http://www.humon.io |
| Moxya,b,i,j,k | Fortiori Design LLC, Hutchinson, MN, USA | Monte Carlo modeling/ CW | LEDs | 4 (25) | 2 | 1 | SmO2, autHb | http://www.moxymonitor.com |
| NIRO-200l | Hamamatsu Photonics K.K., Hamamatsu City, Japan | Multi-distance CW | Laser diodes | 3 (40) | 20 | 2 | TOI, ΔtHb ΔO2Hb/ΔHHb | http://www.hamamatsu.com |
| NIRO-200NXm | Hamamatsu Photonics K.K., Hamamatsu City, Japan | Multi-distance CW | Laser diodes | 3 (40) | 20 | 2 | TOI, ΔtHb/ΔO2Hb/ΔHHb | http://www.hamamatsu.com |
| TRS-20g | Hamamatsu Photonics K.K., Hamamatsu City, Japan | TRS | Laser diodes | 3 (30) | 0.5 | 1 | StO2, atHb aO2Hb, aHHb | http://www.hamamatsu.com |
| Imagent | ISS Inc, Champaign, IL, USA | FD | Laser diodes | 2 | 50 | 16–512 | ΔtHb, ΔO2Hb, ΔHHb | http://www.iss.com |
| OxiplexTS | ISS Inc, Champaign, IL, USA | Multi-distance FD | Laser diodes | 2 (40) | 50 | 2 | SO2, atHb, aO2Hb, aHHb, reduced scattering coefficient | http://www.iss.com |
| moorVMS-NIRS | Moor Instruments Ltd, Axminster, UK | Multi-distance CW | LEDs | 2 (50) | 5 | 2 | SO2, auO2Hb auHHb | http://www.moor.co.uk |
| NIMOj,n | Nirox Srl, Borgosatollo (BS), Italy | Multi-distance CW | Laser diodes | 3 (30) | 40 | 2 | SmO2, atHb, aO2Hb, aHHb | http://www.nirox.it |
| BOM-L1TRW | Omegawave, Inc., Fuchu, Japan | Multi-distance CW | Laser diodes | 3 | 1 | 1 | StO2/autHb/auO2Hb/auHHb | http://www.omegawave.co.jp |

a absolute, au arbitrary units, Δ relative changes, tHb total haemoglobin concentration changes, CW continuous-wave spectroscopy, FD frequency-domain spectroscopy, HHb deoxyhaemoglobin concentration changes, LED light-emitting diode, NIRS near-infrared spectroscopy, O2Hb oxyhaemoglobin concentration changes, SD longest source–detector distance, SmO2/SO2/StO2/TOI/TSI muscle oxygen saturation %, TRS, time-resolved spectroscopy, wl wavelengths

aWearable NIRS system with wireless data transmission

bSmartphone-controllable system

cSystem no longer commercially available replaced by PortaLite MKII which samples six channels at 100 Hz

dAccelerometer is available on request

eSystem no longer commercially available replaced by; (a) PortaMon MkIII which samples six channels at 100 Hz, and exports six channels at 100 Hz; (b) Train.Red FYER (smartphone controllable wearable waterproof to IPX7 system with wireless data transmission, and accelerometer; https://train.red/) which samples and exports one channel at 10 Hz; (c) Train.Red PLUS (smartphone-controllable wearable waterproof to IPX7 with wireless data transmission, and accelerometer; https://train.red/) which samples six channels at 100 Hz, and exports one channel at 10 Hz [236]

fHb11 system no longer commercially available replaced by Oxy-Pro (Hb141; smartphone-controllable wearable sweat-proof 20 Hz system with wireless data transmission, accelerometer, and fat layer correction)

gSystem no longer commercially available

hWaterproof to IPX7

iWaterproof to IPX8

jFat-layer correction

kWater spectra with fixed water volume fraction included in the calibration model

lSystem no longer commercially available replaced by NIRO-200NX

mNot commercially available in European Union

nMeasured water absorption used to calculate optical path length