Table S1. Summary of apparatus qualification results

| Parameter | Acceptance Criteria |  | Results |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Mean ± Tolerance | Range of variation | Mean | Range of variation | Pass/ Fail |
| Internal orifice diameter | 15.00 ± 0.75 mm | ≤ 0.45 mm | 14.91 mm | 0.12 mm | Pass |
| Cell Capacity | 7.00 ± 0.35 mL | ≤0.21 mL | 7.94 mL | 0.04 mL | Fail |
| Receptor medium temperature | 32°C ± 1 °C | - | 32.3 °C | 0.5 °C | Pass |
| Stirring speed | 600 ± 60 rpm | ≤12 rpm | 612 rpm | 6.9 rpm | Pass |
| Dispensed sample volume | 200 ± 15 μL | - | 200 μL | 5 μL | Pass |

Table S2. Summary of performance verification test (PVT) results

| Parameter | Acceptance criteria | Results | Pass/ Fail |
| --- | --- | --- | --- |
| Intra-run variability | Intra-run CV run 1 (n=6) < 15% | 3.55% | Pass |
|  | Intra-run CV run 2 (n=6) < 15% | 1.71% | Pass |
| Inter-run variability | Inter-run CV between runs 1 and 2 (n=12) < 15% | 1.29% | Pass |
| ‘Sameness’ assessment | 90% CI between 75 – 133.33% | Lower limit: 100.59% | Pass |
|  |  | Upper limit: 101.36% |  |

Table S3. Summary of IVRT method validation results

| Parameter | Acceptance criteria | Results | Pass/ Fail |
| --- | --- | --- | --- |
| Membrane inertness | Recovery: ≥ 95% | 99.51 ± 0.14 % | Pass |
| Receptor medium solubility | 5-10-fold the maximum CLZ concentration | 3577.5 µg/mL = > 9-fold the maximum CLZ concentration | Pass |
| Linearity | R2 ≥ 0.95 | R2 = 0.99 | Pass |
| Precision and reproducibility | CV for intra- and inter run variability < 15% | Intra-run CV = 1.63% Inter-run CV = 3.98% | Pass |
| Robustness | Mean release rates of the runs with minor alterations deviate by < 15% of the nominal mean release rate | Lower temperature (30°C): CV = 5.78% Higher temperature (34°C): CV = 7.13% Lower stirring rate (540rpm): CV = 1.29% Higher stirring rate (660rpm): CV = 0.08% | Pass |
| Sensitivity | Mean release rate of 0.5% CLZ < 1% CLZ < 1.5% CLZ | 4.98μg/cm2/min1/2 < 8.00μg/cm2/min1/2 < 10.04μg/cm2/min1/2 | Pass |
| Specificity | R2 ≥ 0.9 | R2 = 0.96 | Pass |
| Selectivity | 90% CI for 0.5% and 1.5% test creams should fall outside the limits of 75 – 133.33% | 1% vs 0.5% CLZ creams: LL: 59.63% UL: 66.29% 1% vs 1.5% CLZ creams: LL: 119.12 UL: 134.01% | Pass Pass |
| Supplemental selectivity (inequivalence) | Q1, Q2 inequivalent creams: 90% CI should lie outside the limits of 75 – 133.33% | RLD (1% CLZ) vs. Q1, Q2 inequivalent Cream A (1% CLZ): LL: 36.28% UL: 40.25% | Pass |
| Supplemental selectivity (equivalence) | Q1, Q2 equivalent cream: 90% CI should lie within the limits of 75 – 133.33% | RLD (1% CLZ) vs. Q1, Q2 equivalent Cream B (1% CLZ): LL: 92.30% UL: 102.40% | Pass |
| Discrimination (inequivalence) | Negative controls: 90% CI for 1.4% and 0.7% test creams should lie outside the limits of 75 – 133.33% | 1% vs 1.4% CLZ creams: LL: 135.81% UL: 173.55% 1% vs 0.7% CLZ creams: LL: 71.53% UL: 75.57% | Pass Pass |
| Discrimination (equivalence) | Positive control: 90% CI for 1% test cream vs itself should lie within the limits of 75 – 133.33% | 1% vs 1% CLZ creams LL: 95.83% UL: 110.11% | Pass |

LL- Lower limit, UL- Upper limit