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Table SM1. Data compilation: 68Ga elution yields and 68Ge contents.
	Date
	Days post Calibration
	68Ga Elution Yield (%)
	% 68Ge in the elution

	5/29/2018
	41
	55.0
	0.000019

	5/30/2018
	42
	59.6
	0.000027

	5/31/2018
	43
	59.4
	0.000021

	6/1/2018
	44
	60.1
	0.000009

	6/4/2018
	47
	60.9
	0.000007

	6/4/2018
	47
	51.0
	0.000019

	6/5/2018
	48
	55.0
	

	6/6/2018
	49
	56.0
	

	6/7/2018
	50
	55.0
	0.000036

	6/11/2018
	54
	57.0
	0.0000265

	6/12/2018
	55
	57.0
	0.000014

	6/12/2018
	55
	59.7
	

	6/13/2018
	56
	55.0
	

	6/14/2018
	57
	57.0
	

	6/15/2018
	58
	57.0
	

	6/18/2018
	61
	62.8
	0.000025

	6/19/2018
	62
	57.0
	

	6/19/2018
	62
	61.2
	

	6/20/2018
	63
	58.0
	

	6/21/2018
	64
	59.0
	

	6/22/2018
	65
	59.0
	

	6/27/2018
	70
	63.7
	0.000026

	6/28/2018
	71
	60.0
	

	6/29/2018
	72
	59.0
	

	7/1/2018
	74
	63.1
	

	7/2/2018
	75
	58.0
	

	7/3/2018
	76
	59.0
	0.000005

	7/5/2018
	78
	53.0
	0.000006

	7/5/2018
	78
	65.0
	

	7/6/2018
	79
	57.0
	

	7/9/2018
	82
	58.0
	0.0000048

	7/10/2018
	83
	57.0
	

	7/11/2018
	84
	55.0
	0.000024

	7/11/2018
	84
	58.3
	0.0000076

	7/12/2018
	85
	58.0
	

	7/13/2018
	86
	58.0
	

	7/16/2018
	89
	61.3
	0.00002

	7/17/2018
	90
	58.0
	

	7/18/2018
	91
	58.0
	0.000028

	7/19/2018
	92
	58.0
	0.000035

	7/20/2018
	93
	59.0
	

	7/22/2018
	95
	61.7
	

	7/23/2018
	96
	57.0
	

	7/24/2018
	97
	63.7
	0.000028

	7/25/2018
	98
	58.0
	0.0000378

	7/26/2018
	99
	63.3
	

	7/27/2018
	100
	59.0
	

	7/30/2018
	103
	54.0
	0.000006

	7/30/2018
	103
	65.2
	

	7/31/2018
	104
	59.0
	

	8/1/2018
	105
	59.0
	0.000029

	8/1/2018
	105
	50.0
	

	8/2/2018
	106
	62.2
	0.000005

	8/3/2018
	107
	58.0
	

	8/3/2018
	107
	52.0
	

	8/6/2018
	110
	55.0
	0.0000318

	8/6/2018
	110
	61.0
	

	8/7/2018
	111
	59.0
	

	8/8/2018
	112
	59.0
	0.000005

	8/8/2018
	112
	60.3
	0.000011

	8/9/2018
	113
	59.0
	0.000014

	8/10/2018
	114
	60.0
	

	8/10/2018
	114
	55.0
	

	8/12/2018
	116
	61.6
	

	8/13/2018
	117
	58.0
	0.000017

	8/14/2018
	118
	58.0
	

	8/15/2018
	119
	61.8
	0.000005

	8/16/2018
	120
	59.0
	0.000042

	8/17/2018
	121
	59.0
	

	8/19/2018
	123
	61.3
	

	8/20/2018
	124
	59.0
	

	8/21/2018
	125
	59.0
	

	8/22/2018
	126
	57.0
	

	8/23/2018
	127
	53.0
	

	8/24/2018
	128
	59.0
	

	8/27/2018
	131
	57.0
	0.000038

	8/27/2018
	131
	65.0
	

	8/28/2018
	132
	59.0
	

	8/29/2018
	133
	62.3
	0.000021

	8/30/2018
	134
	59.0
	

	9/4/2018
	139
	64.3
	

	9/5/2018
	140
	59.0
	

	9/7/2018
	142
	55.0
	

	9/7/2018
	142
	63.6
	

	9/10/2018
	145
	54.0
	

	9/10/2018
	145
	64.1
	

	9/11/2018
	146
	59.0
	

	9/12/2018
	147
	59.0
	0.000019

	9/13/2018
	148
	59.0
	

	9/17/2018
	152
	55.0
	0.000014

	9/17/2018
	152
	63.0
	0.000023

	9/18/2018
	153
	60.0
	0.000021

	9/19/2018
	154
	60.0
	0.000023

	9/20/2018
	155
	64.9
	

	9/21/2018
	156
	59.0
	0.00002

	9/24/2018
	159
	54.0
	0.000026

	9/24/2018
	159
	64.6
	

	9/26/2018
	161
	63.1
	

	9/27/2018
	162
	56.0
	

	9/28/2018
	163
	57.0
	

	10/2/2018
	167
	61.5
	0.000036

	10/2/2018
	167
	61.6
	

	10/3/2018
	168
	57.0
	

	10/4/2018
	169
	57.0
	

	10/5/2018
	170
	58.0
	

	10/7/2018
	172
	61.1
	

	10/8/2018
	173
	57.0
	

	10/10/2018
	175
	56.0
	

	10/11/2018
	176
	56.0
	0.0000279

	10/12/2018
	177
	56.0
	

	10/15/2018
	180
	60.8
	

	10/18/2018
	183
	61.7
	0.000029

	10/19/2018
	184
	55.0
	

	10/24/2018
	189
	59.5
	

	10/25/2018
	190
	58.0
	0.000022

	10/29/2018
	194
	54.0
	0.000021

	10/29/2018
	194
	59.9
	

	10/30/2018
	195
	61.5
	

	10/31/2018
	196
	47.0
	0.000024

	10/31/2018
	196
	48.0
	

	11/1/2018
	197
	58.0
	0.000047

	11/1/2018
	197
	60.1
	

	11/2/2018
	198
	57.0
	

	11/5/2018
	201
	61.3
	0.000039

	11/5/2018
	201
	61.1
	

	11/6/2018
	202
	59.0
	0.000025

	11/7/2018
	203
	59.0
	0.000026

	11/7/2018
	203
	59.4
	

	11/8/2018
	204
	59.0
	0.000036

	11/9/2018
	205
	59.0
	0.000028

	11/12/2018
	208
	55.0
	0.000034

	11/12/2018
	208
	62.5
	

	11/13/2018
	209
	58.0
	0.000032

	11/14/2018
	210
	59.0
	0.000031

	11/14/2018
	210
	46.0
	

	11/19/2018
	215
	55.0
	0.0000295

	11/19/2018
	215
	62.0
	

	11/20/2018
	216
	57.0
	0.000028

	11/21/2018
	217
	59.0
	

	11/26/2018
	222
	54.0
	0.000025

	11/26/2018
	222
	62.6
	

	11/27/2018
	223
	60.0
	

	11/28/2018
	224
	59.0
	0.000035

	11/28/2018
	224
	50.0
	

	11/29/2018
	225
	59.0
	

	11/30/2018
	226
	59.0
	

	12/3/2018
	229
	55.0
	

	12/3/2018
	229
	62.2
	

	12/4/2018
	230
	58.0
	0.000031

	12/4/2018
	230
	62.7
	

	12/5/2018
	231
	58.0
	

	12/5/2018
	231
	48.0
	

	12/6/2018
	232
	57.0
	

	12/7/2018
	233
	61.8
	0.000046

	12/7/2018
	233
	61.3
	

	12/10/2018
	236
	51.0
	

	12/11/2018
	237
	57.0
	

	12/14/2018
	240
	58.0
	0.000035

	12/14/2018
	240
	49.0
	

	12/18/2018
	244
	56.0
	0.000038

	1/2/2019
	259
	55.0
	0.000039

	1/2/2019
	259
	60.4
	0.000039

	1/3/2019
	260
	61.4
	0.00003

	1/4/2019
	261
	59.0
	0.000036

	1/7/2019
	264
	51.0
	0.000053

	1/7/2019
	264
	61.0
	

	1/8/2019
	265
	57.0
	0.000058

	1/9/2019
	266
	58.0
	0.000054

	1/10/2019
	267
	57.0
	

	1/10/2019
	267
	54.0
	

	1/11/2019
	268
	57.0
	0.000038

	1/14/2019
	271
	52.0
	0.000054

	1/14/2019
	271
	62.7
	

	1/15/2019
	272
	56.0
	

	1/16/2019
	273
	62.8
	

	1/17/2019
	274
	43.0
	

	1/18/2019
	275
	57.0
	

	1/22/2019
	279
	54.0
	

	1/22/2019
	279
	61.9
	

	1/23/2019
	280
	57.0
	0.000071

	1/24/2019
	281
	58.0
	

	1/25/2019
	282
	58.0
	0.000057

	1/28/2019
	285
	54.0
	0.000043

	1/28/2019
	285
	65.2
	

	1/29/2019
	286
	63.8
	

	1/30/2019
	287
	56.0
	

	1/31/2019
	288
	62.8
	

	2/1/2019
	289
	57.0
	0.000058

	2/4/2019
	292
	54.0
	

	2/4/2019
	292
	63.6
	

	2/5/2019
	293
	58.0
	0.000066

	2/6/2019
	294
	58.0
	

	2/7/2019
	295
	59.0
	

	2/11/2019
	299
	52.0
	0.000062

	2/11/2019
	299
	61.6
	

	2/12/2019
	300
	62.8
	

	2/13/2019
	301
	58.0
	

	2/14/2019
	302
	58.0
	0.000059

	2/14/2019
	302
	61.5
	0.000044

	2/15/2019
	303
	56.0
	

	2/20/2019
	308
	59.0
	

	2/21/2019
	309
	58.0
	0.000066

	2/22/2019
	310
	59.0
	

	2/22/2019
	310
	53.0
	

	2/25/2019
	313
	55.0
	

	2/25/2019
	313
	64.3
	

	2/27/2019
	315
	60.0
	0.000033

	2/28/2019
	316
	51.0
	

	2/28/2019
	316
	65.2
	

	3/4/2019
	320
	55.0
	0.000039

	3/5/2019
	321
	59.0
	0.000037

	3/6/2019
	322
	60.0
	

	3/7/2019
	323
	60.0
	0.000048

	3/7/2019
	323
	63.9
	

	3/8/2019
	324
	60.0
	0.000035

	3/13/2019
	329
	63.8
	

	3/14/2019
	330
	60.0
	0.000047

	3/18/2019
	334
	54.0
	

	3/18/2019
	334
	61.4
	



Table SM2. Results from ICP-MS for metal contamination.
	Metal contaminant concentration per elution number in ng/ml (ppb)

	[bookmark: _Hlk8799676]Elution #
	Al
	Cr
	Mn
	Fe
	Co

	41
	4.17±2.15
	1.20±0.90
	0.00±0.00
	25.93±3.14
	0.00±0.00

	77
	25.99±7.91
	0.00±0.00
	0.00±0.00
	22.54±1.73
	0.13±0.47

	111
	14.17±3.64
	1.24±1.74
	0.19±0.42
	32.82±13.99
	0.11±0.28

	200
	20.01±8.81
	6.55±2.26
	0.74±0.50
	43.41±4.69
	0.86±0.98

	322
	7.55±9.71
	0.00±0.00
	0.00±0.00
	22.15±19.27
	0.00±0.00

	ALL
	14.38±8.91
	1.80±2.73
	0.19±0.32
	29.37±8.94
	0.22±0.36

	Elution #
	Ni
	Cu
	Zn
	Ga
	Ge

	41
	5.00±8.83
	7.80±4.06
	52.26±5.83
	1.03±0.30
	0.06±0.01

	77
	2.18±0.82
	6.35±0.77
	49.10±9.16
	3.75±0.67
	0.06±0.01

	111
	9.18±6.10
	5.68±0.66
	69.35±49.43
	1.49±1.21
	0.06±0.01

	200
	17.15±3.50
	5.83±0.50
	57.81±26.14
	1.69±0.99
	0.07±0.01

	322
	5.67±2.98
	3.35±5.34
	10.71±13.35
	0.92±0.68
	0.03±0.06

	ALL
	7.84±5.77
	5.80±1.61
	47.85±22.15
	1.77±1.15
	0.06±0.02

	Metal contaminant concentration per elution fraction in ng/ml (ppb)

	Fraction
	Al
	Cr
	Mn
	Fe
	Co

	1
	11.30±17.94
	1.12±4.46
	0.11±1.11
	20.42±21.38
	0.25±0.45

	2
	14.38±6.70
	1.79±1.53
	0.00±0.00
	38.04±14.22
	0.00±0.00

	3
	14.50±8.26
	1.61±3.01
	0.00±0.00
	29.53±9.22
	0.70±1.26

	4
	14.38±12.74
	1.43±2.88
	0.00±0.00
	28.05±8.53
	0.18±0.36

	5
	15.46±10.44
	1.49±2.39
	0.00±0.00
	29.70±8.46
	0.00±0.00

	6
	16.24±8.24
	2.59±4.27
	0.00±0.00
	30.46±11.88
	0.00±0.00

	ALL
	14.38±16.82
	1.80±5.01
	0.19±0.43
	29.37±56.28
	0.22±2.73

	Fraction
	Ni
	Cu
	Zn
	Ga
	Ge

	1
	13.17±10.92
	5.20±8.34
	46.82±46.90
	1.65±1.36
	0.03±0.07

	2
	7.37±5.48
	5.97±0.47
	57.90±31.21
	2.29±1.26
	0.06±0.01

	3
	7.43±7.39
	5.33±0.25
	54.05±36.32
	2.35±1.80
	0.06±0.01

	4
	6.00±5.58
	5.50±0.44
	40.48±17.08
	1.51±1.20
	0.05±0.02

	5
	5.97±5.13
	5.83±0.63
	51.07±38.44
	1.65±1.20
	0.06±0.01

	6
	7.07±8.04
	6.99±0.51
	36.76±18.76
	1.20±1.17
	0.06±0.01

	ALL
	7.84±26.92
	5.80±6.49
	47.85±81.02
	1.77±4.54
	0.06±0.12



Table SM3. Compilation of [68Ga]Ga-PSMA-11 Syntheses.
	Synthesis #
	Date
	Day Post-calibration
	Radiochemical Yield
	Final Product (MBq)

	1
	6/4/2018
	46
	91%
	1494.8

	2
	6/5/2018
	47
	90%
	1620.6

	3
	6/6/2018
	48
	88%
	1616.9

	4
	6/7/2018
	49
	90%
	1639.1

	5
	6/11/2018
	53
	90%
	1672.4

	6
	6/12/2018
	54
	90%
	1653.9

	7
	6/13/2018
	55
	90%
	1613.2

	8
	6/14/2018
	56
	91%
	1668.7

	9
	6/15/2018
	57
	91%
	1642.8

	10
	6/19/2018
	61
	89%
	1613.2

	11
	6/20/2018
	62
	89%
	1631.7

	12
	6/21/2018
	63
	89%
	1642.8

	13
	6/22/2018
	64
	90%
	1672.4

	14
	6/28/2018
	70
	92%
	1694.6

	15
	6/29/2018
	71
	90%
	1665

	16
	7/2/2018
	74
	87%
	1535.5

	17
	7/3/2018
	75
	89%
	1594.7

	18
	7/6/2018
	78
	92%
	1583.6

	19
	7/9/2018
	81
	90%
	1542.9

	20
	7/10/2018
	82
	87%
	1494.8

	21
	7/11/2018
	83
	90%
	1454.1

	22
	7/12/2018
	84
	90%
	1546.6

	23
	7/13/2018
	85
	89%
	1509.6

	24
	7/17/2018
	89
	51%
	873.2

	25
	7/18/2018
	90
	92%
	1550.3

	26
	7/19/2018
	91
	90%
	1517

	27
	7/20/2018
	92
	90%
	1524.4

	28
	7/23/2018
	95
	93%
	1505.9

	29
	7/25/2018
	97
	91%
	1487.4

	30
	7/27/2018
	99
	88%
	1446.7

	31
	7/30/2018
	102
	83%
	1239.5

	32
	7/31/2018
	103
	62%
	991.6

	33
	8/1/2018
	104
	88%
	1202.5

	34
	8/3/2018
	106
	91%
	1450.4

	35
	8/6/2018
	109
	95%
	1420.8

	36
	8/7/2018
	110
	86%
	1357.9

	37
	8/8/2018
	111
	88%
	1406

	38
	8/9/2018
	112
	90%
	1431.9

	39
	8/10/2018
	113
	92%
	1468.9

	40
	8/10/2018
	113
	92%
	1376.4

	41
	8/13/2018
	116
	89%
	1406

	42
	8/14/2018
	117
	88%
	1409.7

	43
	8/16/2018
	119
	91%
	1450.4

	44
	8/17/2018
	120
	91%
	1428.2

	45
	8/20/2018
	123
	93%
	1450.4

	46
	8/21/2018
	124
	88%
	1398.6

	47
	8/22/2018
	125
	90%
	1339.4

	48
	8/23/2018
	126
	89%
	1239.5

	49
	8/24/2018
	127
	90%
	1424.5

	50
	8/27/2018
	130
	91%
	1383.8

	51
	8/28/2018
	131
	89%
	1391.2

	52
	8/30/2018
	133
	89%
	1361.6

	53
	9/5/2018
	139
	91%
	1376.4

	54
	9/7/2018
	141
	89%
	1269.1

	55
	9/10/2018
	144
	88%
	1209.9

	56
	9/11/2018
	145
	89%
	1343.1

	57
	9/12/2018
	146
	89%
	1339.4

	58
	9/13/2018
	147
	90%
	1357.9

	59
	9/17/2018
	151
	90%
	1250.6

	60
	9/18/2018
	152
	90%
	1372.7

	61
	9/19/2018
	153
	89%
	1346.8

	62
	9/21/2018
	155
	90%
	1328.3

	63
	9/24/2018
	158
	90%
	1206.2

	64
	9/27/2018
	161
	89%
	1254.3

	65
	9/28/2018
	162
	90%
	1258

	66
	10/3/2018
	167
	91%
	1265.4

	67
	10/4/2018
	168
	89%
	1221

	68
	10/5/2018
	169
	90%
	1232.1

	69
	10/8/2018
	172
	91%
	1239.5

	70
	10/10/2018
	174
	90%
	1206.2

	71
	10/11/2018
	175
	90%
	1169.2

	72
	10/12/2018
	176
	87%
	1150.7

	73
	10/19/2018
	183
	93%
	1172.9

	74
	10/25/2018
	189
	89%
	1165.5

	75
	10/29/2018
	193
	90%
	1087.8

	76
	10/31/2018
	195
	89%
	936.1

	77
	10/31/2018
	195
	90%
	965.7

	78
	11/1/2018
	196
	89%
	1147

	79
	11/2/2018
	197
	90%
	1139.6

	80
	11/6/2018
	201
	89%
	1147

	81
	11/7/2018
	202
	91%
	1154.4

	82
	11/8/2018
	203
	91%
	1154.4

	83
	11/9/2018
	204
	90%
	1154.4

	84
	11/12/2018
	207
	90%
	1054.5

	85
	11/13/2018
	208
	90%
	1124.8

	86
	11/14/2018
	209
	90%
	1128.5

	87
	11/14/2018
	209
	94%
	921.3

	88
	11/19/2018
	214
	91%
	1069.3

	89
	11/20/2018
	215
	90%
	1098.9

	90
	11/21/2018
	216
	90%
	1139.6

	91
	11/26/2018
	221
	89%
	1021.2

	92
	11/27/2018
	222
	90%
	1124.8

	93
	11/28/2018
	223
	90%
	1121.1

	94
	11/28/2018
	223
	89%
	932.4

	95
	11/29/2018
	224
	89%
	1102.6

	96
	11/30/2018
	225
	92%
	1135.9

	97
	12/3/2018
	228
	91%
	1036

	98
	12/4/2018
	229
	90%
	1091.5

	99
	12/5/2018
	230
	88%
	1043.4

	100
	12/5/2018
	230
	91%
	895.4

	101
	12/6/2018
	231
	89%
	1039.7

	102
	12/10/2018
	235
	88%
	910.2

	103
	12/11/2018
	236
	91%
	1069.3

	104
	12/14/2018
	239
	91%
	1058.2

	105
	12/14/2018
	239
	91%
	906.5

	106
	12/18/2018
	243
	89%
	1006.4

	107
	1/2/2019
	258
	93%
	984.2

	108
	1/4/2019
	260
	89%
	984.2

	109
	1/7/2019
	263
	89%
	847.3

	110
	1/8/2019
	264
	88%
	939.8

	111
	1/9/2019
	265
	90%
	958.3

	112
	1/10/2019
	266
	88%
	943.5

	113
	1/10/2019
	266
	89%
	906.5

	114
	1/11/2019
	267
	85%
	906.5

	115
	1/14/2019
	270
	86%
	836.2

	116
	1/15/2019
	271
	87%
	902.8

	117
	1/17/2019
	273
	89%
	710.4

	118
	1/18/2019
	274
	85%
	880.6

	119
	1/22/2019
	278
	88%
	873.2

	120
	1/23/2019
	279
	89%
	928.7

	121
	1/24/2019
	280
	87%
	921.3

	122
	1/25/2019
	281
	86%
	906.5

	123
	1/28/2019
	284
	88%
	858.4

	124
	1/30/2019
	286
	90%
	910.2

	125
	2/1/2019
	288
	89%
	906.5

	126
	2/4/2019
	291
	89%
	862.1

	127
	2/5/2019
	292
	87%
	891.7

	128
	2/6/2019
	293
	88%
	888

	129
	2/7/2019
	294
	87%
	899.1

	130
	2/11/2019
	298
	86%
	784.4

	131
	2/13/2019
	300
	87%
	865.8

	132
	2/14/2019
	301
	88%
	876.9

	133
	2/15/2019
	302
	88%
	851

	134
	2/20/2019
	307
	89%
	873.2

	135
	2/21/2019
	308
	88%
	854.7

	136
	2/22/2019
	309
	88%
	862.1

	137
	2/22/2019
	309
	89%
	791.8

	138
	2/25/2019
	312
	86%
	799.2

	139
	2/27/2019
	314
	90%
	902.8

	140
	2/28/2019
	315
	91%
	765.9

	141
	3/4/2019
	319
	90%
	817.7

	142
	3/5/2019
	320
	88%
	847.3

	143
	3/6/2019
	321
	91%
	880.6

	144
	3/7/2019
	322
	88%
	858.4

	145
	3/8/2019
	323
	90%
	869.5

	146
	3/14/2019
	329
	89%
	851

	147
	3/18/2019
	333
	90%
	777

	148
	3/21/2019
	336
	92%
	828.8

	149
	3/22/2019
	337
	87%
	788.1

	150
	3/25/2019
	340
	84%
	715.21

	151
	3/26/2019
	341
	88%
	821.4

	152
	3/29/2019
	344
	89%
	780.7

	153
	4/1/2019
	347
	88%
	758.5

	154
	4/2/2019
	348
	89%
	784.4

	155
	4/4/2019
	350
	87%
	777

	156
	4/5/2019
	351
	87%
	762.2

	157
	4/9/2019
	355
	86%
	706.7

	158
	4/10/2019
	356
	88%
	758.5

	159
	4/11/2019
	357
	88%
	680.8

	160
	4/16/2019
	362
	86%
	740

	161
	4/17/2019
	363
	87%
	754.8

	162
	4/19/2019
	365
	89%
	773.3

	163
	4/22/2019
	368
	88%
	695.6

	164
	4/23/2019
	369
	88%
	747.4

	165
	4/24/2019
	370
	87%
	736.3

	166
	4/24/2019
	370
	89%
	669.7

	167
	4/25/2019
	371
	88%
	754.8

	168
	4/25/2019
	371
	86%
	703.37

	169
	4/26/2019
	372
	90%
	762.2

	170
	4/29/2019
	375
	89%
	743.7

	171
	4/30/2019
	376
	88%
	740

	172
	5/1/2019
	377
	86%
	714.1

	173
	5/2/2019
	378
	89%
	743.7

	174
	5/2/2019
	378
	88%
	658.6

	175
	5/7/2019
	383
	85%
	625.3



Table SM4. Drug product release criteria for [68Ga]Ga-PSMA-11
	Test
	Acceptance Criteria
	Results

	Appearance
	Clear, colorless solution. free from particulate matter
	conforms

	Radionuclidic identity
	62-74 min
	66-70 min

	Radionuclidic purity
	511 ± 51 keV
	conforms

	Radiochemical identity
	RT of [68Ga]Ga-PSMA-11should be ± 10% the RT of Ga-PSMA-11 std
	RT std: 7.16 min
RT FDP: 7.37 min
Δ RT: 2.93%

	Radiochemical purity
	≥ 96%
	> 99%

	Chemical purity
	≥ 90%
	> 99%

	68Ge impurity
	< 0.001%
	< 0.0000069%

	pH
	4.5-8.0
	5.0-5.5

	Bubble Point
	≥ 50 psi
	52-58 psi

	LaL (pyrogenicity) test
	<29 EU/mL
	<5.00 EU/mL

	sterility
	Pass (no microbial growth)
	pass



Figure SM5. Typical QC Chromatogram for [68Ga]Ga-PSMA-11.
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