Additional File 1

Figures

Figure S1. External calibration curve for cadmium determination by GFAAS using the P1 furnace program.

Figure S2: Drying temperature optimization for ASW

Figure S3: Pyrolysis temperature optimization for ASW

Figure S4: Optimization of atomization temperature for ASW

Figure S5: Optimization of the atomization of the atomization temperature for ABW

Impact of P2 programme on CRMs

Black line- Analyte absorbance during furnace program P1

Blue line- Analyte absorbance during furnace program P2

Figure S6: Impact of the P2 programme on CRMs.

Tables

Table S1:  P2 Temperature Progam for ASW

| Step | Name | Temp.[°C] | Ramp[°C/s] | Hold [s] | Time [s] | Gas |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Int | Add |  |
| 1 | Drying | 80 | 6 | 20 | 27.2 | Max |  | Stop |
| 2 | Drying | 90 | 3 | 20 | 23.3 | Max |  | Stop |
| 3 | Drying | 150 | 5 | 10 | 14.0 | Max |  | Stop |
| 4 | Pyrolysis | 350 | 50 | 20 | 24.8 | Max |  | Stop |
| 5 | Pyrolysis | 800 | 300 | 10 | 11.5 | Max |  | Stop |
| 6 | AZ* | 800 | 0 | 6 | 6.0 | Stop |  | Stop |
| 7 | Atomize | 1400 | 1500 | 3 | 3.4 | Stop |  | Stop |
| 8 | Clean | 2450 | 500 | 4 | 6.1 | Max |  | Stop |

Table S2:  P3 Temperature Progam for ABW

| Step | Name | Temp.[°C] | Ramp[°C/s] | Hold [s] | Time [s] | Gas |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  | Int | Add |  |
| 1 | Drying | 80 | 6 | 20 | 27.2 | Max |  | Stop |
| 2 | Drying | 90 | 3 | 20 | 23.3 | Max |  | Stop |
| 3 | Drying | 150 | 5 | 10 | 14.0 | Max |  | Stop |
| 4 | Pyrolysis | 350 | 50 | 20 | 24.8 | Max |  | Stop |
| 5 | Pyrolysis | 800 | 300 | 10 | 11.5 | Max |  | Stop |
| 6 | AZ* | 800 | 0 | 6 | 6.0 | Stop |  | Stop |
| 7 | Atomize | 1500 | 1500 | 3 | 3.4 | Stop |  | Stop |
| 8 | Clean | 2450 | 500 | 4 | 6.1 | Max |  | Stop |

| Table S3: External Calibration using CRMs under P1 program |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Concentration µg/L | Absorption Mean |  |  |  |  | SD | RSD % |
|  | Trial 1 | Trial 2 | Trial 3 |  |  |  |  |
| 0 | -0.005 | -0.005 | -0.005 | -0.005 |  | 0.0000 | 4 |
| 2.5 | 0.052 | 0.058 | 0.059 | 0.056 |  | 0.0037 | 7 |
| 5 | 0.120 | 0.107 | 0.114 | 0.113 |  | 0.0065 | 6 |
| 10 | 0.230 | 0.237 | 0.230 | 0.233 |  | 0.0040 | 2 |

| Table S4: Drying Temperature optimization for ASW |  |  |  |
| --- | --- | --- | --- |
| Drying Temperature (°C) | Mean Absorbance | Standard Deviation | RSD (%) |
| 110 | 0.206 | 0.0097 | 5 |
| 150 | 0.233 | 0.0096 | 4 |
| 200 | 0.200 | 0.0001 | 0 |

| Table S5: Pyrolysis Temperature Optimization for ASW |  |  |  |
| --- | --- | --- | --- |
| Pyrolysis Temperature (°C) | Mean Absorbance | Standard Deviation | RSD (%) |
| 700 | 0.148 | 0.0085 | 6 |
| 800 | 0.234 | 0.0082 | 3 |
| 900 | 0.210 | 0.0158 | 8 |
| 1000 | 0.183 | 0.0070 | 4 |

| Table S6: Atomization Temperature Optimization for ASW |  |  |  |
| --- | --- | --- | --- |
| Atomization Temperature (°C) | Mean Absorbance | Standard Deviation | RSD (%) |
| 1200 | 0.008 | 0.0009 | 12 |
| 1300 | 0.050 | 0.0012 | 2 |
| 1400 | 0.234 | 0.0121 | 5 |
| 1500 | 0.183 | 0.0070 | 4 |

Table S7: Atomization Temperature Optimization for ABW

| Atomization Temperature ⁰C | ABW 10 µg/L Cd | Mean Absorbance | Standard Deviation | RSD % |
| --- | --- | --- | --- | --- |
| 1300 | 1.95 | 0.0434 | 0.0017 | 4 |
| 1400 | 6.78 | 0.1565 | 0.0351 | 22 |
| 1500 | 9.84 | 0.2283 | 0.0017 | 1 |
| 1600 | 0.02 | <LOD | - | - |

Table S8: GAPI comparison

| GAPI Parameter for Greenness level | Optimized GFAAS |  | Tetrahydroborate Reductive Precipitations (Nakashima et al. 1987) (23) |  | Preconcentration on Silica-Immobilized 8-Hydroxyquinoline (Nakashima et al. 1988) (24) |  | Solvent Extraction with PAPT (Lee et al. 1995) (25) |  | Dynamically Coated Column (C18) (Akatsuka et al. 1998) (26) |  | Automated Online Preconcentration with APDC and C18 Column (Colbert et al. 1998) (27) |  | Solid Phase Extraction (SPE) with Iminodiacetic Resin (Grotti et al. 2002) (28) |  | Solvent Extraction with DDTC + CPE (Yuan et al. 2004) (29) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Preparation. | Immediate filtration, acidification (ultrapure nitric acid) |  | Reductive Precipitation. |  | Adsorption on silica-immobilized sorbent. |  | Solvent extraction. |  | Dynamic coating on C18 column. |  | Online preconcentration. |  | Solid phase extraction. |  | Solvent extraction + CPE. |  |
| Reagents used | No any reagent used |  | Tetra hydroborate. |  | 8-Hydroxyquinoline. |  | Picolinaldehyde-4-phenyl-3-thiosemicarbazone (PAPT) |  | Coating reagents. |  | APDC. |  | Iminodiacetic Resin. |  | Diethyldithiocarbamate (DDTC), Triton X-114. |  |
| Solvent Consumption | Minimal Eco-friendly. |  | MinimalEco-friendly . |  | Minimal Eco-friendly . |  | Uses hazardous organic solvents. |  | Minimal Eco-friendly |  | Minimal Eco-friendly . |  | Minimal Eco-friendly |  | Moderate (combination of solvent extraction and CPE). |  |
| Waste Generation | Minimal hazardous waste |  | Precipitate waste. |  | Sorbent waste. |  | Hazardous organic solvent nonbiodegradable extraction media |  | Column waste. |  | Reagent and column waste. |  | Nonbiodegradable media |  | Mixed waste (solvents, surfactants). |  |
| Energy Consumption | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy . |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  |
| Matrix Tolerability | High-TDS tolerance; optimized for interferences |  | Effective, but requires precise pH control. |  | High sensitivity; pH control needed. |  | Struggles with high TDS, interferences. |  | Requires precise coating control. |  | Prone to interference from other metals. . |  | Influenced by matrix complexity. |  | High sensitivity; |  |
| Safety | Reduced hazardous solvents |  | Standard chemical handling. |  | Standard handling. . |  | Hazardous organic solvents, poses. |  | Standard chemical handling. |  | Standard chemical handling. |  | Stndard chemical handling. |  | Uses hazardous DDTC and Triton X-114. |  |
| Overall Suitability/GAPI Score | High |  | Moderate |  | Moderate |  | Low. |  | Moderate. |  | Moderate. |  | Moderate. |  | Moderate. |  |

| GAPI Parameter for Greenness level | CPE with Triton X-114 and 5-Br-PADAP (Xiao et al. 2007) (30) |  | Ultrasound-Assisted DLLME (Tang and Zheng, 2005) (31) |  | Hollow Fiber SLM Extraction (Peng et al. 2007) (32) |  | CPE Coupled with DLLME (Yazdi and Es’haghi 2011) (33) |  |  | Deep Eutectic Solvents (DES) in DLLME (Zhang et al. 2016) (34) |  | SPE with Modified Silica Gel (Yuan et al. [26], Moghimi & Yari 2019) (35) |  | Ionic Liquid Solvent Bar Micro-Extraction (IL-SBME) (Herce-Sesa, López-López, and Moreno, 2017) (36) |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Sample Preparation. | Cloud Point Extraction (CPE). |  | DLLME via ultrasound. |  | SLM Extraction. |  |  | CPE + DLLME. |  | DLLME using DES. |  | Solid phase extraction. |  | Microextraction with ionic liquid solvent bar. |  |
| Reagents used | Triton X-114 Br-PADAP. (safer surfactants). |  | Organic solvents, dispersers. |  | Organic solvents in membrane. |  |  | Organic solvents, surfactants. |  | Deep Eutectic Solvents (DES). |  | Modified Silica Gel. |  | Ionic liquids. |  |
| Solvent Consumption | Low. |  | Low. |  | Low. |  |  | Low. |  | Very Low. |  | Very Low. |  | Very Low.. |  |
| Waste Generation | Environmentally friendlier surfactants. |  | Reduced, but still some organic waste. |  | Membrane waste. |  |  | Mixed waste. |  | Reduced, more biodegradable. |  | Reduced, but sorbent regeneration issues.. |  | Reduced, but involves ionic liquid disposal. |  |
| Energy Consumption | Heating for cloud point.. |  | Ultrasound energy. |  | Graphite furnace use. High Energy |  |  | Heating for cloud point, vortex energy. |  | Heating for DES preparation/extraction. |  | Graphite furnace use. High Energy |  | Graphite furnace use. High Energy |  |
| Matrix Tolerability | Needs better selectivity. Moderate Greenness. |  | Variable ultrasonic settings. |  | High enrichment factor; needs further validation. |  |  | Effective but complex. |  | Environmentally friendly; variable recovery rates; application in high-TDS aaremains underexplored. |  | Effective for trace levels; regeneration issues. |  | High sensitivity and selectivity; complex ionic liquid handling and matrix interference potential. |  |
| Safety | Environmentally friendly. |  | Reduced hazardous solvents |  | Not specified. |  |  | Not specified. Moderate Greenness. |  | "Less toxic" and "more biodegradable". |  | Standard chemical handling. |  | Complex ionic liquid handling. |  |
| Overall Suitability/ GAPI Score | High |  | Moderate |  | Moderate |  |  | Moderate. |  | High Potential (promising but still needs validation for high-TDS routine use). |  | Moderate. |  | Moderate |  |